## Slide 1: Bunch length monitoring with Electro-optical sampling at CLEAR
E. Senes, A. Schloegelhofer, S. Mazzoni, T. Lefevre

## Slide 2
The CLEAR facility
Past implementations
Experiment proposal
1
2
3
Discussion
4

## Slide 3: The CLEAR facility

## Slide 4: CLEAR@CERN
Electron test beamline for internal and external users
Energy 60-220 MeV
Beam parameters
CERN-PHOTO-201711-271-18
Charge 0.01 – 1.5 nC
Bunch length (1σ) 100 fs – 5 ps
Rep. rate 0.8 – 10 Hz
Full details in D. Gamba et al., The CLEAR user facility at CERN, NIMA 909 (2018) pg. 480-483

## Slide 5: CLEAR@CERN
Interesting building blocks:
DUMP
eGUN
3x ACCELERATING SECTIONS
RF DEFLECTOR
SPECTROMETER
USER STATION #1
PREVIOUS TEST AREA

## Slide 6: CLEAR@CERN
Interesting building blocks:
DUMP
eGUN
3x ACCELERATING SECTIONS
RF DEFLECTOR
SPECTROMETER
USER STATION #1
PREVIOUS TEST AREA
Deflecting cavity + screen
Destructive measurement
Resolution ~100 fs sigma
Existing bunch length measurement:

## Slide 7: CLEAR@CERN
Interesting building blocks:
DUMP
eGUN
3x ACCELERATING SECTIONS
RF DEFLECTOR
SPECTROMETER
USER STATION #1
PREVIOUS TEST AREA
In just few slides …
Previous EOS measurements in vacuum

## Slide 8: CLEAR@CERN
Interesting building blocks:
DUMP
eGUN
3x ACCELERATING SECTIONS
RF DEFLECTOR
SPECTROMETER
USER STATION #1
PREVIOUS TEST AREA
IN-AIR TEST AREA
In air, at the end of the beamline
Proposed EOS new measurement

## Slide 9: In-vacuum implementation
With some slides taken from M. Bergamaschi, others from S. Mazzoni, a few stolen from I. Gorgysyan … I even did some by myself !

## Slide 10: In-vacuum system
B. Steffen et al. Proc. DIPAC09 (2009)
Laser reflected from the back surface of the crystal (non-metallic coating)
Redesigned chamber installed in 2019

## Slide 11: In-vacuum system
B. Steffen et al. Proc. DIPAC09 (2009)
Laser reflected from the back surface of the crystal (non-metallic coating)
Redesigned chamber installed in 2019

## Slide 12: Results
First measurement in 2020
M. Bergamaschi
Resolution limited by the 3 mm-thick crystal (0.7 ps)
ZnTe crystal

## Slide 13: In-air setup proposal

## Slide 14: Why a new design ?
The in-vacuum design is quite difficult to align.
Need for a setup that allows for DEOS.
It is now possible to leverage the recently installed fibre infrastructure.

## Slide 15: Laser system
The 780 nm laser system produces a chirped pulse of ~10 ps (1σ)
Pulse avg. power < 200 mW
The pulse is sent to the setup in a PM fibre. The pulse stretching in the fibre is to be measured.
Pulse repetition 13 ns
LASER UNIT
DELAY
STAGE
GRATINGS
POCKELS
CELL
IN-FIBRE
COUPLING
IN-FIBRE
LASER

## Slide 16: Bunch length measurement setup
CRYSTAL
LASER
BEAM
POLARISER
COLLIMATOR
COLLIMATOR
HWP
IN-FIBRE POLARIZING BEAM SPLITTER
SM
POL1
POL2
LASER SYSTEM
PM FIBRE
Laser polarisation defined by the PM fibre orientation (and launch collimator rotation)
Launch collimator to keep the beam size at the crystal of ~1-2 mm
HWP
QWP
Laser launch
HWP and polariser to set the polarisation at the crystal

## Slide 17: Bunch length measurement setup
CRYSTAL
LASER
BEAM
POLARISER
COLLIMATOR
COLLIMATOR
HWP
IN-FIBRE POLARIZING BEAM SPLITTER
SM
POL1
POL2
LASER SYSTEM
PM FIBRE
OTR/YAG screen mount for beam alignment.
HWP
QWP
EO crystals
3D printed holder for 3 crystals + beam screen.
BEAM
LASER
EO
CRYSTALS
OTR
ZnTe crystals in inventory (110)
(110) cut
4, 2, 1, 0.5, 0.3 mm thickness
VIEW FROM HWP
CRYSTAL
BEAM

## Slide 18: Bunch length measurement setup
CRYSTAL
LASER
BEAM
POLARISER
COLLIMATOR
COLLIMATOR
HWP
IN-FIBRE POLARIZING BEAM SPLITTER
SM
POL1
POL2
LASER SYSTEM
PM FIBRE
Motorised QWP and HWP to tune the fibre coupling.
QWP
HWP
Coupling back to fibre
Polarizing beam splitter in-fibre splits the 2 orthogonal components.
Multimode fibres from the beam splitter output

## Slide 19: How does it look like
LASER
CRYSTAL ARRAY
3D STAGE
LAUNCH
COLLIMATOR
HWP
FIXED POLARISER
MOTORISED
QWP
MOTORISED
HWP
2+1D
MOTORISED
COLLIMATOR
IN-FIBRE POLARISING BEAM SPLITTER
TO OPTICS LAB
FROM
OPTICS
LAB
BEAM

## Slide 20: Timing alignment setup
CRYSTAL
LASER
BEAM
POLARISER
COLLIMATOR
COLLIMATOR
HALF WAVEPLATE
IN-FIBRE POLARIZING BEAM SPLITTER
SM
POL1
POL2
LASER SYSTEM
PM FIBRE
90/10 IN-FIBRE SPLITTER
TIMING FIBRE
10% of the laser is split to a multimode fibre intercepting the beam at 45°.
The beam generates Cherenkov light intercepting the fibre.
A 12 GHz photodiode is used to check align the laser and beam pulses. Streak camera is also an option.

## Slide 21: How does it look like
LASER
CRYSTAL ARRAY
VERTICAL
STAGE
LAUNCH
COLLIMATOR
90/10
SPLITTER
TO OPTICS LAB
BEAM
FROM
OPTICS
LAB
MULTIMODE
FIBRE AT 45°
Vertical stage up:
The fibre produces Cherenkov light to the 12 GHz photodiode
Vertical stage down:
The fibre brings the unmodulated laser reference to the spectrometer

## Slide 22: Spectrometer (current status)
Developed for classic spectrally decoded EO sampling
Single fibre input
Past resolution for ps pulses:
0.29 nm = 0.47 ps (closed slit)
0.35 nm = 0.57 ps (open slit)
IMAGE
INTENSIFIER
PMT
FIBRE
LAUNCHER
GRATING
SLIT
BEAM
SPLIT.

## Slide 23: Spectrometer (after the upgrade)
Being developed for the DEOS technique
Improved enclosure
4-to-1 input fibre bundle 1) Unmodulated pulse 2) Polarisation #1 3) Polarisation #2 4) Calibration lamp
Transplant in progress … (next week)
Resolution to be assessed and possibly improved after the realignment
IMAGE
INTENSIFIER
FIBRE BUNDLE
LAUNCHER
THORLABS
BFL44LS01
GRATING

## Slide 24: Thanks a lot for your help !!!
eugenio.senes@cern.ch

## Slide 25: Backup