## Slide 1: Fast Quality Assurance (QA) Tests for LGAD Production
p. Skomina, G. Kramberger, j. debevc, a. Howard, i. Mandić, b. Hiti
JSI Ljubljana

## Slide 2: Contents
Motivation
Devices & readout
Measurements & analysis
Results & extracted parameters
Conclusions

## Slide 3: Motivation
During production phase of sensors for HGTD extensive Irradiation Tests (IT) will be done on each batch (50-200 wafers per batch)
Typically several tests per batch planned
Radiation hardness of LGADs is very sensitive to small variation in production (1% diff. in [B] makes a difference)
IT are meant to provide info on radiation hardness
The outcome of the IT will largely depend on statistics. Methods should therefore be:
  simple & fast, to allow large number of samples measured with little resources
  such that allow extraction of large number of sensor parameters from a single measurement
  available at several sites
CC/timing measurements with 90Sr are slow & allow limited number of tests
Proposal: use a Scanning-TCT set-up and a 1x2 LGAD array w/ opening in IP region
Speed up the tests and provide reliable information on gain (and other important parameters)
Idea from studies of PiN and LGAD structures shown before

## Slide 4: Devices used
Instead of separate PiN and LGAD structures
1x2 LGAD (IPD measurement) w/ opening in passivation in ATLAS HGTD QC-TS
Gain only present on pads, no gain in IP region
Single measurement, multiple parameters extracted
Much faster check of irradiated samples than CC/timing
G=1
G>1
FBK4.0 1x2 sample
QC-TS 1x2 LGAD sample
G>1
x
y
z

## Slide 5: Readout of the LGADs
HV
Bias-T
GND
FBK4-W18
(deep C-GL)
1.5e15 neqcm-2
20°C, 190V
Sensor mounted in simple Al housing and wire-bonded (two bonds)
Single amplifier is used (explored also option with two amps)
Scan over the opening in the IP region using an IR laser (1064nm)

## Slide 6: Measurement – before irradiation
Depletion of “no gain” region – full depletion of active layer
Full depletion of the device
Depletion of gain layer – appearance of signal in GL (agrees with CV)
Gain increases
FBK4.0 W18, 20°C, Voltage scan

## Slide 7: Measurement – after irradiation
FBK4.0 W18, 1.5e15 neqcm-2
20°C, Voltage scan

## Slide 8: Analysis
Full depletion of “no gain” region – full depletion of active layer
Depletion of gain layer – appearance of signal in GL (agrees with CV)
Full depletion of the device
Gain increases
Extract parameters: CC in gap, CC on pad, Ileak, Igen, G, VGL, VFD, IPD

## Slide 9: Analysis – temperature scan
FBK4.0 W18, 1.5e15
Average of several points
Scaled to correct for different laser intensities & absorption of IR light in Si
No significant difference between 20°C and -30°C in important parameters (apart from of course Ileak)
20°C
-30°C
20°C
-30°C
20°C
-30°C
20°C
-30°C

## Slide 10: Extraction of parameters
One scan  several parameters measured
VGL, VFD extracted the same way as from CV
Gives good picture of radiation hardness of entire device
Analysis done with root script (can produce file ready for reporting to DB)
VGL ≈ 35V
VFD ≈ 75V
VFD ≈ 55V

## Slide 11: Conclusions
Use TCT and a 1x2 LGAD array w/ opening in IP region on the ATLAS QC-TS for QA-IT
At room temperature a full test takes around 2h – much faster than CC/timing
An analysis script will produce the output for the database with all extracted parameters: Ileak, Igen, G, VGL, VFD, IPD
Very low values of gain can be measured at room temperature (with CC in 90Sr gains at room temperature can not be observed due to noise)
Need to assure a clear correlation between gain measured for a sample with CC/timing using 90Sr at -30°C and ≈VSEB and the gain measured with TCT at room temperature
Once the correlation is established we monitor the differences with respect to the “reference” sample

## Slide 12: Thank you for your attention

## Slide 13: Backup Slides

## Slide 14: Experimental setup
Scanning TCT
XYZ tables
IR laser – focused beam of (8-10µm in waist)
HV and readout through Bias-T
Both pads connected to one amplifier
Operating conditions:
Temp. control from 40°C to -30°C
Insulation cap
Closed environment
Dry air (dew point well below measuring temperature)

## Slide 15: IP and gain measurement – 2 amps
A good agreement with measured charge in CC setup.
FBK4.0 – W9 Type9 (not irradiated) at 100V and 293K
G=1
G=9
HV

## Slide 16: FBK4.0 correlation of TCT with 90Sr
FBK4.0 – W18 deep GL design at 400 V and 243 K , Type9 IP
1.5e15 cm-2
G~7
G~1
G=7 ; Q~4 fC
Note UCSC CC measurement shows 4fC for that sample type with 90Sr!

## Slide 17: Addition to: Extraction of parameters
Ileak = Igen x G; G = G@VFD(gap)
Usually we don’t know the gain, but with this method we do and we measure the leakage current. So we can also extract the generation current.
At low gain values, the current gain is the same as charge gain (up to about G=4 for this fluence).