## Slide 1: Screen Testing for Beam Instrumentation at CERN
S.Burger – CERN / 20th of June 2023
Workshop: Low Density Materials for Beam Instrumentation Workshop
https://indico.cern.ch/event/1275649/

## Slide 2
Content
BTV Instrument Principle
Screen Light Emission Processes
Testing of Nanotube Material (not exhaustive)
  BNNT (CLEAR + HiRadMat)
  CNT (HiRadMat)
Conclusion

## Slide 3
 Beam Profile / Size
 Beam Position
Screen based instruments are at any trajectory change – injection & extraction lines but also in spectro / target / dump and transfer lines
 ~200 devices installed
Location of BTV instruments
CERN accelerator chain map
Beam Observation Principle: BTV Instrument

## Slide 4
Screen Types  2 main light emission processes
Scintillation screen:
    Photons of a characteristic spectrum are emitted following the absorption of radiation (== interacting with energized particles). Scintillation is an inherent molecular property in () molecules and arises from their electronic structures (Wikipedia). Isotropic emission.
    Material example:
      Ceramic Alumina, Al2O3:CrO2
      Powder P43, P47 (on substrate)
      & Plastic, Quartz, YAG, LAG, etc…
Optical Transition Radiation (OTR) screen:
    Is a form of electromagnetic radiation emitted when a charged particle passes through inhomogeneous media, such as a boundary between two different media (Wikipedia). This process emits photons in the visible and is proportional to the reflective index of the material. Specific angular emission.
    Material example:
      Ti, Al, Graphite, SiC, GlassyC, etc…
Efficiency Scintillation vs OTR ~ few 1e3
Low intensity  use of scintillation
High intensity  use of OTR

## Slide 5: Screen Thermal and Mechanical Damage Examples
Screen testing (2005) on 4.2MeV/u lead ions, 100μA, 600μs beam every 1.2s
Low Energy Ion Beam
High density
440GeV
Proton Beam
26Gev Proton Beam
~7e12p/cm2 in air
Chromox
Ti

## Slide 6: BNNT Material tested in CLEAR (1)
Reference: Aries Materials Workshop, June 22, 2021 Kevin Jordan, Boron-based nano-tubes as scintillators and further applications
e- in air
YAG Chromox BNNT
2022/02

## Slide 7: BNNT Material tested in CLEAR (1)
Light yield (normalized to Chromox):
Chromox 1 (IT 3ms)
YAG 5 (IT 3ms  >> 70 ns decay time)
BNNT 0.025 (IT 34us  >> few ns decay time)
Beam size:
Chromox and YAG have similar profiles
BNNT seems to have tails  in air, parasitic light (luminescence) of the same amplitude order as the signal !
1e7
2e6
5e4

## Slide 8
Screens as discovered after pumping down and back to atmospheric pressure
BNNT screen tests in HiRadMat (1)
May 2022
Screens as installed before pumping down
BNNT 1mg/cm2
BNNT 2mg/cm2
Glassy C
Chromox
In vacuum HiRadMat BTV
In vacuum system to mitigate parasitic light
Beam density range :
Minimum <1.2e11 p/mm2
Maximum 5.52e14 p/mm2

## Slide 9
Chromox 250um
OD 3
GlassyC
OD 1
BN-NT
OD 0
Viewer digital gain 22
Viewer digital gain 18
Single bunch _~9e10ppp_ Optics 0.5mm2
BNNT screen tests in HiRadMat (2)
May 2022
Ti
OD 0
Beam Size |  |  | 
0.46 x 0.61 | 0.47 x 0.60 | 0.47 x 0.63 | 0.5 x 0.63
Normalized Amplitude |  |  | 
200 | 1 | 0.05 | x
 Very good behavior

## Slide 10
New support to better stretch the material and ensure a flat surface.
BNNT screen tests in HiRadMat (3)
October 2022

## Slide 11
25b / NF
0.33x0.42
0.31x0.40
0.29x0.39
BNNT 60b / NF
216b / NF
2nd BNNT 216b / NF
288b / NF
Step + 4000
Saturated
BNNT 288b / OD1
Step -4000
Measurement seems OK but screen seems to be burnt by this beam impact !
 Confirms that damage is seen on the next extracted beam measurement
Image of the screen after the high intensity beam extraction test
Measurement OK
BNNT screen tests in HiRadMat (4)
October 2022 Optics is 0.25mm2

## Slide 12
 | 1 | 12 | 25 | 60 | 216 | 288
 | 4.80E+11 | 5.76E+12 | 1.20E+13 | 2.88E+13 | 1.04E+14 | 1.38E+14
 | 1.92E+12 | 2.30E+13 | 4.80E+13 | 1.15E+14 | 4.15E+14 | 5.53E+14
 | 1 | 12 | 25 | 60 | 216 | 288
 | 4.80E+11 | 5.76E+12 | 1.20E+13 | 2.88E+13 | 1.04E+14 | 1.38E+14
 | 1.92E+12 | 2.30E+13 | 4.80E+13 | 1.15E+14 | 4.15E+14 | 5.53E+14
ρ =
Beam density
BNNT gets damage
BNNT screen tests in HiRadMat (5)
October 2022

## Slide 13
Thanks to Anastasia Mikhalchan and Juan-Jose Vilatela
2 types of CNT: ‘standard’ & high purity
Very difficult to install the screen even on the new support made for stretching of the screen.
The screen can not be ‘drilled’ by hand. When screwed, the material is pulled inside its support
CNT screen tests in HiRadMat (1)
October 2022

## Slide 14
14:10:29
14:11:09
14:11:48
14:12:28
14:13:07
Screen not stretched
Amplitude is varying
Screen movements !?
2022_10_26
CNT Single bunch - 0.25mm
CNT Measurements (1)

## Slide 15
Screen not stretched
Amplitude is varying
Screen movements !?
Degradation visible ??
2022_10_26
CNT 60 bunches - 0.25mm
15:37:58
15:38:36
7500
4200
Screen not stretched
Amplitude is varying
Seems more stable ??
2022_10_26
CNT 25 bunches - 0.25mm
15:04:41
15:05:19
300
with OD1
3150
15:05:58
3150
CNT Measurements (2)

## Slide 16
2022_10_26
CNT 216 bunches - 0.25mm
16:22:07
16:22:46
15000
9000
17:09:17
14500
Screen not stretched
Amplitude is varying
Screen movements !?
Degradation visible ??
2022_10_26
CNT 288 bunches - 0.25mm
17:09:17
24500
Target damaged area
Target ‘virgin’ area
CNT Measurements (3)

## Slide 17
Looks like stretched : due to heating ??
Difficult to estimate the beam density which starts damaging the material and the one that destroys it…
Before
irradiation
After
irradiation
 Clear impacts seen on the material
Limit for CNT damage ??
CNT Measurements (3)

## Slide 18
New setup for 2023
CNT beam test, Next step
Laser cut of the hole for easier clamping of the screen
Laser cut for holes clearance
High purity and ‘standard’ CNT type installed ready for beam
Installation of the 2 types of screen ready for beam test

## Slide 19
New setup for 2023
Conclusion
BNNT and CNT material were tested on different beams (CLEAR and /or HiRadMat)
BNNT gives results in very good agreement with the well known Chromox or GlassyC screens
Due to its low yield, probably not adapted to low intensity beam.
Can not sustain full power of HiRadMat but looks ok up to a few 1e14p/mm2
 good performance for a scintillation screen type
CNT could not be tested correctly as the screen could not be installed flat or stretched enough for OTR measurement (angular dependent)
Nevertheless, light is observed (~same level as other OTR screens)
Can not sustain full power of HiRadMat but looks ok up to 1e14p/mm2
 Test to be redone soon with new setup and high purity material
Main advantages of NanoTube material as screens is the very little amount of material that sees the beam compared to standard screens which limits the radiation impact around the device
With 200 BTVs installed, we constantly look at new ideas or material to improve the performance of our instrumentation  all ideas, new material to be tested is more than welcome !!

## Slide 20: Thanks For Your Attention !

## Slide 21
CPS BTV
SPS BTV
TL LHC BTV
LHC BTV
ADT BTV
Most of the time in vacuum system

## Slide 22
60cm
14mm
Screen examples

## Slide 23
Beam Measurement Examples Using Screens

## Slide 24: Screen testing at HiRadMat in 2016
Screen materials installed for testing with a 440GeV/c p beam in air.
Testing yield and resolution, not specifically robustness…