## Slide 1: In the framework of PhD thesis “Optimization of ion beam parameters for very high dose rate ( FLASH ) radiotherapy”
Analytical pencil beam dose and dose-rate engine development for FLASH therapy studies of heavy ion and very high energy electron beams
Kristaps Palskis ( RTU IPPAT, CERN )
under the supervision of
Dr. Mariusz Sapinski ( PSI )
Prof. Joao Seco ( DKFZ )
Prof. Toms Torims ( RTU IPPAT )
Dr. Maurizio Vretenar ( CERN )

## Slide 2: Outline
[ 2 ]
Recap on FLASH therapy – clinical potential and dosimetric constraints
Motivation for current work – previous work and perspective of low LET particles
Considerations for input data set and methodology for Monte Carlo simulations
Dose model for ion beams:
  IDD normalization. Analytical model comparison for lateral profiles
  Beyond physical dose: analytical model for calculation of MKM input parameters and LET
Differences between VHEE and ion beams. Dose model for VHEE beams:
  Development of IDD normalization approach
  Analytical model development for lateral profiles and LET calculation
Considerations for the dose-rate engine
Discussion on achieved precision with an analytic model. Conclusions
Outlook on future work

## Slide 3: Introduction: What is FLASH therapy ?
Biologically observed differential effect – when irradiation with ionizing irradiation is performed at ultra-high dose rates (delivery timescale of miliseconds) adverse effects in healthy tissue are greatly reduced while DNA damage in cancerous tissue remains the same
[ 3 ]
[ 1 ] Vozenin M-C, De Fornel P, Petersson K, et al. The Advantage of FLASH Radiotherapy Confirmed in Mini-pig and Cat-cancer Patients. Clinical Cancer Research 2019
[ 2 ] Schüller A, Heinrich S, Fouillade C, et al. The European Joint Research Project UHDpulse – Metrology for advanced radiotherapy using particle beams with ultra-high pulse dose rates. Physica Medica 2020
[ 1 ]
[ 2 ]
No clear explanation of FLASH effect has been developed, while various proposals exist: transient oxygen depletion, radical recombination, metabolism of peroxidized compounds and others
Reduced number of treatment fractions (hypofractionation)
Reduced treatment side-effects due to radiation toxicity
Breathing motion interplay effects are mitigated
The FLASH effect
Mechanism of FLASH effect
Clinical advantages

## Slide 4: Introduction: Requirements to achieve FLASH effect
Conventional delivery (energy switching + PBS) – not feasible
Novel delivery approaches are considered – transmission beams, 3D-range modulators and magnetically focused beams
[ 4 ]
[ 3 ] Montay-Gruel P, Acharya MM, Gonçalves Jorge P, et al. Hypofractionated FLASH-RT as an Effective Treatment against Glioblastoma that Reduces Neurocognitive Side Effects in Mice. Clinical Cancer Research 2021
[ 4 ] Verhaegen F, Wanders R-G, Wolfs C, Eekers D. Considerations for shoot-through FLASH proton therapy. Phys. Med. Biol. 2021
[ 5 ] Jolly S, Owen H, Schippers M, Welsch C. Technical challenges for FLASH proton therapy. Physica Medica 2020
[ 3 ]
[ 4 ]
[ 5 ]
Threshold dose rate
Delivery must be done with at least 40 Gy/s
Threshold dose
A dose of at least 8 – 10 Gy must be delivered
Two dosimetric requirements
FLASH effect has been mostly observed with electron beams, with fewer studies – proton and X-ray beams. Data with heavier ions (helium, carbon) - sparse
Radiation type for FLASH effect
Ion FLASH delivery challenges

## Slide 5: Results so far: Model for FLASH dose threshold estimation
[ 5 ]
Due to the favourable physical characteristics of ion beams, it is of interest to study the feasibility of coupling ion beam therapy with FLASH effect – both from biological and technological perspective
Collaboration between CERN-NIMMS, DKFZ (BioMedical Physics in Radiation Oncology group) and RTU IPPAT
Study on ion therapy FLASH
Cellular DNA damage
Saturation of
solvated electron concentration
Saturation of
cellular DNA damage
FLASH effect
FLASH mechanism model by prof. Joao Seco
Experimental basis: Recent experiments show decreased oxygen consumption rate at ultra-high dose rates, reaching a saturation state
[ 7 ] [ 8 ]
Main idea – threshold concentration of solvated electrons as surrogate for dose threshold
Physical implementation of the model – part of my thesis
How does ion used affect this threshold?
FLASH dose threshold model
[ 6 ] Le Caër S. Water Radiolysis: Influence of Oxide Surfaces on H2 Production under Ionizing Radiation. Water 2011
[ 7 ] Jansen J, Beyreuther E, García-Calderón D, et al. Changes in Radical Levels as a Cause for the FLASH effect: Impact of beam structure parameters at ultra-high dose rates on oxygen depletion in water. Radiotherapy and Oncology 2022
[ 8 ] Jansen J, Knoll J, Beyreuther E, et al. Does FLASH deplete oxygen? Experimental evaluation for photons, protons, and carbon ions. Medical Physics 2021
[ 6 ]

## Slide 6: Results so far: Comparison of different ions
[ 6 ]
Model was used to study FLASH dose threshold for different ion types. Results have been presented:
Poster at FRPT’22
Assessing particle FLASH-RT effectiveness by modelling solvate electron production dependency on particle LET
Poster at PTCOG’23
FLASH-RT DOSE THRESHOLD EFFECT ESTIMATION FOR PARTICLE BEAMS BY MODELLING SOLVATE ELECTRON PRODUCTION DEPENDENCY ON PARTICLE LET 3D DISTRIBUTION
1H
3He
4He
6Li
7Li
9Be
10B
12C
14N
16O
20Ne
According to the developed model, lower physical dose threshold is required for FLASH effect with lighter ions
Solvated electron concentration distributions are more favourable (larger portions above threshold) for lighter ions
FLASH effect dose threshold depends not just on ion type, but irradiation geometry and FLASH «prescription point»
Conclusions so far . . .

## Slide 7: Results so far: Low LET particles and delivery perspective
[ 7 ]
Protons
Helium-4 ions
Carbon ions
Both in 1D and 3D study performed, model has shown lower, clinically more acceptable threshold doses for FLASH effect with lighter ions – low LET radiation
Another, novel low LET modality - VHEE
Current study has explored only general considerations – only single field irradiation without a specific delivery method
Multi-field irradiation? Spot scanning ?
Dose rate distribution?
To adequately compare ion and VHEE irradiation – multi-field treatments must be studied due to dose distributions of VHEE
General TPS is necessary for further studies
Implications of the current results

## Slide 8: Motivation and aim of this sub-study
[ 8 ]
Motivation for the sub-study
For further studies of ion beam FLASH and comparison with VHEE modality:
a general, non-Monte Carlo TPS should be developed – pencil beam convolution method
the developed TPS should include LET (solvated electron concentration) and biological dose
modelling capabilities
applicable for dosimetric parameter calculation in multi-field irradiation scenario
account for delivery method differences
dose-rate distribution calculation possibility
Goal of the sub-study
To develop an analytical dose calculation model and dose-rate engine applicable
for a pencil beam convolution based treatment planning system (TPS) for
heavy ion and VHEE beams

## Slide 9: Monte Carlo simulation set-up – input for analytical fit
[ 9 ]
Geant4 environment, RTU HPC computing resources
Choice of physics list settings:
  Electromagnetic – emstandard_opt4
  Hadronic – QGSP_BIC_HP
  Additional – G4RadioactiveDecayPhysics, …
0.5 mm range cuts and 1 mm step limit
Cylindrical geometry ( 20 cm, 40 cm)
Material: G4_WATER (I=75 eV)
106 particles simulated per energy
Energy deposition as a 2D histogram with 0.1 mm resolution (depth and radial distance)
Simulation set-up
According to previous study - interest in light ions and VHEE
For ions:
1H, 3He, 4He, 6Li, 7Li and 12C (for reference)
Which ions ?
Pristine pencil-beam is considered
No specific beamline set-up is assumed
No initial energy spread – can be implemented by interpolation of base data set
No initial beam size – can (and practically is) implemented in convolution with fluence map
No initial beam divergence - can be implemented in the convolution process + ray-tracing
Beam set-up
[ 9 ] Arce P, Bolst D, Bordage MC, et al. Report on G4-Med, a Geant4 benchmarking system for medical physics applications developed by the Geant4 Medical Simulation Benchmarking Group. Med Phys. 2021

## Slide 10: Monte Carlo simulation set-up – input for analytical fit
[ 10 ]
Calculate energy-range relationship by solving Bethe-Bloch equation
As suggested in [ 10 ] – for practical treatment planning energies can be chosen in discrete steps
For conventional and 3D range modulator delivery:
Energies corresponding to
RH2O = 20 to 360 mm (step=20)
For «shoot-through» delivery:
Two transmission energies with
RH2O = 500 and 600 mm
How many initial energies ?
[ 10 ] Krämer M, Jäkel O, Haberer T, Kraft G, Schardt D, Weber U. Treatment planning for heavy-ion radiotherapy: physical beam model and dose optimization. Phys. Med. Biol. 2000

## Slide 11: Biologically relevant parameter simulation
[ 11 ]
LET – important for biological effect assesment and solvated electron concentration calculation
Geant4 lookup function – ComputeElectronicDEDX
Score for all charged particles (electrons included)
Output – dose weighed LET
Linear Energy transfer
MKM input parametrs are neccessary for biological dose calculation
For biological dose calculation – Microdosimetric Kinetic Model (MKM) used – HSG cell line assumed
Saturation-corrected dose mean specific energy - z*sat
Data available only for Z<6 – calculate z*sat kinetic energy dependency - Kiefer–Chatterjee ion track parametrization [ 11 ]
Calculated data – lookup function created
Score for all ions (electrons excluded)
MKM input parameters
[ 11 ] Kase Y, Kanai T, Matsufuji N, Furusawa Y, Elsässer T, Scholz M. Biophysical calculation of cell survival probabilities using amorphous track structure models for heavy-ion irradiation. Phys. Med. Biol. 2007
For separation of projectile and target fragmentation reaction products :
Score with Z ≤ Zprojectile – projectile fragmentation + secondaries of same type
Score with Z ≤ Zprojectile + 1– light target fragments
No Z differentiation – all target fragments

## Slide 12: Ion beams: IDD normalization
[ 12 ]
Basis of pencil beam convolution algorithm
Integrated depth dose
Lateral profile
IDD(z) is normalized with respect to depth – IDD(z)  IDD(z/Rmax) - stored as lookup table.
IDD for intermediate energies (ranges) via linear interpolation. Good agreement shown – difference below 1-2 %
IDD normalization for lookup function
Protons
12C ions

## Slide 13: Ion beams: Choice of fitting functions for lateral profiles
[ 13 ]
3 combinations were considered:
Double Gaussian + Logistic, Double Logistic + Gaussian and Triple Gaussian
Large part of TPSs – single Gaussian approximation – insufficient description of tails due to large-angle single scattering events
Double or triple Gaussian – better description
Following [ 12 ] – logistic functions can also be considered - better description of tails
Lateral dose profile functions
[ 12 ] Zhang H, Dai Z, Liu X, et al. A novel pencil beam model for carbon-ion dose calculation derived from Monte Carlo simulations. Physica Medica 2018
Lateral dose profile analytical description

## Slide 14: Ion beams: Workflow of the fitting routine
[ 14 ]
Due to large number of free parameters – non-unique solutions can appear in the curve fitting process
Dedicated fitting routine was adapted
Single Gaussian fit – non-specific starting point
Triple function fit
Specific starting point ( approach similar to [ 12 ] [ 13 ] ):
f1 – σ1 = 0.6 * σstart and w1 = 0.8
f2 – σ2 = 2 * σstart and w2 = 0.12
f3 – σ3 = 3 * σstart and w3= 0.08
Extract σstart
[ 13 ] Hirayama S, Takayanagi T, Fujii Y, et al. Evaluation of the influence of double and triple Gaussian proton kernel models on accuracy of dose calculations for spot scanning technique. Medical Physics 2016
Parametrization of primary contribution
Extract min(σ1 , σ2 , σ3) - parametrize by Highland-like function
Re-optimization with fixed σ1 corresponding to the primary particle contribution
(ROI – 0.01% of CAX dose)
Lateral spread due to MCS of primary particles can be parametrized analytically with Moliere or Highland formula
For each ion type a function was used to parametrize Gaussian spread parameter (σ) dependendence with energy and depth
 | A | B | C | D
1H | 0.192 | 1.029 | 0.191 | 0.746
3He | 0.130 | 0.996 | 0.156 | 0.637
4He | 0.097 | 1.002 | 0.184 | 0.728
6Li | 0.088 | 0.978 | 0.192 | 0.667
7Li | 0.083 | 0.970 | 0.176 | 0.703
12C | 0.027 | 1.053 | 0.307 | 1.196
Highland-like function

## Slide 15: Ion beams: Comparison of fit accuracy achieved
[ 15 ]
Example comparison of fitting functions
R = 30 cm
@ 250 mm depth
RMS residual error comparison
R = 30 cm

## Slide 16: Ion beams: Choice of analytical function for lateral profile
[ 16 ]
Generally, better agreement with data shown when using a combination of double Gaussian and a logistic function for analytic description of lateral dose profile
Lookup tables created for TPS – depth dependence of the 3 spread parameters and relative weights

## Slide 17: Ion beams: LET and MKM parameter modelling approach
[ 17 ]
[ 14 ] Hirayama S, Matsuura T, Ueda H, et al. An analytical dose‐averaged LET calculation algorithm considering the off‐axis LET enhancement by secondary protons for spot‐scanning proton therapy. Medical Physics 2018
Both MKM input parameters and LET have off-axis dependence
Approach developed for proton beams [ 14 ] has been adapted and expanded for ions
Calculated lateral dose profile functions used as input for LET and specific energy distribution model
Comparison was made – «halo» modelled by two independent functions or a sum of them
Lateral dose profile functions
Multiplied by dose deposited – higher weight @CAX
Multiplied by step function – ignore low dose envelope
Considerations for cost function
4He 242 MeV/u
CAX
4He 242 MeV/u
Profile @ 320 mm

## Slide 18: Ion beams: LET and MKM parameter analytical model
[ 18 ]
Lookup tables created for TPS – depth dependence of the 3 LET and 3 z*sat values
Z ≤ Zprojectile scoring for final data set – heavy target fragments introduce noise
4He 150.7 MeV/u
6Li 190.6 MeV/u

## Slide 19: Dosimetrical comparison of VHEE and ion beams
[ 19 ]
Very High Energy Electrons – electron beams with E > 50 MeV
From CERN perspective – 250 MeV (CLEAR facility)
Less prone to patient set-up and motion uncertainties compared to photons and heavy ions
Technologically less demanding for FLASH dose rates
From perspective of our model – low LET particle
What is VHEE ? Advantages?
Physics lists settings – same as for ion beams
MKM parameters not scored – RBE = 1
LET scoring – differentiate separetely heavy photonuclear reaction products
Input energies:
  50 to 100 MeV in steps of 10 MeV
  100 to 300 MeV in steps of 25 MeV
Considerations in MC setup

## Slide 20: VHEE beams: Developing IDD normalization approach
[ 20 ]
VHEE – no prominent dose maximum peak as with ions
From experience with clinical electron (< 20 MeV) beams:
  dose maximum depth – physically not as significant as for ions
  Parameters of importance – R90, R80, R50 and Rp
  All three parameters extracted and compared for IDD normalization, best agreement and ease of calculation with using practical range Rp
IDD(z)  IDD(z/Rp)
Practical range

## Slide 21: VHEE beams: Choice of fitting functions for lateral profiles
[ 21 ]
No pencil beam convolution based TPS exists for VHEE
( accuracy of this aspect to be discussed )
VHEE exhibits pronounced field size dependence due to lateral scattering equilibrium – lateral dose profile modelling accuracy of high importance
Different combinations of fitting functions considered:
  Double Gaussian
  Double Logistic
  Double Gaussian + Logistic
  Double Logistic + Gaussian
  Triple Logistic
Lateral dose profile functions
A combination of double Gaussian and a logistic function chosen for analytic description of lateral dose profile

## Slide 22: VHEE beams: Empirical function for parametrization
[ 22 ]
Current parametrization is of empirical nature – further investigations are needed, possibly by separating energy deposition events by primary, secondary, tertiary (and higher order) electrons and electrons created by interactions of bremsstrahlung photons
With current data set (no differentiation of energy deposition events by physical process) parametrization with Highland-like function does not ensure sufficient fit quality
At this stage, an empirical function was used for parametrization of third component – contribution by logistic function
with free parameter values calculated as:
A = 2.233
B = 0.362
C = 0.921
D = 2.366
E = 1.361
Empirical function
for parametrization
300 MeV
@150mm
300 MeV
@350mm

## Slide 23: VHEE beams: Choice of fitting functions for lateral profiles
[ 23 ]
Lookup tables created for TPS – depth dependence of the 3 spread parameters and relative weights

## Slide 24: VHEE beams: LET analytical model
[ 24 ]
Same approach as with ions
250 MeV
CAX
250 MeV
Profile @150mm
250 MeV
Scoring without heavy fragments from photonuclear reactions

## Slide 25: Dose-rate engine: How to define and implement?
[ 25 ]
[ 15 ] Deffet S, Hamaide V, Sterpin E. Definition of dose rate for FLASH pencil‐beam scanning proton therapy: A comparative study. Medical Physics 2023
As mentioned – dose rate is crucial parameter to deliver FLASH therapy and assess it’s effectiveness
Different definitions exist calculation [ 15 ]:
  Pencil beam scanning dose rate
  Percentile dose rate
  Maximum percentile dose rate
  Average dose rate
  Dose-averaged dose rate
No clear identification, which formalism is most significant from biology perspective
Although – average dose rate / dose-averaged dose rate could be more significant estimator
How to plan dose-rate ?
In pencil beam convolution – separate the spot delivery pattern in temporal domain.
The developed dose engine is a starting point !
How to implement ?
[ 15 ]
[ 15 ]

## Slide 26: Discussion. Conclusions
[ 26 ]
An analytic dose engine has been developed for a pencil beam convolution based calculation for proton, heavier ion (3He, 4He, 6Li, 7Li, 12C) and VHEE beams for application in FLASH therapy
Dose engine is capable of LET (and solvated electron yield) and biological dose calculation
Accuracy of the developed dose engine (compared to Monte Carlo) :
  Integrated depth dose: < 1% in entrance region, < 2-3 % in Bragg peak region
  Lateral dose profile: < 1% for regions with dose higher than 0.01% of CAX
  LET and biological modelling parameters: < 1% for regions with dose higher than ~ 1 % of CAX
A novel approach was used for lateral distribution definition - sum of 2 Gaussian and a logistic distribution.
For studies of clinical cases (inhomogeneities due to different human tissue) pencil beam convolution algorithm limitation should be explored and taken into account, especially for VHEE beams
As base data set is Monte Carlo data – MC accuracy limitations itself (especially LET and biological dose)

## Slide 27: Future outlook
[ 27 ]
Fully implement the developed dose engine for dose rate calculation (various definitions)
Integrate the dose engine into a TPS:
  use existing software as open-source academic TPS - matRad
  or couple with any dose-influence matrix calculation algorithm
Perform multi-field irradiation (clinical scenario) studies for heavy ion and VHEE based FLASH therapy
  optimum spot size and spacing parameters
  possibly – spot delivery sequence
  usage with different delivery techniques – transmission beams and/or 3D range modulators
Possibly – explore the usage of the developed dose engine for analytical calculations for quadrupole focused beams

## Slide 28
Thank you for your attention ! Open for questions !