## Slide 1
Operation and Performance of the CMS Muon System for LHC Run 3
G. Pugliese
INFN & Politecnico of Bari
On behalf of the CMS Collaboration
ICHEP, Prague 17-24 July 2024

## Slide 2: The CMS detector
Tracker
Muon Spectrometer
Electromagnetic calorimeter
Hadron calorimeter
Weight: 14000 t
Length: 28.7 m
Diameter: 15 m
Magnetic field: 3.8 T
Superconducting Solenoid
Niobium titanium coil carrying 18.000 A
Steel Return Yoke
Barrel Region
Endcap
Endcap

## Slide 3: The CMS Muon Spectrometer
Muon system: three gaseous technologies for muon identification, timing and momentum measurement
Muon acceptance: |η| < 2.4
Cathode Strip Chambers (CSC)
540 trapezoidal chambers, 500k channels
24 number of hits
Spatial resolution≈50 ÷140 mm
Time resolution ≈ 3 ns
Resistive Plate Chambers (RPC)
540 trapezoidal endcap chambers
480 rectangular barrel chambers
k channels
6 (4) number of hits
Spatial resolution ≈1 cm
Time resolution ≈ 1.5 ns
Gas Electron Multiplier (GE1.1 installed in 2021):
72 Super-Chambers, consisting of two triple-GEM
2 number of hits
Spatial resolution≈100 mm
Time resolution ≈ 10 ns
Drift Tubes (DT)
250 chambers, 170k channels
44 number of hits
Spatial resolution≈100 mm
Time resolution ≈ 2 ns
four
See more results in A. Sharma and S. Buontempo’s talks

## Slide 4: The CMS and LHC schedule
1997 Muon Project TDR
Phase 1
Phase 2
2010 Start of LHC
2013 The EU Strategy Report for High Energy Physics approved the HL-LHC as priority project
2017 The Phase-2 Upgrade of the CMS Muon Detector TDR
R&D and production
2012 Higgs Discovery
Run 1 & Run 2
2005-2008 Muons Chambers produced and installed

## Slide 5: LHC and HL-LHC schedule
Long Shutdown 2 (2019-2021):
Extensive work to upgrade the CMS detector because of the increase of energy and luminosity in Run 3 and beyond
New CSC electronics and GEM chambers in the Muon Spectrometer
https://home.cern/press/2022/CMS-upgrades-LS2
Phase 1
Phase 2
Notes: Maintenance and repair activities
Hardware interventions to improve detector performance at HL-LHC (anticipating some work needed after LS3)
Installation of Muon and BRIL demonstrators
CMS detector re-commissioned with comic muons, pilot beam, splashes and collisions

## Slide 6: LHC and HL-LHC schedule
Run3
Phase 1
Phase 2
1299 papers published
5th July 2022: start of Run 3 & first stable beam energy record of 13.6 TeV
Notes: Maintenance and repair activities
Hardware interventions to improve detector performance at HL-LHC (anticipating some work needed after LS3)
Installation of Muon and BRIL demonstrators
CMS detector re-commissioned with comic muons, pilot beam, splashes and collisions

## Slide 7: Muon Operation in Run3
Excellent CMS data-taking efficiency >92% in RUN3 physics collisions
Muon System:
  Smooth operation with a minor contribution to Luminosity loss (4%, due to CSC readout)
  Stable active channels fraction:
Key to success: prompt intervention of the experts during all beam-off and Technical Access time in case of failure
CSC
~99%
DT
~98.6%
Year | 2018 | 2022 | 2023 | 2024
% of RPC active channel | 96.5 | 89.6 | 87.7 | 82.6
Since 2022, all RPC leaky chambers (located only in the Barrel region) have been disconnected to reduce GHG emissions and to use the new RPC recuperation system efficiently
GEM: ~92%
GEM dead channels are mainly due an electronics issue (VTRx outgassing is causing damage in the optical connection) that will be fixed in LS3
Notes: •DT : 98.57 %
•CSC : ~ 99 %
•RPC : ~ 83%
- Barrel : ~ 69%
- Endcap : ~ 99%
•GEM : ~92%
For 2024
Lumi loss
CSC 3.9 % rest negligible
For Run3
CSC 1.7%
GEM 1.0%
DT and RPC negligible
The VTRx photodiode outgassing is causing damage in the optical connection so these channels are dead and we will not be able to fix until LS3
There is a nice explanation of the VTRX problem/solution in the May 2022 EP seminar by Brendan Regnery: https://indico.cern.ch/event/1108957/

## Slide 8: DT Segment Efficiencies in Run 3
DT segment efficiency: stable in time and more than 99% with few exceptions due to known hardware problems (one chamber in Wheel -2 Sector 8 MB1 is not working because of a gas connection problem and will be fixed over the coming YETS)
Segment efficiency is studied with Tag &Probe method on di-muons

## Slide 9: DT Segment Efficiencies in Run 3
See more results in C. Battilana’s talk
Excellent DT segment performance up to the maximum LHC instantaneous luminosity (~2.1*1033 cm-2s-1)

## Slide 10: CSC Segment Efficiency in 2024
CSC Segment efficiency: more than 98% of the CSCs is operating at close to 100% efficiency. Few chambers with lower efficiency are due to known reasons (electronics board failures, that cannot be fixed without access to the chambers, or accessional temporary failures

## Slide 11: RPC performance in Run 3 (Barrel region)
Barrel RPC efficiency and cluster size distributions in Run2 (2018) and Run3 (2022, 2023, and 2024)
RPC performance is measured using the Segment Extrapolation Method where DT/CSC Segments (in the Barrel/Endcap) that belong to a standalone muon track with timing corresponding to RPC readout BX windows are selected and extrapolated to the plane of a given RPC
Stable Barrel RPC performance in RUN 3 and in agreement with previous LHC Runs

## Slide 12: RPC performance in Run 3 (Endcap region)
Endcap RPC efficiency and cluster size distribution in Run2 (2018) and Run3 (2022, 2023, and 2024)
Stable Endcap RPC performance in RUN 3 and in agreement with previous runs

## Slide 13: Local trigger and TwinMux performance
In the Barrel region, DT and RPC information is processed in two stages to provide optimal online reconstruction inputs to Level-1 Muon Trigger Track Finders.
In the first layer, TwinMux boards match RPC hit clusters with DT Local Trigger segments in order to:
  recover DT inefficiencies using RPC-only primitives (MB1 and MB2 stations only)
  improve BX identification efficiency by exploiting the RPC's excellent time resolution
DT local trigger and TwinMux efficiency are both stable as a function of LHC instantaneous luminosity
TwinMux efficiency is bigger than the DT local trigger efficiency of:
  ~3% in the MB1 and MB2
  ~1.5% in the MB3 and MB4
Notes: …. making the best out of DT spatial resolution and RPC time resolution
RPC timing used to assign the bunch-crossing of DT segments without 8 fired layers

## Slide 14: GEM calibration in 2024
Calibration runs were taken in 2024 at different HV settings and Frontend chip (VFAT) configurations:
○ Pre-amplifier [low, medium, high] gain
○ Comparator mode [ARM, CFD].
GEM analysis done using events with standalone (STA) muons (pT > 10 GeV, with at least 15 hits in the muon system, and χ2 < 5) and with hits in the CSC companion station (i.e. accept a track through GE1/1 only if it contains ME1/1 hits)
GE1/1 efficiency vs. current
GE1/1 efficiency map

## Slide 15: GEM performance in RUN3
More on GE1.1 operations in S. Calzaferri’s talk
A new HV setting was applied in the middle of June 2024
A significant increase in the GE1.1 efficiency has been measured with HV-optimized and High gain Constant Fraction Discriminator
Fully validated the alignment for trigger capability: the Banding Angle distributions after the alignments are close to the ideal case

## Slide 16: Background in 2024
A New Forward Shielding (NFS) was designed to reduce background in the cavern detected by the muon detectors. In the last shutdown, it was installed only on the negative side
Significant reduction of the currents and hit rates observed mostly in the outermost layer of the Muon System (MB4 and ME4)
DT Background currents and rates versus instantaneous luminosity in one MB4 sector of W-2 as measured in 2024 and 2023
56 % reduction
52 % reduction
Notes: which will receive from now
on a much lower dose than before.
240 tons of additional shielding material:
  90% Concrete
  10% Steel

## Slide 17: Background in 2024
The effect gradually decreases going negative side (where the shielding is installed) to the opposite side
No effect in the internal stations (MB1, MB2 and MB3)
The effect will be fully symmetric after the installation of the NFS on the positive side, thus dramatically reducing the detector aging process
NFS position
Notes: which will receive from now
on a much lower dose than before.

## Slide 18: Conclusions
The CMS Muon system is operating extremely well, delivering good triggers and data for physics:
Negligible % contribution to CMS luminosity loss
Stable fraction of active channels
After 14 years of LHC running with increasing instantaneous luminosity and 20 years since construction, the Muon detector performance remains within specifications both as a triggering and as a reconstruction system:
Muon chamber performance is stable with no degradation observed
Excellent and stable DT and CSC Segment and local trigger efficiency
Excellent efficiency for the new GE1.1 station: 94%
Significant reduction of the LHC background measured after the installation of the new shielding in early 2024
NEWS: published the paper Development of the CMS detector for the CERN LHC Run 3

## Slide 19: Thanks!
Credits
to CMS People

## Slide 20: Backup slides

## Slide 21: RUN3 DT Hits Efficiencies
DT hits efficiency exceeds 96.9 % (few exceptions are due to known hardware problems) and it is stable
In 2024, one chamber only in Wheel -2 Sector 8 MB1 is not working because of a gas connection problem: this problem is reported as a light blue bin in the plot and will be fixed over the coming YETS.

## Slide 22: CMS Upgrade Project
The CMS detector has to be upgraded to cope with expected HL-LHC conditions (highest rate, fluence and pileup ever achieved) for new measurements and new physics searches

## Slide 23: LHC and HL-LHC schedule
Today
10 years of running at higher rates and radiation doses
Phase 1
Phase 2
Notes: So far LHC has delivered 5% of the total planned integrated luminosity. By the end of RUN3, it will be double
Major HL – LHC challenges wrt. LHC:
The instantaneous luminosity will increase of a factor 3-4
The Pile-Up (PU) (interaction proton – proton collisions in each Bunch Crossing) up to 5 times
The integrated luminosity up to 10 times
higher rates and radiation doses (radiation damage of detector and electronics components)

## Slide 24: Muon Reconstruction
Local reconstruction
Performed within single chamber
Tracker track reconstruction
Standalone Muon reconstruction
Performed using DT/CSC segments & RPC hits
Global muon reconstruction (out side –in): a standalone muon is propagated to match a tracker track. If matching is positive a global fitting is performed.
Tracker Muon (inside – outside): a tracker track is propagated to muon system and qualified as muon if matching with standalone or one segment.

## Slide 25: CSC Electronics Upgrade motivation
On-chamber and off-chamber electronics to be replaced in order to handle the CMS trigger requirements at HL-HC

## Slide 26: LS2 CSC Upgrade activity
The on-detector Refurbishment of Electronics in LS2

## Slide 27: Muon Upgrade
New GEM and RPC detectors needed to improve efficiency reconstruction and trigger performance at HL_LHC
NEW