## Slide 1: SM18 validation of RFD CM Prototype for SPS tests
Katarzyna Turaj on behalf of WP4 and SY-RF-SRF
13th HL-LHC Collaboration Meeting, Vancouver, 25-28.09.2023

## Slide 2: RFD CM parameters
Voltage 3.4 MV/cavity
Epeak 40 MV/m
Bpeak 70 mT
Frequency 400.79 MHz
Q0 10^10
Qext 5 x 10^5
Cavity tuning ±100 kHz
Temperature 2.0 K
RF power 40 kW
Eng spec. 2043014
Courtesy: T. Capelli

## Slide 3: Goals of RFD CM cold test in SM18
Testing of RFD CM is a crucial validation step of the cryomodule assembly process.
Validation of the cryomodule at 2K is required before RFD CM installation in the SPS.
Courtesy: A. Kosmicki

## Slide 4: Infrastructure: M7 test stand
Dedicated bunker M7 to allow full validation of CC CM.
The test infrastructure will be compatible with both RFD and DQW cryomodules.
  This test stand will be used to qualify all HL-LHC CM (11).
Manufacturing and Inspection Plan : CRAB RFD CM TESTING SM18  EDMS 2756481
Courtesy: A. Kosmicki

## Slide 5: SM18 planning changes
Due to delays in production and a short window for installation in the SPS during EYETS 2023-24 the testing program in SM18 was significantly reduced (*).
  Test flow and steps were adjusted.
  Duration reduced from original 4 months to 1.5 month.
The main goal is to perform Vacuum/Cryo-System leak tests at 2K.
All measurements and tests at 2K were reduced to minimum.
  No conditioning of the RF power couplers.
  No single cavities performance tests.
  RF tests only include measurements of the fundamental frequency, HOM spectrum and tuner response.
Testing in SM18
Removal of DQW
Installation of RFD CM in SPS
No works during the Christmas break
(*) 12th HL-LHC Collaboration Meeting, : RFD-CM reception at CERN & acceptance tests

## Slide 6: Workflow of RFD CM in SM18
CM Reception
Transfer to M7
Alignment
RF Warm Measurement
Pump & purge
Cool Down
Low Power Measurement
Connection
(VAC/CRG/RF)
RF Measurement
LLRF Test
Thermal cycle and final
Warm-up
Heat Loads
RF Conditioning
RGA
Ready to Transport
Disconnection
Leak Test
RGA
cancel

## Slide 7: SM18 planning
Cool-down
2K
Warm-up
Delivery and inspections
Leak checks
Jumper and safety group extension welding
Alignment
Pressure test
Final checks
Leak checks
Disconnection

## Slide 8: Conclusions
The test infrastructure is ready to enable validation of the RFD CM.
A validation and testing program is being developed. Includes procedures, reports and full data logging.
Due to delays in production, the test program was reduced.
The main goal of the SM18 test is to verify Vacuum/Cryo-System leak tightness at 2K.
Only low power RF measurements are intended.
Courtesy: T. Capelli
Notes: LLRF installation on going RF Power: Reserved space ready for installation Cryogenics: Infrastructure in place.

## Slide 9: Thank you very much!

## Slide 10: RFD CM Test Program
Manufacturing and Inspection Plan: CRAB RFD CM TESTING SM18  EDMS 2756481
Tests (Eng. spec. EDMS2043014) | Comments
RF Frequencies of deflecting mode | RF VNA measurements (after string assembly) Target frequencies, 400 MHz: , HOMs:
Q0-Vt curve including power dissipated at 4.1MV at 2K | Calorimetric measurements
Quench field value | 
Lorentz Force Detuning test | 
dF/dp test | 
HOM frequencies & Q’s | With particular care to frequencies around 760 MHz for RFD and 960 MHz for DQW
Output power through HOM coupler | 
Test on modes around 760 MHz | 
External coupling verification of ancillaries | External quality factor of several HOMs External quality factor of the field antenna (at fundamental mode frequency)
Field emission onset (Emitted radiation from cavity) | < 50 µSv/hr2
Multipacting levels | Sustained RF full power within vacuum limit of 1 10-8 mbar
Effect of thermal cycling (15K-2K) | 3 cycles. The variation of all the parameters in this table shall be within a range of +/- 5%
Tests (Eng. spec. EDMS2043014) | Comments
RF Frequencies of deflecting mode | 
Q0-Vt curve including power dissipated at 4.1MV at 2K | Calorimetric measurements
Quench field value | 
Lorentz Force Detuning test | 
dF/dp test | 
HOM frequencies & Q’s | With particular care to frequencies around 760 MHz for RFD and 960 MHz for DQW
Output power through HOM coupler | 
Test on modes around 760 MHz | 
External coupling verification of ancillaries | External quality factor of several HOMs External quality factor of the field antenna (at fundamental mode frequency)
Field emission onset (Emitted radiation from cavity) | < 50 µSv/hr2
Multipacting levels | Sustained RF full power within vacuum limit of 1 10-8 mbar
Effect of thermal cycling (15K-2K) | 3 cycles. The variation of all the parameters in this table shall be within a range of +/- 5%
Notes: RF frequencies of deflecting mode
Q0-Vt curve including power dissipated at 4.1 MV for 2 K operation and quench field
Lorentz Force Detuning
dF/dp test
HOM frequencies & Q’s values
Output power through HOM couplers
External coupling verification of ancillaries
Field emission onset ( <50uSv/h2)
Multipacting levels
Effect of thermal cycling (up to 80K)

## Slide 11: RFD cryomodule ports/interfaces
FSI – DN63 CF flange
Total Qty : 16
Pick up antenna feedthrough
Total Qty : 2
Insulation vacuum instrumentation – Iso K DN100
Total Qty : 1
Insulation vacuum pressure relief – Iso K DN160
Total Qty : 1
Instrumentation port
3 Flanges ISO K DN100
Beam vacuum instrumentation 1/2
Total Qty : 4
Drawing under preparation
Beam vacuum instrumentation 2/2
Total Qty : 4
See drawings XXXXX
HOMS feedthrough
Total Qty : 4
Insulation vacuum pumping port 1/2
Total Qty : 4– iso K DN100
Insulation vacuum pumping port 2/2
Total Qty : 4– iso K DN100
FSI – DN63 CF flange
Total Qty : 16
Level gauges feedthrough
2x / 2x ConFlat (CF) DN40
Cryogenic safety port for pressure measurement
1 Flanges ISO K DN63
Cryogenic jumper
Iso K DN400
Total Qty : 1
Instrumentation port
1 Flange ISO K DN100
RF connexions
SY/RF
Alignment and position monitoring
BE/GM
Insulation vacuum
TE/VSC
Cryogenic
TE/CRG
Beam vacuum
TE/VSC
2750
1040
875