## Slide 1: Actions for the SPS BWS upto and including YETS 23-4
Ray Veness

## Slide 2: Actions before YETS
Machine intervention
  Two spare scanner instruments were calibrated and transferred to BA4 radioactive buffer in case an urgent intervention is needed
  As the wires in BA4 are intact and do not show any signs of degradation, we do not need to make a machine intervention before YETS
Machine operations upto YETS23-4
  Can we/do we want to push the limits further before the ion run/YETS?
Ferrite availability
  F.Carra (Thanks Fed!) has 35 tiles of TT2-111R, with same dimensions as those installed (75x25x6)
    Have been fired, but should consider re-treating
  Other ferrites from, for example, the PS-WCM or from SY-ABT may be available
  Do we need to try and urgently procure other materials?

## Slide 3: Decisions to take before YETS23-4
BA4 BWS
  Do we return the scanners to their original orientations (2xV)?
  Do we add a coupler to the second tank and if so, do we connect it?
  Do we modify the number of ferrites in one or other tank (either both 5 ferrites, or both 6)?
  If the wires are not broken, do we still replace the wires/instruments?
BA5 BWS
  Do we keep both scanners as 2xH (see above)
  What ferrite configuration do we use in BA5?
  Do we add couplers to the BA5 tanks and if so, do we (try to) connect them?

## Slide 4: Actions during YETS23-4
Week 47/2023 (BA5)
  Removal of the BA5 BWS for wire replacement and calibration
  Recover broken wires/instruments for inspection
  Modification of the 2xBA5 tanks
Week 2/2024 (BA5)
  Backup option for tank modification
  Installation the two scanners
Week 3/2024 (BA4)
  Modification of the 2xBA4 tanks
  Installation of the 2 scanners (either re-installation of same scanners, or the two spares currently in BA4 storage)

## Slide 5: Actions for post-YETS23/4
Do we understand the ‘limits’ for the BWS performance in the SPS, based on operational data and simulations to-date?
Are further mitigations possible and/or necessary?
  Optimisation of ferrite mitigation?
  Optimisation of coupler performance?
  Instrument design changes (fork material, geometry change,…)