## Slide 1
Abstract: Novel acceleration schemes pose new challenges to the beam instrumentation required. This contribution presents a novel device to measure the beam position, enabling the discrimination of different co-propagating beams. The method leverages the characteristic properties of the Coherent Cherenkov-Diffraction Radiation (ChDR) emitted from dielectric inserts in the beampipe. The beam discrimination is performed in the frequency domain, exploiting the bunch length difference of the two beams. This device was developed for the AWAKE experiment, where not only an electron beam co-propagates with a more intense proton beam, but also traditional pickups are impacted by the environment polluted with spurious charges from the plasma. The electron beam discrimination takes place in a narrow band around 30 GHz. The overall design and results from the AWAKE experiment are presented. The utilisation of coherent ChDR to distinguish different co-propagating beams is a substantial novelty in the field, pushing the instruments capabilities for novel accelerating technologies, such as plasma-based accelerators.
Title: A dielectric pickup for short bunches

## Slide 2: A dielectric pickup for short bunches
Eugenio Senes, V. Bencini, C. Davut, W. Farabolini, P. Karataev, P. Korysko, M. Krupa, K. Lasocha, S. Liu, T. Lefevre, T. Manson, S. Mazzoni, B. Moser, C. Pakuza, E. Poimenidou, A. Schloegelhofer, A. Topaloudis, B. Spear, L. Verra, V. Verzilov, M. Wendt, G. Zevi Della Porta
International Beam Instrumentation Conference IBIC 2023, Saskatoon, Canada
11 September 2023

## Slide 4: Outline
Motivation
The AWAKE experiment
Co-propagating beams diagnostics
Coherent Cherenkov diffraction radiation
Pickup design
Test results
Operational electronics
Future outlook

## Slide 5: Motivation – AWAKE Experiment
AWAKE: Proton-driven plasma acceleration
SPECTROMETER
DUMP
SPS protons
ELECTRON LINAC
Electron BPMs
Proton BPMs
Accelerated Beam diagnostics
Plasma cell
LASER
PULSE

## Slide 6: Motivation – AWAKE Experiment
SPECTROMETER
DUMP
SPS protons
ELECTRON LINAC
Electron BPMs
Proton BPMs
Accelerated Beam diagnostics
Plasma cell
LASER
PULSE
INJECTORS
1
SPECTROMETER
DUMP
SPS protons
ELECTRON LINAC
Electron BPMs
Proton BPMs
Accelerated Beam diagnostics
Plasma cell
LASER
PULSE
The AWAKE ingredients
Proton beam 48 nC, 250 ps-σ
Electron beam 600 pC, 4 ps-σ
Laser beam 120 fs, 450 mJ
Rb vapour (later)

## Slide 7: Motivation – AWAKE Experiment
PROTON MICROBUNCHING AND ELECTRON ACCELERATION
2
The proton bunch is broken into a train of microbunches via the self modulation process in the plasma.
See AWAKE Collab., Phys. Rev. Lett. 122, 054802 (2019)
Captured electrons in the plasma wakefields are accelerated.
See AWAKE Collab., Nature 561, 363-367 (2018)

## Slide 8: Motivation – Multiple beams
SPECTROMETER
DUMP
SPS protons
ELECTRON LINAC
Electron BPMs
Proton BPMs
Accelerated Beam diagnostics
Plasma cell
LASER
PULSE
The trajectory control of the laser, electron and proton beam is essential to avoid instabilities and study reproducibility !
Due to the larger intensity, the proton beam saturates all the other diagnostics when present. Need for an electron position measurement in the common beamline.
Common beamline

## Slide 9: Cherenkov Diffraction radiation
Cherenkov Diffraction Radiation is a type of polarization radiation induced in dielectrics when a charged particle passes in proximity to the surface
Very well-defined properties, including the radiation front direction at the Cherenkov angle θCh
See: M. V. Shevelev and A. S. Konkov, JETP 118, 501-511 (2014)
Recent interest to realise non-intercepting diagnostics
See:
Beam
θCh
ChDR emission
h
A. Curcio et al., Phys. Rev. Accel. Beams 23 (2020) 022802 R. Kieffer et al., Phys. Rev. Lett. 121 (2018) 0548802 T. Lefevre et al., Cherenkov Diffraction Radiation as a tool for beam diagnostics, Conf. Proc. IBIC 2019 (2019), Malmo, Sweden
K. Lasocha et al., Experimental Verification of Several Theoretical Models for ChDR Description, Conf. Proc. IPAC’22 (2022), Bangkok, Thailand
εr (ω)

## Slide 10: Kacper test slide
Show linearity

## Slide 11: Coherent Radiation
For any radiation mechanism, we know that the radiation intensity spectrum is
At low frequencies, in the coherent region of the spectrum, one can assume the single particle emission in quasi-phase, hence:
Bunch intensity
Single particle term
Bunch form factor (here Gaussian)

## Slide 12: Coherent Radiation
Protons
Electrons
Different beams feature different bunch form factors. For the AWAKE case:
!
Gaussian beams assumed here ! Real life is not that simple !

## Slide 13: Coherent Radiation
BPMs for p+ detect at 20 MHz
DOMINANT PROTON AREA < 2 GHz
!
Gaussian beams assumed here ! Real life is not that simple !

## Slide 14: Coherent Radiation
BPMs for p+ detect at 20 MHz
BPMs for e- detect at 434 MHz
DOMINANT PROTON AREA < 2 GHz
!
Gaussian beams assumed here ! Real life is not that simple !

## Slide 15: Coherent Radiation
BPMs for p+ detect at 20 MHz
BPMs for e- detect at 434 MHz
DOMINANT PROTON AREA < 2 GHz
DOMINANT ELECTRON AREA > 2 GHz
!
Gaussian beams assumed here ! Real life is not that simple !

## Slide 16: Coherent Radiation
BPMs for p+ detect at 20 MHz
BPMs for e- detect at 434 MHz
DOMINANT PROTON AREA < 2 GHz
Ideal BPMs for e- detect at 20-30 GHz
DOMINANT ELECTRON AREA > 2 GHz
!
Gaussian beams assumed here ! Real life is not that simple !

## Slide 17: Coherent Radiation
BPMs for p+ detect at 20 MHz
BPMs for e- detect at 434 MHz
DOMINANT PROTON AREA < 2 GHz
Ideal BPMs for e- detect at 20-30 GHz
DOMINANT ELECTRON AREA > 2 GHz
PLASMA FREQUENCY > 100 GHz
!
Gaussian beams assumed here ! Real life is not that simple !

## Slide 18: Coherent Radiation – Non Gaussian Beams
ToDO

## Slide 19: Pickup design – Mechanical description
Cylindrical Radiators are inserted in a vacuum chamber.
A brazed collar assure them to be vacuum tight.
The dielectric bar is left protruding outside, to ease the transition into waveguides.

## Slide 20: Pickup design – Radiation generation
Cherenkov Diffraction Radiation is generated at the dielectric surface exposed to the direct beam field.
CST Studio 2018 e- beam, σ = 5ps 50 GHz simul. BW
20 cells/λ mesh
Dielectric insert
Beam axis
The radiator must be oriented at the Cherenkov angle to limit internal reflections.
Other minor diffraction radiation wavefronts are generated due to the interaction with the radiator edges.

## Slide 21: Pickup design – Radiation propagation
Although broadband radiation is generated at the surface, the radiator dimensions determine the output radiation properties.
For cylindrical radiators surrounded by metallic walls, the base mode cutoff frequency is
Radiator radius
Material selection
Below cutoff, the electric field propagation is exponentially damped.

## Slide 22: Pickup design – Radiation propagation simulation
CST simulation
Analytical formula
CST Studio

## Slide 23: Pickup design – Material selection
A large choice of dielectric materials can be used as radiators, such as PTFE, PEEK, Fused silica, Alumina, …
As εr increases:
Larger Cherenkov angles
The low cutoff frequency decreases
PTFE εr = 2.1
PEEK εr = 3.2
SiO2 εr = 3.8
Al2O3 εr = 9.6
Beam
Comparison of different dielectric materials for cylindrical ChDR radiators. The radiator radius is scaled to maintain the same cutoff frequency.
θCh=46⁰
θCh=56⁰
θCh=59⁰
θCh=71⁰

## Slide 24: Pickup design – Real case
Alumina 99.5% pure (εr = 9.4)
Cylindrical rods 90 mm x Ø 6 mm
Metallic coating ~10 µm thick
Cutoff frequency choice fc = 9.7 GHz

## Slide 25
2019 – PTFE Proof of Principle

## Slide 26
2019 – PTFE Proof of Principle
2020/21 – In-air alumina prototype

## Slide 27
2019 – PTFE Proof of Principle
2020/21 – In-air alumina prototype
2022 – Vacuum version

## Slide 28: Test setup
Beam-based test with 100-600 pC single-bunch electron.
Multi-shot measurement, using the remote-controlled attenuator to keep a constant diode input power and limit non-linear response effects.
Data recorded with a 6 GHz scope in the alcove ~20 m away.

## Slide 29: Test results (1)
Similar response to high and low charge with AWAKE parameters
Sensitivity to electrons only
Linear response in a 5 mm range
Limited test time in AWAKE, preliminary test for the operational electronics

## Slide 30: Test results (2)
Proton-only signals below noise level for 1e11 protons per bunch
Proton signal rejection
No signal shape change for simultaneous protons + electrons
Some signal with 3e11 protons per bunch, however depressed with respect to the electrons. Under investigation.

## Slide 31: Operational electronics
A collaboration with TRIUMF was started to realise high-frequency BPM electronics
The elected operation frequency is 30.5 GHz with a bandwidth of 1 GHz.
The system was realised on the model of the existing electron BPMs, to ease integration.
First module installed for tests in 2022.

## Slide 32: Operational electronics
BEAM
BP FILTER f0 30.5 GHz BW 1 GHz
Analog Front-end
TUNNEL SIDE
2.5 GHz
LO
RF
IF
ALCOVE
28 GHz Generator
28 GHz
10 MHz REF
Digitiser
ChDR BPM Pickup
WR28 WAVEGUIDE
COAXIAL CABLE
RED = eBPM
BLACK = ChDR BPM
Initial tests in 2022 currently ongoing.
Need for a better plot than that.

## Slide 33: Future outlook
ChDR pickups are a promising but young technology !
Large signal output. Can measure charges as low as few pC.
PROs:
Difficult to test in the GHz regime. Quality control of the pickups is complicated.
CONs:
High cutoff frequency defined only by the bunch length. Potential for optical techniques with short bunches. See poster TUP022.
Low cutoff frequency decided by construction.
Broadband emission at high frequency, needs careful control when using waveguide hardware.
Requires the difficult design of a broadband transition between the radiator and the waveguides to the detection.
Large impedance, so far tests in single bunch single pass beamlines. Not suitable for counter-rotating beams.
Allow for beam distinction based on frequency discrimination.

## Slide 34: Potential impact
The impact of ChDR techniques go well beyond AWAKE !
Potential use for PWFA, LWFA, xFELs, …

## Slide 35
Thanks for your attention
Acknowledgements: V. Bencini, P. Bestmann, P. Burrows, A. Cherif, N. Chritin, V. Clerc, C. Davut, W. Farabolini, F. Galleazzi, E. Gschwendtner, E. Guran, P. Karataev, J. Kortesmaa, P. Korysko, M. Krupa, K. Lasocha, S. Liu, T. Lefevre, T. Manson, S. Mazzoni, B. Moser, C. Pakuza, E. Poimenidou, P. Muggli, A. Pardons, C. Pasquino, C. Saury, A. Schloegelhofer, P. Schwartz, E. Senes, A. Topaloudis, B. Spear, L. Verra, F. M. Velotti, C. Vendeuvre, V. Verzilov, M. Wendt, B. Woolley, G. Zevi Della Porta
Questions ? eugenio.senes@cern.ch