## Slide 1
Towards pico-seconds
Time Digitization in FPGAs
Jinhong Wang
(jinhongwang@ustc.edu.cn)
June 22, 2024

## Slide 2: Introduction: why the timing precision matters!
2008 Beijing Olympics, 100m results (men)
World records (by now), 100m: 9.58 (men); 10.49 (women)
Milliseconds!
~10 m/s

## Slide 3: Introduction: why the timing precision matters!
Particles travelling at a speed
Approximating the speed of light!
1 ms (10^-3)  ~ 300 km
1 ns (10^-9)  0.3 m (30 cm)
10 ps (10^-11) 0.003 m (3mm)
I
jet
jet
I
Detectors

## Slide 4: Introduction: why the timing precision matters!
BES III (Beijing Spectrometer III) , separate K/π with 2σ at 1GeV/c  100 ps timing precision for the TOF (time of Flight)

## Slide 5: Introduction: why the timing precision matters!
Measuring muon momentum through its bending trajectories
Monitored Drift Tube
Magnet field bends the trajectory of the muons
Momentum is measured through “inner-middle-outer” layers
Trajectory reconstruction

## Slide 6: Introduction: why the timing precision matters!
3D imaging

## Slide 7: Introduction: Examples of timing techniques
A marine sandglass
(seconds-minutes)
A stop watch
(milliseconds)
A XXXXX?
(picoseconds)
Scaled by 1/10^9
In an FPGA
See “Introduction to FPGAs” by Hannes Sakulin

## Slide 8: Introduction: Essentials of timing techniques
Measuring the length of a rectangle
A
B
Starting line
Finishing line
Count by milliseconds
Measure by millimeters
Count at Finer Intervals
So, let’s start counting, in an FPGA!

## Slide 9: Time Digitization in FPGAs: Counter
1
2
3
Clock
Time
digits
Start
Stop
Clock Manager
Logic blocks
min. period of Clock is limited in FPGAs
typ. in nano-seconds ranges min.
Timing resolution corresponds to Tclock
If Tclock is 3.125 ns (320MHz), min.,
what can we do?

## Slide 10: Time Digitization in FPGAs : dual Counter
1
2
3
Clock
Time
digits
Start
Stop
Clock Manager
1
2
3
Clock
180 deg
If Tclock is 3.125 ns, min., what can we do?
Add a 180 deg (inverted) clock @ 3.125ns
Equivalently, 0.5*Tclock resolution is obtained, i.e., 1.5625 ns
Pursuit in this line…

## Slide 11: Time Digitization in FPGAs : phase interpolation
Four phases of 320 MHz, (0, 90, 180, 270), equivalently ~780 ps resolution
Limited by number of finer clock phases (generated from clock managers)
 plus, difficulty in keeping alignment of clock edges (cons)
Pros: simple, and easy to implement with macro blocks
Achieved resolution: ~ 100s pico-seconds
Progress by now: ~ hundreds ps

## Slide 12: Pico-seconds ?
1: identify finer time intervals
(e.g., if a “component”/unit/cell with ~50 ps delay is found, and there are many of them to be added together uniformly)
2: add up to a “working” range
(e.g., add as many components up to 3.125 ns, the “counter” clock period)
Pioneered by Jinyuan Wu @ Fermilab
Cascade chain (2003) https://ieeexplore.ieee.org/stamp/stamp.jsp?tp=&arnumber=1352025
Inspired by Wu’s idea,
A carry-chain version was invented @ USTC
https://ieeexplore.ieee.org/stamp/stamp.jsp?tp=&arnumber=1610982
Nowadays, the delay of typical carry-chain cell is around 10s picoseconds

## Slide 13: With Carry-chain as the“Delay”
c) Carry chain of a multi-bit adder
a) Carry-in in a Slice
b) Rout in a SLICE

## Slide 14: Forming a chain of“Delay”
CYCLONE
In Altera (Intel) FPGAs

## Slide 15: Forming a chain of "Delay”
In Xilinx (AMD) FPGAs

## Slide 16: Principle: tapped-delay-line approach
Divide a clock period into finer slices (DELAY)
 Forming a tapped-delay line (TDL)
https://ieeexplore.ieee.org/document/5446507

## Slide 17: Determine delay time of “Delay”
Average Bin size: Tclk/n
Or, since bin width is proportional to its counts, calculate the bin width for each bin
“Random”
Signals**

## Slide 18: Determine delay time of “Delay”
Average Bin size: Tclk/n
Or, since bin width is proportional to its counts, calculate the bin width for each bin
“Random”
Signals**

## Slide 19: Timing Performance Evaluation
Generation of a stable time interval, pico-seconds stability
  High resolution signal generator (e.g., AWG)
  Delay line approach
Pulse Generator
FPGA
  Cable delay （5 ns @1m）
  Power splitter
A
B
Power splitter splits input pulse into two branches
Each branch share similar time uncertainty, thus could be reduced by measuring their relative difference
A/B are two “identical” channels in FPGA (TDC)  timing uncertainty for a single channel: 1/sqrt(2)
Bin size (average)
Nonlinearity (DNL/INL)
Timing uncertainty (RMS)
…
Progress by now: ~ 20 ps

## Slide 20: The WaveUnion Approach
In
CLK
Ultra-wide Bins
Ultra-Large Bin in the Carry-chain
Boosting the TDL-TDC approach beyond its cell delay
Ref: https://ieeexplore.ieee.org/document/4775079

## Slide 21: WaveUnion A (Jinyuan Wu @Fermilab)
In
CLK
Wave Union
Launcher
0: Hold
0
1: Unleash
1
+
Ultra-wide Bins
Ultra-wide bins are sub-divided.

## Slide 22: WaveUnion A (Jinyuan Wu @Fermilab)
1: Unleash
1
2
1
2
Device: EP2C8T144C6
Plain TDC:
  Max. bin width: 160 ps.
  Average bin width: 60 ps.
Wave Union TDC A:
  Max. bin width: 65 ps.
  Average bin width: 30 ps.

## Slide 23: WaveUnion A (Jinyuan Wu @Fermilab)
Plain TDC:
  delta t RMS width: 40 ps
  25 ps single hit
Wave Union TDC A:
  deltat RMS width: 25 ps
  17 ps single hit

## Slide 24: WaveUnion B (Jinyuan Wu @Fermilab)
Wave Union
Launcher B
In
CLK
1: Oscillate
0: Hold
Does more oscillation cycles always mean better performance?

## Slide 25: WaveUnion B (USTC)
INV+Delay+MUX
Wave Union Launcher
Principle of the 10-ps FPGA TDC
https://ieeexplore.ieee.org/stamp/stamp.jsp?tp=&arnumber=5941022

## Slide 26: WaveUnion B (USTC)
N times Oscillation
Signal Processing of the Raw TDC time

## Slide 27: WaveUnion B (USTC)
Signal Processing of the multi-averaging TDC
RMS timing precision (σdelay) vs. N
• Non-uniformed distribution of the carry chain delay (σcell )
• Random uncertainty of the oscillation period (σosc )
• Other contributors, e.g. the steady of the clock (σother)
Three possible cases:
• Case 1: σosc << σcell
• Case 2: σosc ≈ σcell The best timing @
• Case 3: σosc >> σcell

## Slide 28: WaveUnion B (USTC)
Simulation and Test
Test: RMS vs. N
Case 1: σosc << σcell
Case 2: σosc ≈ σcell
Case 3: σosc >> σcell
3
2
1
Simulation
Actual implementation falls in to Case 2
Progress by now: ~ 10 ps

## Slide 29: WaveUnion B (USTC)
Simulation and Test
Bin size vs. N
Effective Bin size:
Scales as 1/N
No Averaging
N=4
N=8

## Slide 30: WaveUnion B (USTC)
Pros and Cons
√ Larger N results in smaller bin size, lower timing precision
× Larger N results in larger dead time
~ (N+1) * TCLK ,
Trade-off should be made between TDC timing performance and N
No Averaging
N=4
B:300 ps
A:325 ps

## Slide 31: Outlook
sub-10 ps resolution/precision time digitization already achieved in FPGAs (15 years ago)
 is it possible to overcome the 1 ps barrier?
Timing resolution/precision is just one single measure of timing performance, among linearity, power, resource utilization…
 Nowadays, FPGAs typically not limited by configuration logic resources, thus more attention on power efficiency might be necessary.
TDL approach is the most popular one in achieving sub-10 ps resolution, but not the only one, the choice of architecture is a trade off among power, resource utilization …

## Slide 32: Quiz
Top secrets before class…

## Slide 36: What if TDL is not long enough?
Typically it means clock period is relatively large!
Phase interpolation of clocks with TDL

## Slide 38
ACTEL FPGA
  Flash
    IGLOO、、PROAISC3(E)
  Anti-fuse
    AXCELERATOR、SX_A、RTAX_SSL
TDC
    Flash Buffer
      Bin Size~440ps
    Anti-fuse
      Bin Size ~ 80ps

## Slide 39
ACTEL FPGA：A3PE1500