## Slide 1: Fibre Optic Sensors in ATLAS … Keep ATLAS Dry !

## Slide 2: Humidity in the ITk: Keep ATLAS ITk dry
Novel Fibre Optic Sensor (FOS) system
FOS technology
Rad Hard
Long Term Stability
Sufficiently precise
Need Companion conventional sensor for calibration, then they die
R&D, Procurement, Manufacture, QA, Ongoing development during Run 3
Take charge of feed throughs, cabling
Dashboard at PP1 via OPC Server
R&D
Develop FOS technology further
Bare fibre type selections
Grating writing technology
Functional Coating Technology
Environment Chamber for tests, calibrations
Algorithms for readout, compensation, calibration
Modelling
Continuum Fluid Dynamics (CFD) 3D simulation of fluid flow, thermals, leaks
Operational conditions
Dry-out, normal, fault, bake-out
Future
Manufacture, QA and provide throughout ATLAS ITk : 54 Sensors
More R&D to improve Design, Manufacture, Readout, Performance in long term

## Slide 3: Fibre Optic Sensors : FOS ……. a growth area
ISI Web of Science : fibre + optic + sensor

## Slide 4: Fibre Optic Sensors : FOS ……. many types
(from G Berruti, Forum on Tracking Detector Mechanics 2017 3 -5- July, CPPM Marseille)
Option :
Functional Coating
to target sensitivity to specific measurement parameter.

## Slide 5: Fibre Optic Sensors : FOS ……. advantages
Extreme environments
  Dust, moisture, vacuum, temperature (hot / cold), vibration, EM interference, radiation, aggressive chemicals
Many Sensing Modes
  Strain, temperature, humidity, vibration (including 3-D displacements), many specific bio-molecules, radiation dose
Highly Customisable
Sensitivity (enhanced sensitivity to a single environmental parameter and also relative sensitivity between different environmental parameters)
Dynamic range
Multiplexing factor
Interrogator (Expensive / cheaper)
Type of functional coating (what is sensed)
Parameters of the grating (FBG / LPG / no grating)
Type and treatment of fibre (modes, physics mechanism)
Smart
Light, near microscopic in size, can be embedded, can be mass-produced, durable, stable, wavelength and time encoded (not reliant on absolute intensity), optical readout, (very) remote interrogator, 4IR, IoT, ……

## Slide 6: Fibre Optic Sensors
Credit: Maria Konstantaki, Foundation of Research and Technology - Hellas)
MPT 12 fibre feedthrough
Switch

## Slide 7: Fibre Optic Sensors : Fibre Bragg Grating (FBG)
From https://indico.cern.ch/event/837593/

## Slide 8: Fibre Optic Sensors : Long Period Grating (LPG)
TiO2
functional coating : nano-crystalline layer of TiO2 - absorbs H2O reversibly
Leads to both change in L and also nieff,cl
TiO2
environment
from G. Berruti and A. Cusano
functional coating

## Slide 9: FOS Package and Sacrificial Sensors - overview
5 wire Sacrificial Sensors (58)
LPG (H)
FBG (T)
Length 30cm
FBG (D)
Free end as flylead (6m)
Will later splice to flylead from PP1
>15mm
FOS Package (54)
Mount very close to FOS package holes
HIH4000
Pt10k
QA during integration by OTDR
RH measurement by T and Dose compensation

## Slide 10: R&D and Manufacture at CERN

## Slide 11: Fibre Optic Sensors : Long Period Grating (LPG) design
Gaia Maria Berruti (EP-DT-FS)

## Slide 12: Fibre Optic Sensors : Raw Data
LPG (H)
FBG (T)
FBG (D)
FBG (D)
FBG (T)
LPG (H)

## Slide 13: Fibre Optic Sensors : Placements
Recap – Original Number of Sensors … see below …
now 4 Additional FOS + partner Sacrificial + 4 Additional Sacrificial

## Slide 14: ITk Humidity : CFD : Strips
Showing buoyancy driven internal flow
Flush gas is warm and rises
Fluid temperature visualized on a disk
5 mm away from petals

## Slide 15: ITk Humidity : CFD : Outer Pixels
Similar plot for outer pixels
Inlet flows shown

## Slide 16: ITk Humidity : CFD : Inner Pixels
Similar flow pattern in inner pixels
Horizontal plane
Vertical plane

## Slide 17: Continuum Fluid Dynamics  CFD : model the fluid system
Simulate fluid flow, drying out, vapour leaks, thermal conditions, sensor placement, performance
Example from ITk Inner Pixels volume …… In this example … bake-out conditions.

## Slide 18: ITk Humidity : CFD : OSV
Flushing ports in OSV still need optimization in order to get circulation throughout.

## Slide 19: ITk Humidity : CFD : OSV
More R&D to improve Design, Manufacture, Readout, Performance in long term
R&D
Transition the LPG Fibre to a Rad Hard Fibre
  Currently no site of technology that can write in a grating for LPG in a Rad Hard Fibre
    It needs to be tensioned and also azimuthally rotated during write.
    Requires Femto-second laser, stepper rotation and translation motors
  Produce Functional Coating – nano-TiO2 layer – hygroscopic, reversible, in own lab
Develop further the various compensations
  Requires further radiation resting,
  Calibration work in the Environment Chamber
Maintain Integrity and Monitoring of Humidity in ATLAS
  Longer term committmemnt
Spin out the technology
  To SA
  Nuclear Industry
    First of a Kind, in-core, real-time, on-line, monitoring in Nuclear Power Reactors
    Temperature, strain, water level, TID Dose, NIEL Dose
  Space
  Other