## Slide 1
Carlo Tintori (c.tintori@caen.it) picoTDC meeting @ CERN – January 25th, 2024
picoTDC and FERS system

## Slide 2: Building around the picoTDC
Front-End
Readout
Interfaces
Infrastructure
(power, clock, config…)
DAQ
Multi-board Synchronization
T-ref
make the picoTDC ready for the application!

## Slide 3: FERS A5202: 64/128 channel TDC
TDlink:
Eth/USB
2 In + 2 Out
+12V
picoTDC-1
(64 ch)
to picoTDC-2
(128 ch)
LVDS
inputs
FPGA
uC
Clk In/Out
Data+Sync

## Slide 4: Technical Specifications
TDC: 64/128 channels, LSB = 3.125 ps, dynamic range = 56 bit (extended by FPGA)
Acquisition modes: Common Start/Stop (Tref=Ch0), Trigger Matching, Streaming
Readout: Direct: Eth/USB (slow); TDlink + Concentrator: 1/10 GbE or USB 3.0 (up to ~300 MB/s)
FPGA ToT filter: rejects pulses if ToT < LowCut or ToT > HighCut (remove noise, DCR, saturation…)
Throughput: tested up to ~64 Mcps/board (without filters)
Synchronization: up to 128 boards (16k channels) synchronized by Concentrator via TDlink (fibers). Sync to GPS (White Rabbit soon) on Concentrator. Auxiliary Clock Daisy Chain for better timing.
DeltaT Resolution (*) :
  Same board: typ 5 ps RMS
  Board to board (with clock cables): ~8 ps RMS
  Board to board (without clock cables): ~20 ps RMS
(*) Tested with A5256. Pulse: 0.5 Vpp, 0.8 ns rise time

## Slide 5: Input Adapters and Front-End
A5256: Edge Discriminators
A5255: Mechanical adapter (+ Tref)
A5260: Remotization cable
Custom Flange
naked A5203
Detector and
Front-End
Back Panel
Front Panel

## Slide 6: FERS-5200 architecture
1/10G Eth
USB 3.0
Janus SW
1 TDlink => up to 16 FERS (daisy chain)
1 DT5215 => 128 FERS = 8k/16k channels
3.125 Gb/s TDlink
10/100M Eth
USB 2.0
Custom Flange
Desktop Evaluation Setup:
Low Cost, Plug & Play
Easy ASIC integration on FERS
Readout
+
Slow Control
+
Synchronization
Global Time
Large Readout Systems:

## Slide 7: A5204: Radioroc + picoTDC
64 channel all-in-one readout for SiPM: preamp + shapers + peak-ADC + discriminator + TDC + HV bias. Radioroc self-triggers connected in parallel to TDC (for timing) and FPGA (for counting/triggering)
Dynamic Range: 2000 p.e.
Shaping Time: 20÷1200 ns (for Peak ADC)
HV: 20÷80 V pos, 10 mA
Timing Resolution: 55 ps @ 1 p.e.
PHA by ToT: 1 % linearity up to 2000 p.e.
Radioroc
picoTDC
FPGA
Same PCB compatible with:
Psiroc (GEM, Silicon Strips): A5205
Liroc (LIDAR): A52xx

## Slide 8: FERS Roadmap
A5202
Weeroc Citiroc
A5204
Weeroc Radioroc +
CERN picoTDC
A5203
CERN picoTDC
A5205
Weeroc Psiroc +
CERN picoTDC
SiPM
SiPM high-res.
3 ps TDC
GEMs, Si strip
and other low gain detectors

## Slide 9
Test Results:
use of ToT for Walk Correction and Energy Reconstruction

## Slide 10: Test Setup
AFG 3252
Step Attenuator
DT5203+A5256
Pulse: Amp = 3.85 V, Rise = ~2.5 ns
Janus SW
Start on Ch0 with fixed amplitude. Stop on Ch1 and Ch2 (dual threshold) with variable amplitude (max = 3.85 V). Delay = 13 ns.
Sweep: acquire ToT and T at different amplitudes (from 0 to 54 dB, 3 dB step)
Fit points and build ToT-Walk and ToT-Ampl curves
Use curves to correct Walk from ToT (replace CFD)
Use curves to get Amplitude from ToT (make ADC from TDC)

## Slide 11: Calibration curves (Low-Thr = 5 mV, High-Thr = 300 mV)
Poly-Fit
Amp-ToT
Amp-Walk
ToT-Walk
ToT-Amp

## Slide 12: Walk Correction
T (ns)
3850 mV
1217 mV
385 mV
122 mV
38 mV
12 mV
Walk Corrected Peak: =18 ps
Acquired pulses at 6 different amplitudes over a 50 dB dynamic range
The walk causes ~2 ns spread on T: 6 separate peaks appear on the histogram. Timing resolution totally destroyed!
T corrected by ToT using a 5th order polynomial fit of the ToT-Walk points taken at lower threshold (5 mV)
Corrected T histogram presents one single peak:
18 ps RMS over 50 dB dynamic range

## Slide 13: Amplitude Reconstruction
3850 mV
1217 mV
385 mV
122 mV
38 mV
12 mV
Ampl | Centroid | Sigma | Rel Res % | INL %FSR
12.2 | 11.3 | 0.1 | 1.1 | -0.16
38.0 | 41.1 | 1.3 | 3.1 | -0.05
121.8 | 120.9 | 4.0 | 3.3 | -0.14
385.0 | 403.0 | 0.4 | 0.1 | +0.39
1217.5 | 1215.4 | 7.6 | 0.6 | +0.00
3850.0 | 3831.0 | 29.4 | 0.8 | -0.03
ToT-Amp
Vin > 300 -> use High Thr curve (orange) Vin < 300 -> use Low Thr curve (blue)
12
38
122
385
1217
3850
Thr = 5 mV

## Slide 14: Conclusions
The walk correction (mimic CFD) is possible even with a single threshold, with excellent results: 18 ps RMS on a 50 dB dynamic range
The amplitude reconstruction (mimic ADC) requires at least 2 thresholds (2 TDC channels). Linearity = ~0.4%. Resolution = ~3%. Possible improvement with a more accurate threshold setting
Problem: how to build the calibration curves in a real case?
  If possible, use test pulse to generate outputs with known amplitude
  Acquire pulses with a waveform digitizer and emulate the amplitude sweep in software. Need a precise model of the discriminator.
  Machine learning ???
Next steps:
  Sweep emulation of the pulse generated by the pulse generator
  Test with real preamps and detectors

## Slide 15
Use Case:
Picotech ProVision PET scanner

## Slide 16
Picotech ‘ProVision’ PET scanner
picoTDC
2 detectors: each detector uses 12 picoTDC ASICs (768 channels)
need precise timing and TOT measurement
Need a system without any deadtime

## Slide 17
Some details of the Picotech detector module
Pico2023 8 channel differential amplifier-discriminator with amplitude encoded into pulse width

## Slide 18: Readout architecture
(software coincidence trigger)
Need low threshold for precise timing, however SiPMs are noisy - therefore introduce TOT cut of 96 ns (any pulse shorter than 96 ns rejected) - use paired mode readout of picoTDC and reject small TOT in TDC cards.
to do : incorporate coincidence filter in concentrator unit

## Slide 19
‘Paired mode’ readout of picoTDC
19 bit time with 8 bit width (TOT) (timing lsb 3.125 ps ;ΤΟΤ 1.6 ns bins)
[note 219 x 3.125 ps = 1.638 μs - internal trigger period on A5203 card]
‘Empty’ region thanks to TOT filter

## Slide 20
Crosstalk
Time channel 1 - Time channel 2
~ 4% of events where two channels fire within +- 90 ps … there can be an average shift of ~ 30 ps (later channel made earlier)
Possible causes
front-end electronics
PCB/cables transporting signals
picoTDC
study in progress
Tdc bins (3.125ps)

## Slide 21: Thank you for your attention

## Slide 22: Backup slides

## Slide 23: FERS-5200: generic block diagram
ADC
ASIC:
PREAMP
SHAPER
DISCRIM
TRIGGERS
ACQ
LOGIC
READOUT
INTERF
HV BIAS
CONFIG
SYNC
DT5215
ETH 10/100T
USB 2.0
1 GbE
10 GbE
USB 3.0
TDlink
up to 128 FE cards
A5204
TDC
Concentrator
External Sync
(GPS, ExtClk, etc)
COMMON INFRASTRUCTURE
DETECTOR SPECIFIC
MEM
Direct Connect: easy but slow (~2 MB/s)
TDlink: Multi-board Readout + Sync + Config
Up to 300 MB/s

## Slide 24: A5204: block diagram