## Slide 1
Beam Instrumentation & Diagnostics Part 1
CAS Introduction to Accelerator Physics
Santa Susanna, 1st of October 2024
Peter Forck
Gesellschaft für Schwerionenforschnung (GSI)
p.forck@gsi.de
Beam Instrumentation:
Functionality of devices & basic applications
Beam Diagnostics:
Usage of devices for complex measurements

## Slide 2
Demands on Beam Diagnostics
Diagnostics is the ’sensory organs’ for a real beam in a real environment.
(Referring to lecture by Volker Ziemann about ‘Detecting imperfections to enable corrections’)
Different demands lead to different installations:
Quick, non-destructive measurements leading to a single number or simple plots
Used as a check for online information. Reliable technologies have to be used
Example: Current measurement by transformers
Complex instruments for severe malfunctions, accelerator commissioning & development
The instrumentation might be destructive and complex
Example: Emittance determination, tune measurement
General usage of beam instrumentation:
Monitoring of beam parameters for operation, beam alignment & accel. development
Instruments for automatic, active beam control
Example: Closed orbit feedback at synchrotrons using position measurement by BPMs

## Slide 3
Demands on Beam Diagnostics
Diagnostics is the ’sensory organs’ for a real beam in a real environment.
(Referring to lecture by Volker Ziemann about ‘Detecting imperfections to enable corrections’)
Non-invasive ( = ’non-intercepting’ or ‘non-destructive’) methods are preferred:
The beam is not influenced  the same beam can be measured at several locations
The instrument is not destroyed due to high beam power
Instruments could be different for:
Transfer lines with single passage  synchrotrons with multi-passages
Electrons are mostly relativistic  protons are at the beginning non-relativistic
Remark:
Most instrumentation is installed outside of rf-cavities to prevent for signal disturbance

## Slide 4
Typical Installation of a Beam Instrument
Accelerator tunnel:
Local electronics room:
 action of the beam to the detector
 low noise pre-amplifier and first signal shaping
 analog treatment, partly combining other parameters
 digitalization, data bus systems (GPIB, VME, cPCI, µTCA...)
‘long cable’
typ.  100 m
concrete wall
‘long cable’ out
LAN
ADCs
Quadrupole
Linear Ampl.
Beam
Long cable

## Slide 5
Typical Installation of a Beam Instrument
Control room:
 visualization and storage
 parameter setting
Modern trend:
High performance ADC
& digital signal processing
‘long cable’
typ.  100 m
concrete wall
Accelerator tunnel:
 action of the beam to the detector
 low noise pre-amplifier and first signal shaping
Local electronics room:
 analog treatment, partly combining other parameters
 digitalization, data bus systems (GPIB, VME, cPCI, µTCA...)
Beam alignment
by ‘turning knobs’
Happy people at
LHC Sept. 10, 2008
http://www.lhc-facts.ch/

## Slide 6
Outline of the Lectures
The ordering of the subjects is oriented by the beam quantities:
Part 1 of the lecture on electro-magnetic monitors:
Current measurement
Beam position monitors for bunched beams
Part 2 of the lecture on transverse and longitudinal diagnostics:
Profile measurement
Transverse emittance measure
Measurement of longitudinal parameters

## Slide 7
Measurement of Beam Current
The beam current and its time structure the basic quantity of the beam:
It this the first check of the accelerator functionality
It has to be determined in an absolute manner
Important for transmission measurement and to prevent for beam losses.
Different devices are used:
Transformers: Measurement of the beam’s magnetic field
Non-destructive
No dependence on beam type and energy
They have lower detection threshold.
Faraday cups: Measurement of the beam’s electrical charges

## Slide 8
Magnetic Field of the Beam and the ideal Transformer
Beam current of Npart charges with velocity 
cylindrical symmetry
→ only azimuthal component
Example: I =1μA, r =10cm  Bbeam= 2pT, earth Bearth= 50T
Ibeam
vout
R
Torus to guide the magnetic field
 Loaded current transformer
I1/I2= N2/N1  Isec = 1/N · Ibeam
Inductance of a torus of μr
Goal of torus: Large inductance L
and guiding of field lines.
Definition: U = L · dI/dt
Idea: Beam as primary winding and sense by secondary winding

## Slide 9
Fast Current Transformer FCT (also called Passive Transformer)
Simplified electrical circuit of a passively loaded transformer:
beam
A voltages is measured: U = R · Isec = R /N · Ibeam ≡ S · Ibeam
with S sensitivity [V/A] to determine beam current Ibeam
equivalent to transfer function or transfer impedance Z
Equivalent circuit for analysis of sensitivity and bandwidth
(disregarding the loss resistivity RL)

## Slide 10
Response of the Fast Current Transformer: Droop Time
beam current
signal
baseline
Baseline: Ubase  1 - exp(-t/droop)
positive & negative areas are equal
Bunch train:
Time domain description:
A transformer is a ac-coupled device,
i.e. it does not transfer dc-currents
Characterization with droop time:
τdroop= 1/( 2πflow ) = L/R
Effect for bunched beam: Shift of ‘baseline’

## Slide 11
Example for Fast Current Transformer
For bunch beams e.g. during accel. in a synchrotron
typical bandwidth of 2 kHz < f < 1 GHz
 10 ns < tbunch < 1 μs is well suited
From
Company Bergoz
Inner / outer radius | 70 / 90 mm
Permeability | µr  105 for f < 100 kHz µr  1/f above
Windings | 10
Sensitivity | 4 V/A for R = 50 
Droop time droop = L/R | 0.2 ms
Rise time rise= | 1 ns
Bandwidth | 2 kHz ... 500 MHz
Inner / outer radius | 70 / 90 mm
Permeability | µr  105 for f < 100 kHz µr  1/f above
Windings | 10
Sensitivity | 4 V/A for R = 50 
Droop time droop = L/R | 0.2 ms
 | 1 ns
Bandwidth | 2 kHz ... 500 MHz
Example: U73+ from 11 MeV/u ( = 15 %) to 350 MeV/u within 300 ms (displayed every 0.15 ms)
0
1.5
 200 mm
FCT
injection
extraction
synchrotron
Example: GSI Fast Current Transformer FCT:

## Slide 12
Example for Fast Current Transformer
For bunch beams e.g. during accel. in a synchrotron
typical bandwidth of 2 kHz < f < 1 GHz
 10 ns < tbunch < 1 μs is well suited
Inner / outer radius | 70 / 90 mm
Permeability | µr  105 for f < 100 kHz µr  1/f above
Windings | 10
Sensitivity | 4 V/A for R = 50 
Droop time droop = L/R | 0.2 ms
Rise time rise= | 1 ns
Bandwidth | 2 kHz ... 500 MHz
Inner / outer radius | 70 / 90 mm
Permeability | µr  105 for f < 100 kHz µr  1/f above
Windings | 10
Sensitivity | 4 V/A for R = 50 
Droop time droop = L/R | 0.2 ms
 | 1 ns
Bandwidth | 2 kHz ... 500 MHz
Example: U73+ from 11 MeV/u ( = 15 %) to 350 MeV/u within 300 ms (displayed every 0.15 ms)
0
1.5
FCT
injection
synchrotron
More examples see lecture
‘Longitudinal Beam Dynamics’
by Frank Tecker & Heiko Damerau
Example: GSI Fast Current Transformer FCT:
Numerous application e.g.:
Transmission optimization
Bunch shape measurement
Input for synchronization
of ‘beam phase’
extraction

## Slide 13
How to measure the DC current? The current transformer discussed sees only B-flux changes.
The DC Current Transformer (DCCT)  magnetic saturation of two torii.
Modulation of the primary windings forces both torii into saturation
twice per cycle
Sense windings measure the modulation signal and cancel each other.
But with the Ibeam, the saturation is shifted and Isense is not zero
Compensation current adjustable until Isense is zero once again
Remark:
Same principle installed in power supplier
The dc Transformer DCCT
Depictive statement:
A single transformer needs varying beam. The trick is to ‘switch two transformers’!
Notes: 13

## Slide 14
The dc Transformer
Modulation without beam:
typically about 9 kHz to saturation → no net flux
Modulation with beam:
  saturation is reached at different times, → net flux
Net flux: double frequency than modulation
Feedback: Current fed to compensation winding
for larger sensitivity
Two magnetic cores: Must be very similar.
Remark: Same principle used for power suppliers
Notes: 14

## Slide 15
Application for dc transformer:
Observation of beam behavior with typ. 20 µs time resolution → the basic operation tool
Measurement with a dc Transformer
Important parameter:
Detection threshold:  1 µA
(= resolution)
Bandwidth: f = dc to 20 kHz
Rise-time: trise = 20 µs
Temperature drift: 1.5 µA/0C
 compensation required.
injection
extraction
DCCT
synchrotron
Example: The DCCT at GSI synchrotron
U73+ accelerated from
11. 4 MeV/u ( = 15.5%) to 750 MeV/u ( = 84 %)
For slow extraction: See lecture ‘Injection and Extraction’ by Pablo Arrutia
Notes: 15

## Slide 16
Measurement of Beam Current
Transformers: Measurement of the beam’s magnetic field
Non-destructive
No dependence on beam type and energy
They have lower detection threshold.
Faraday cups: Measurement of the beam’s electrical charges
They are destructive
For low energies only
Low currents can be determined.

## Slide 17
Semi-classical approach:
Projectiles of mass M collide
with free electrons of mass m
If M >> m then the relative energy transfer is low
many collisions required many elections participate
proportional to target electron density
low straggling for the heavy projectile i.e. ‘straight trajectory’
If projectile velocity   1 low relative energy change of projectile ( is Lorentz factor)
I is mean ionization potential including kinematic corrections I  Zt  10 eV for most metals
Strong dependence an projectile charge Zp as
Constants: NA Advogadro number, re classical e- radius, me electron mass, c velocity of light
Bethe-Bloch formula:
(simplest formulation)
Excurse: Energy Loss of Protons & Ions
Maximum energy transfer from projectile M to electron me:
e-
e-
`δ-ray’
beam, charge Zp
mass M
Material
Zt , At , t
Notes: 17

## Slide 18
Source
LINAC, Cycl.
Synchrotron
Excurse: Energy Loss of Protons & Ions in Copper
Numerical calculation for ions
with semi-empirical model e.g. SRIM
Main modification
 Cups only for
Ekin < 100 MeV/u due to R < 10 mm
Range:
with approx. scaling
Bethe-Bloch formula:
(simplest formulation)
Approximation e.g.
Notes: 18

## Slide 19
19
Excurse: Secondary Electron Emission caused by Ion Impact
Energy loss of ions in metals close to a surface:
Closed collision with large energy transfer:  fast e- with Ekin > 100 eV
Distant collision with low energy transfer  slow e- with Ekin  10 eV
‘diffusion’ & scattering with other e-: scattering length Ls  1 - 10 nm
at surface  90 % probability for escape
Secondary electron yield and energy distribution comparable for all metals!
 Y = const. * dE/dx (Sternglass formula)
Electrons per ion
Different targets:
E.J. Sternglass, Phys. Rev. 108, 1 (1957)
beam
Ls  10 nm
e-
e-
`δ-ray’
Material
Zt , At , t
Notes: 19

## Slide 20
Faraday Cups for Beam Charge Measurement
The beam particles are collected inside a metal cup
 The beam’s charge are recorded as a function of time.
The cup is moved in the beam pass  destructive device
Currents down to 10 pA with bandwidth of 100 Hz!
To prevent for secondary electrons leaving the cup
Magnetic field: The central field is B  10 mT
for  mm.
or Electric field: Potential barrier at the cup entrance U  1 kV.
50mm
permanent magnet
HV electrode
radius rc

## Slide 21
Realization of a Faraday Cup at GSI LINAC
The Cup is moved into the beam pass.
LINAC
RFQ
source
Cup: beam stopped
pneumatic
drive
vacuum flange
here 150 mm
Faraday Cup
60 mm
bellow
compression
for movement
 1m
beam

## Slide 22
Summary for Current Measurement
Transformer:  measurement of the beam’s magnetic field
 Magnetic field is guided by a high μ toroid
 Types: FCT  large bandwidth, Imin 30 µA, BW = 10 kHz ... 500 MHz
[ACT : Imin 0.3 µA, BW = 10 Hz .... 1 MHz, used at proton LINACs ]
DCCT: two toroids + modulation, Imin 1 µA, BW = dc ... 20 kHz
 Non-destructive, used for all beams
Faraday cup:  measurement of beam’s charge,
 Low threshold by I/U-converter: Ibeam > 10 pA
 Totally destructive, used for low energy beams only
Active transformer ACT
Fast Transformer FCT
DC transformer DCCT
Company Bergoz
Resolution limit
Notes: 22

## Slide 23
Example  Synchrotron Light Facility ALBA
LINAC 100 MeV
Booster 3 GeV
Storage ring
C = 268 m
LINAC 100 MeV
Storage Ring:
3 GeV
Booster Ring:
0.1  3 GeV
Layout:
Storage ring length: 268 m
Electron energy: 3 GeV
Top-up injection
Max. beam current: 0.4 A
3rd generation Spanish synchr. light facility in Barcelona

## Slide 24
FCT
DCCT
FCUP
AE
BCM
Transfer line
1 FCT
1 FCUP
3 BCM
BOOSTER
1 FCT
1 DCCT
1 AE
Transfer line
2 FCT
Storage Ring
1 FCT
1 DCCT
1 AE
Beam current measurements:
Several in transport lines
One per ring
Abbreviation:
FCT: Fast Current Transformer
DCCT: dc transformer
FCUP: Faraday Cup
AE: Annular Electrode
BCM: Bunch Charge Monitor
Remark:
AE: Annular Electrode
i.e. circular electrode acting
like a high frequency pick-up
From U. Iriso, ALBA
Appendix: The Synchrotron Light Facility ALBA: Current Meas.
LINAC
100 MeV
DCCT
FCT
beam

## Slide 25
Outline:
Signal generation  transfer impedance
Capacitive button BPM for high frequencies
Electronics for position evaluation
BPMs for measurement
Summary
Pick-Ups for bunched Beams
A Beam Position Monitor is an non-destructive device for bunched beams
It delivers information about the transverse center of the beam:
Trajectory: Position of an individual bunch within a transfer line or synchrotron
Closed orbit: Central orbit averaged over a period much longer than a betatron oscillation
Single bunch position: Determination of parameters like tune, chromaticity, -function
Remarks: - BPMs have a low cut-off frequency  dc-beam can’t be monitored
- The abbreviation BPM and pick-up PU are synonyms

## Slide 26
Ibeam
Time Domain  Frequency Domain: Instrumentation
Time domain: Recording of a voltage as a function of time:
Frequency domain: Displaying of a voltage as a function of frequency:
Instrument:
Spectrum Analyzer
Instrument:
Oscilloscope
courtesy company Keysight
voltage
time
frequency
See lecture ‘Time and Frequency Domain Signals’ by Hermann Schmickler
Photos and idea by Piotr Kowina
Fields seems unstructured
Other perspective:
Fields is structured

## Slide 27
Ibeam
Time Domain  Frequency Domain: Mathematics
Time domain: Recording of a voltage as a function of time:
Frequency domain: Displaying of a voltage as a function of frequency:
Mathematics for function :
Fourier Transformation:
Fourier Transformation:
Contains amplitude & phase
as the values are (complex in math. sense)
The same information is displayed differently
Law of Convolution: For a convolution in time:
  convolution be expressed as multiplication of FTs
See lecture ‘Time and Frequency Domain Signals’ by Hermann Schmickler

## Slide 28
Pick-Ups for bunched Beams
The image current at the beam pipe is monitored on a high frequency basis
i.e. the ac-part given by the bunched beam.
Beam Position Monitor BPM is the most frequently used instrument!
For relativistic velocities,
the electric field is transversal:

## Slide 29
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
V
-
-
-
-
-
-
-
Principle of Signal Generation of a BPMs, centered Beam
Animation by Rhodri Jones (CERN)
The image current at the wall is
monitored on a high frequency basis
i.e. ac-part given by the bunched beam.

## Slide 30
Model for Signal Treatment of capacitive BPMs
Equivalent circuit
The wall current is monitored by a plate or ring inserted in the beam pipe:
Ubeam
ground
CC
C
R
Ibeam
capacitive coupling
 image current
electronics elements
Uim
At a resistor R the voltage Uim from the image current is measured.
Goal: Connection from beam current to signal strength by transfer impedance Zt ( )
in frequency domain: Uim(ω) = R · Iim(ω) = Zt(ω) · Ibeam(ω)
Result: i.e. complex function
geometry
stray capacitance
frequency response

## Slide 31
Example of Transfer Impedance for Proton Synchrotron
The high-pass characteristic for typical synchrotron BPM:
Uim(ω)=Zt(ω)·Ibeam(ω)
Large signal strength for long bunches  high impedance
Smooth signal transmission important for short bunches  50 
Parameter linear-cut BPM at proton synchr.:
C = 100pF, l = 10cm,  = 50%
fcut = ω/2π = (2πRC)-1
for R = 50   fcut = 32 MHz
for R = 1 M  fcut = 1.6 kHz
Remark: For   0 it is Zt  0 i.e. no signal is transferred from dc-beams e.g.
 de-bunched beam inside a synchrotron
 for slow extraction through a transfer line

## Slide 32
Calculation of Signal Shape (here single Bunch)
The transfer impedance is used in frequency domain! The following is performed:
1. Start: Time domain Gaussian function Ibeam(t) having a width of σt
2. FFT of Ibeam(t) leads to the frequency domain Gaussian Ibeam(f) with
3. Multiplication with Zt(f) with fcut = 32 MHz leads to Uim(f) = Zt(f) ∙ Ibeam(f)
4. Inverse FFT leads to Uim(t)
Fourier trans.
inverse Fourier trans.
Remark: Time domain processing via convolution or filters (FIR and IIR) are possible

## Slide 33
Calculation of Signal Shape: repetitive Bunch in a Synchrotron
Fourier spectrum is concentrated at acceleration harmonics
with single bunch spectrum as an envelope.
Bandwidth up to typically 10*facc
Synchrotron filled with 8 bunches accelerated with facc= 1 MHz
BPM terminated with R=50   facc<< fcut :
Parameter: R=50   fcut=32 MHz, all buckets filled
C=100pF, l=10cm,  =50%, σt=100 ns  σl=15m
See lecture ‘Time and Frequency Domain Signals’ by Hermann Schmickler

## Slide 34
Calculation of Signal Shape: Bunch Train with empty Buckets
Fourier spectrum is more complex, harmonics are broader due to sidebands
Synchrotron during filling: Empty buckets, R=50 :
BPM terminated with R=50   facc<< fcut
Parameter: R=50   fcut=32 MHz, 2 empty buckets
C=100pF, l=10cm,  =50%, σt=100 ns  σl=15m

## Slide 35
Calculation of Signal Shape: Filtering of Harmonics
Remark: For numerical calculations, time domain filters (FIR and IIR) are more appropriate
Generally: Ztot(ω) = Hcable(ω) ∙ Hfilter(ω) ∙ Hamp(ω) ∙ … ∙Zt(ω)
Effect of filters, here 4th order Butterworth bandpass filter:
and
Only few frequency components leading to ‘ringing’ due to sharp cutoff
Other filter types more appropriate
Parameter: R=50 , 4th order Butterworth filter at flow=1.8 MHz & fhigh=2.2 MHz
C=100pF, l=10cm,  =50%, σt=100 ns  σl=15m

## Slide 36
-
-
-
-
-
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
V
-
-
-
-
-
-
-
-
+
+
+
+
+
-
-
+
-
+
-
+
-
+
-
-
-
-
+
+
+
-
+
+
-
+
-
-
+
-
+
-
-
-
-
-
+
+
+
-
+
+
-
-
+
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
Animation by Rhodri Jones (CERN)
Principle of Signal Generation of a BPMs: off-center Beam
The image current at the wall is
monitored on a high frequency basis
i.e. ac-part given by the bunched beam.

## Slide 37
Principle of Position Determination by a BPM
The difference voltage between plates gives the beam’s center-of-mass
most frequent application
S(ω,x) is called position sensitivity, sometimes the inverse is used k(ω,x)=1/S(ω,x)
S is a geometry dependent, non-linear function,
Units: S=[%/mm], sometimes S=[dB/mm] or k=[mm].
It is at least:
y from ΔU = Uup - Udown
Typical desired position resolution:
x  0.1 ... 0.3  x of beam width
Example: One turn = 4 bunches @ 35 MeV/u
Uright
Ureft

## Slide 38
Outline:
Signal generation  transfer impedance
Capacitive button BPM for high frequencies
used at most proton LINACs and electron accelerators
Electronics for position evaluation
BPMs for measurement of closed orbit, tune and further lattice functions
Summary
Pick-Ups for bunched Beams

## Slide 39
2-dim Model for a Button BPM
‘Proximity effect’: larger signal for closer plate
Ideal 2-dim model: Cylindrical pipe  image current density
via ‘image charge method’ for ‘pencil’ beam:
Image current: Integration over finite button size:
,  = 0o


a
r
button
 =0

## Slide 40
Ideal 2-dim model: Non-linear behavior and hor-vert coupling:
Sensitivity S converts signal to position
with S = [%/mm] or [dB/mm]
i.e. S is the derivative of the curve , here Sx = Sx(x, y) i.e. non-linear.
For this example: central part S=7.4%/mm  k=1/S=14mm
button
 real position
• measured position
Position Map
Horizontal plane
2-dim Model for a Button BPM


a
r
button


## Slide 41
Button BPM Realization
LINACs, e--synchrotrons: 100 MHz < frf < 3 GHz  bunch length  BPM length
 50  signal path to prevent reflections
Example: LHC-type inside cryostat:
24 mm, half aperture a = 25 mm, C = 8 pF
 fcut= 400 MHz, Zt = 1.3  above fcut
Ø24 mm

5cm
N-Connector
Button electrode
Ceramic vacuum insulation
Courtesy C. Boccard (CERN)

## Slide 42
Simulations for Button BPM at Synchrotron Light Sources
Result: - Non-linearity and xy-coupling occur in dependence of button size and position
- Can be corrected by polynomial interpolation for beams much smaller than chamber
real beam position x [mm]
y = 20 mm
y = 10 mm
y = 0 mm
20
Sx(center) = 7.8 %/mm
for |x|<5mm & y=0mm
Sy(center) = 7.2 %/mm
for |y|< 5mm & x=0mm
button 2
position x [mm]
30
button 1
button 3
button 4
30
position x [mm]
button 1
button 2
button 4
button 3
28 mm
72 mm
 7 mm
button 1
button 2
button 3
button 4
18 mm
y = 20 mm
y = 10 mm
y = 0 mm
1 mm
steps
1 mm steps
y
x
A.A. Nosych et al.,
IBIC’14
Optimization:
Horizontal distance
Size of buttons
Horizontal: 
Vertical: y  y
Example: Simulation for ALBA light source for 72 x 28 mm2 chamber
linear-cut BPM
stripline BPM

## Slide 43
Outline:
Signal generation  transfer impedance
Capacitive button BPM for high frequencies
used at most proton LINACs and electron accelerators
Electronics for position evaluation
analog signal conditioning to achieve small signal processing
BPMs for measurement of closed orbit, tune and further lattice functions
Summary
Pick-Ups for bunched Beams

## Slide 44
Broadband Signal Processing
Hybrid or transformer close to beam pipe for analog ΔU & ΣU generation or Uleft & Uright
Attenuator/amplifier
Filter to get the wanted harmonics and to suppress stray signals
ADC: digitalization  followed by calculation of of ΔU /ΣU
.
Advantage: Bunch-by-bunch observation possible, versatile post-processing possible
Disadvantage: Resolution down to  100 µm for large aperture, i.e. 0.1% of aperture,
resolution is worse than narrowband processing, see below
Challenge: Precise analog electronics with very low drift of amplification

## Slide 45
General: Noise Consideration
1. Signal voltage given by:
2. Thermal noise voltage given by:
Example: GSI-LINAC with frf=36 MHz
Signal-to-noise ΔUim/Unoise is influenced by:
Input signal amplitude
Thermal noise from amplifiers etc.
Bandwidth Δf
 Restriction of frequency width
as the power is
concentrated at harmonics n frf

## Slide 46
Narrowband Processing for improved Signal-to-Noise
Narrowband processing equals heterodyne receiver (e.g. AM-radio or analog spectrum analyzer)
Attenuator/amplifier
Mixing with accelerating frequency fLO
 IF-output: signal with difference frequency fIF = fLO – fRF
Bandpass filter of the mixed signal (e.g at 10.7 MHz)
Rectifier: synchronous detector
ADC: digitalization → followed calculation of ΔU/ΣU
Advantage: Spatial resolution about 100 time better than broadband processing
Disadvantage: No turn-by-turn diagnosis, due to mixing = ’long averaging time’
Digital
correspondence:
I/Q demodulation

## Slide 47
Analog versus Digital Signal Processing
Digital receiver as modern successor of super heterodyne receiver
Basic functionality is preserved but implementation is very different
Digital transition just after the amplifier & filter or mixing unit
Signal conditioning (filter, decimation, averaging) on FPGA
Advantage of DSP: Versatile operation, flexible adoption without hardware modification
Disadvantage of DSP: non, good engineering skill requires for development, expensive
Analog frequency
translator
Analog demodulator
and filter
ADC
BPM analog
signal
digital
output
Analog frequency
translator
ADC
Digital
filter
Digital
Signal Proc.
digital
output
BPM analog
signal
Modern digital processing
Traditional analog processing
Analog
Digital
Modern instrumentation uses digital techniques with extended functionality.

## Slide 48
Comparison of BPM Readout Electronics (simplified)
Type | Usage | Precaution | Advantage | Disadvantage
Broadband | p-sychr. | Long bunches | Bunch structure signal Post-processing possible Required for transfer lines with few bunches | Resolution limited by noise
Narrowband | all synchr. | Stable beams >100 rf-periods | High resolution | No turn-by-turn Complex electronics
Digital Signal Processing | all | ADC sample typ. 250 MS/s | Very flexible & versatile High resolution Trendsetting technology for future demands | Basically non! Limited time resolution by ADC  under-sampling Man-power intensive
 Stripline BPM

## Slide 49
Outline:
Signal generation  transfer impedance
Capacitive button BPM for high frequencies
used at most proton LINACs and electron accelerators
Electronics for position evaluation
analog signal conditioning to achieve small signal processing
BPMs for measurement of closed orbit, tune and further lattice functions
frequent application of BPMs
Summary
Pick-Ups for bunched Beams

## Slide 50
Trajectory:
The position delivered by an individual bunch within a transfer line or a synchrotron.
Main task: Control of matching (center and angle), first-turn diagnostics
Example: LHC injection 10/09/08 i.e. first day of operation !
Courtesy R. Jones (CERN)
-10
-10
horizontal
vertical
Monitor number (530 BPMs on 27 km)
10
-10
0
x [mm]
10
-10
0
y [mm]
Tune values at LHC: Qh = 64.3, Qv = 59.3
Trajectory Measurement with BPMs

## Slide 51
Closed Orbit Feedback: Typical Noise Sources
Courtesy M. Böge, PSI, N. Hubert, Soleil
Beam movement:
Short term (min to 10 ms):
Traffic
Machine (crane) movements
Water & vacuum pumps
50 Hz main power net
Medium term (day to min):
Movement of chambers
due to heating by radiation
Day-night variation
Tide, moon cycle
Long term ( > days):
Ground settlement
Seasons, temperature variation
1
0.1
100
10
1000
Frequency [Hz]
Power spectral density [mm2/Hz]
0.01

## Slide 52
Close Orbit Feedback: BPMs and magnetic Corrector Hardware
Orbit feedback: Synchrotron light source  spatial stability of light beam
Example: SLS-Synchrotron at Villigen, Switzerland
Feedback loop:
1. Position from all BPMs
2. Calculation of corrector setting via
Orbit Response Matrix
3. Change of magnet setting
1.’ New positon measurement …..
 regulation time down to 10 ms
 Role od thumb:  4 BPMs per betatron wavelength
From M. Böge, PSI
Uncorrected orbit: typ.
Corrected orbit: typ. up to  100 Hz bandwidth!
Position from all BPMs
Calculation of corrector strength
Setting of correctors
feedback
Acc. optics
Orbit Response Matrix: See lecture ‘Imperfections and Corrections’ by Volker Ziemann
Conclusion

## Slide 53
Graphics by R. Singh, GSI
Tune Measurement: General Considerations
Beam particle’s in-coherent motion  center-of-mass stays constant
Coherent excitations are required for the detection by a BPM
Excitation of all particles by rf  coherent motion
 center-of-mass variation turn-by-turn i.e. center acts as one macro-particle
Notes: Beam with several particles with randomly distributed phases and amplitude will give only very small fluctuations due to finite number of particles (schottky fluctuations) proportional to sqrt(N)
Beam excitation aligns the particle phases, and allows to measure center-of-mass before the beam decohores due to finite spread in frequency. Thus regular excitation needed. Types of excitation were studied as part of this work, Kick excitation (a wide band, or single time impulse excitation), requires high power, and significantly dilutes the phase space.
Chirp excitation, slow method, difficult to use in day to day operations, BTF excitation, to study the amplitude and phase response of the beam.
A compromise of the both the above methods, band limited (or pseudo random noise) for regular excitation without significant blow-up, usable along the ramp.

## Slide 54
Tune Measurement: The Kick-Method in Time Domain
The beam is excited to
coherent betatron oscillation:
 Beam position measured
each revolution (’turn-by-turn’)
 Fourier Trans. gives non-integer tune q.
Short kick compared to revolution.
Here: N = 200 turn  q > 0.003
(tune spreads can be q  0.001!)
The de-coherence time limits the resolution:
N non-zero samples
 General limit of discrete FFT:
Decay is caused by de-phasing,
not by decreasing
single particle amplitude.
See lecture ‘Time and Frequency Domain Signals’ by Hermann Schmickler

## Slide 55
Tune Measurement: Frequency Chirp Measurement
SIS-synchrotron: A wide scan with
both sidebands at h=25th-harmonics:
From the position of the sidebands q = 0.306 is determined.
Advantage: High resolution for tune and tune spread (also for de-bunched beams)
Disadvantage: Long sweep time (up to several seconds).
A detailed scan for the lower sideband
 beam acts like a driven oscillator:
Principle: Slowly scan of the excitation frequency  beam acts as driven oscillator!
(sometimes refer to as Beam Transfer Function BTF measurement)


## Slide 56
Tune Measurement: Gentle Excitation with Wideband Noise
Instead of a sine wave, noise with adequate bandwidth can be applied
 beam picks out its resonance frequency:
Broadband excitation with white noise
of  10 kHz bandwidth
Turn-by-turn position measurement
Fourier transformation of the recorded data
 Continues monitoring with low disturbance
Example: Vertical tune within 4096 turn
duration ≃ 15 ms
at GSI synchrotron 11 → 300 MeV/u in 0.7 s
Advantage:
Fast scan with good time resolution
Disadvantage: Lower tune resolution
vertical tune versus time
Qy  0.02
time [ms]
vert. fractional tune qy
750
600
400
0
200
0.2
0.25
0.3
excitation
noise band
Q  0.05
vertical tune at fixed time  15ms
vertical fractional tune qy
0.2
0.25
0.35
0.3
amplitude [a.u.]
U. Rauch et al., DIPAC 2009
Conclusion

## Slide 57
Chromaticity ξ: Change of tune for off-momentum particle
Two step measurement procedure:
1. Change of momentum p by detuned rf-frequency
2. Excitation of coherent betatron oscillations
and tune measurement
(kick-method, BTF, noise excitation):
Chromaticity Measurement from Closed Orbit Data
Plot of ΔQ/Q as a function of Δp/p
 slope is dispersion ξ.
Example: Measurement at LEP:
From M Minty, F. Zimmermann,
Measurement and Control of charged Particle Beam,
Springer Verlag 2003
Conclusion

## Slide 58
From J. Wenninger: CAS on BD, CERN-2009-005 & J. Wenninger CERN-AB-2004-009
Dispersion Measurement from Closed Orbit Data
Dispersion D(si ): Change of momentum p by detuned rf-cavity
 Position reading at one location :
 Result from plot of xi as a function of Δp/p  slope is local dispersion D(si )
Example: Dispersion measurement D(s) at BPMs at CERN SPS
Theory-experiment correspondence
after correction of
BPM calibration
quadrupole calibration
: model expectation
green histogram: measurement
6.9 km
See lecture ‘Imperfections and Corrections’ by Volker Ziemann
Conclusion

## Slide 59
Intra-Bunch Observation
Courtesy Kevin Li, CAS Proceedings 2022
See lecture ‘Collective Effects’ by Kevin Li
High band-width measurements delivers:
Bunch shape given by the sum of two plates
Intra-bunch movement of the center by
A single trace corresponds to one bunch passage
Time 4.25 ns
Example: Single bunch observation on turn-by-turn basis with beam excitation at SPS
Goal: Monitoring instabilities
Conclusion
stripline BPM

## Slide 60
LTB:
3 BPM
BTS:
4 BPM
SEXC
STORAGE RING:
7 or 8 BPM/cell
(120 BPMs in total)
1 SEXC
BOOSTER:
11 BPM/quadrant
(44 BPMs in total)
1 SMES
1 SEXC
SEXC
SMES
Beam position:
Many locations!
Frequent operating tool
For position stabilization
i.e. closed orbit feedback
Tune: Qx = 18.18 & Qy = 8.37
From U. Iriso, ALBA
Appendix: Synchrotron Light F. ALBA: Position, tune etc. Measure
Button BPMs  BPM
Meas. Stripline SMES
Exc. Stripline  SEXC
7 mm
High: 28 mm
BLM

## Slide 61
Summary Pick-Ups for bunched Beams
The electric field is monitored for bunched beams using rf-technologies
(’frequency domain’). Beside transformers they are the most often used instruments!
Differentiated or proportional signal: rf-bandwidth ↔ beam parameters
Proton synchrotron: 1 to 100 MHz, mostly 1 M  proportional shape
LINAC, e−-synchrotron: 0.1 to 3 GHz, 50   differentiated shape
Important quantity: Transfer impedance .
Types of capacitive pick-ups:
Linear-cut (p-synch.), button (p-LINAC, e−-LINAC and synch.)
Position reading: Difference signal of two or four pick-up plates (BPM):
Non-intercepting reading of center-of-mass → online measurement and control
Synchrotron: Fast reading, ‘bunch-by-bunch’ trajectory, slow reading  closed orbit
 Synchrotron: Excitation of coherent betatron oscillations  tune q,  , (s), D(s)...
Remark: BPMs have high pass characteristic  no signal for dc-beams
Thank you for your attention!

## Slide 62
Backup slides

## Slide 63
Example: The DCCT at GSI synchrotron
63
The dc Transformer Realization
dc transformer
2 cores mounted
ac transformers
(two types)
magnetic shield
200 mm flange
ceramic gap
Torus radii | ri = 135 mm ro=145 mm
Torus thickness | d = 10 mm
Torus permeability | µr = 105
Saturation inductance | Bsat = 0.6 T
Number of windings | 16 for modulation & sensing 12 for feedback
Resolution | Iminbeam = 2 µA
Bandwidth | f = dc .... 20 kHz
Rise time constant | rise = 10 µs
Temperature drift | 1.5 µA/oC
Notes: 63

## Slide 64
Shielding of a Transformer
Task of the shield:
The image current of the walls have to be bypassed by a gap and a metal housing.
This housing uses µ-metal and acts as a shield of external B-field
(remember: Ibeam = 1 μA, r = 10 cm  Bbeam = 2pT, earth field Bearth = 50 µT)
vacuum pipe  150 mm
ceramic gap
magnetic shield &
current bypass
transformer
torus
beam
courtesy Company Bergoz
beam
 200 mm
ceramic
gap

## Slide 65
Outline:
Signal generation  transfer impedance
Capacitive button BPM for high frequencies
used at most proton LINACs and electron accelerators
Capacitive linear-cut BPM for low frequencies
used at most proton synchrotrons due to linear position reading
Electronics for position evaluation
BPMs for measurement of closed orbit, tune and further lattice functions
Summary
Pick-Ups for bunched Beams

## Slide 66
Linear-cut BPM for Proton Synchrotrons
Frequency range: 1 MHz < frf < 100 MHz  bunch-length >> BPM length.
Signal is proportional to actual plate length:
In ideal case: linear reading
Linear-cut BPM:
Advantage: Linear, i.e. constant position sensitivity S
 no beam size dependence
Disadvantage: Large size, complex mechanics
high capacitance
Size: 200x70 mm2

## Slide 67
Technical Realization of a linear-cut BPM
Quadrupole
Linear Ampl.
Beam
Technical realization at HIT synchrotron of 46 m length for 7 MeV/u 440 MeV/u
BPM clearance: 180x70 mm2, standard beam pipe diameter: 200 mm.
70mm
180mm
Linear Ampl.
Horizontal
electrode
Right
channel

## Slide 68
70mm
180mm
Linear Ampl.
Horizontal
electrode
Right
channel
Technical Realization of a linear-cut BPM
Technical realization at HIT synchrotron of 46 m length for 7 MeV/u 440 MeV/u
BPM clearance: 180x70 mm2, standard beam pipe diameter: 200 mm.

## Slide 69
Comparison linear-cut and Button BPM
 | Linear-cut BPM | Button BPM
Precaution | Bunches longer than BPM | Bunch length comparable to BPM
BPM length (typical) | 10 to 20 cm length per plane | 1 to 5 cm per button
Shape | Rectangular or cut cylinder | Orthogonal or planar orientation
Bandwidth (typical) | 0.1 to 100 MHz | 100 MHz to 5 GHz
Coupling | 1 M or 1 k (transformer) | 50 
Cutoff frequency (typical) | 0.01… 10 MHz (C=30…100pF) | 0.3… 1 GHz (C=2…10pF)
Linearity | Very good, no x-y coupling | Non-linear, x-y coupling
Sensitivity | Good, care: plate cross talk | Good, care: signal matching
Usage | At proton synchrotrons, frf < 10 MHz | All electron acc., proton Linacs, frf > 100 MHz
Remark: Other types are also some time used: e.g. strip-line, wall current monitors,
inductive antenna, BPMs with external resonator, cavity BPM, slotted wave-guides etc.
electronics

## Slide 70
Outline:
Signal generation  transfer impedance
Capacitive button BPM for high frequencies
used at most proton LINACs and electron accelerators
Stripline BPMs for hight frequencies
used at colliders with counter-propagating beams
Electronics for position evaluation
BPMs for measurement of closed orbit, tune and further lattice functions
Summary
Pick-Ups for bunched Beams

## Slide 71
Stripline BPM: General Idea
For short bunches, the capacitive button deforms the signal
Relativistic beam   1  field of bunches nearly TEM wave
Bunch’s electro-magnetic field induces a traveling pulse at the strips
Assumption: Bunch shorter than BPM, Zstrip = R1 = R2=50  and vbeam = cstrip
LHC stripline BPM, l = 12 cm
From C. Boccard, CERN

## Slide 72
Stripline BPM: General Idea
For relativistic beam with   1 and short bunches:
Bunch’s electro-magnetic field induces a traveling pulse at the strip
Assumption: lbunch<< l, Zstrip=R1=R2=50  and vbeam=cstrip
Signal treatment at upstream port 1:
Signal at downstream port 2: Beam induced charges cancel with traveling charge from port 1
 Signal depends on direction  can distinguish between counter-propagation beams
t=0: Beam induced charges at port 1:
 half to R1, half toward port 2
t=l/c: Beam induced charges at port 2:
 half to R2, but due to different sign,
it cancels with the signal from port 1
 half signal reflected
t=2∙l/c: reflected signal reaches port 1
If beam repetition time equals 2∙l/c: reflected preceding port 2 signal cancels the new one:
 no net signal at port 1

## Slide 73
Stripline BPM: Transfer Impedance
The signal from port 1 and the reflection from port 2 can cancel  minima in Zt .
For short bunches Ibeam(t)  Ne ∙ (t):
Zt show maximum at l=c/4f=λ/4 i.e. ‘quarter wave coupler’ for bunch train
 l has to be matched to vbeam
No signal for l=c/2f=λ/2 i.e. destructive interference with subsequent bunch
Around maximum of |Zt|: phase shift φ=0 i.e. direct image of bunch
fcenter=1/4 ∙ c/l ∙ (2n-1). For first lope: flow=1/2∙fcenter, fhigh=3/2 ∙ fcenter i.e. bandwidth 1/2∙fcenter
Precise matching at feed-through required t o preserve 50  matching.
Stripline length l=30 cm, α=100

## Slide 74
Stripline BPM: Transfer Impedance
The signal from port 1 and the reflection from port 2 can cancel  minima in Zt .
For bunches of length :
Stripline length l=30 cm, α=100
Zt(ω) decreases for higher frequencies
If total bunch is too long ±3σt > l destructive interference leads to signal damping
Cure: length of stripline has to be matched to bunch length
Further advantage: Linear phase propagation  good for coupled bunch feedback

## Slide 75
Comparison: Stripline and Button BPM (simplified)
 | Stripline | Button
Idea | traveling wave | electro-static
Requirement | Careful Zstrip = 50  matching | 
Signal quality | Less deformation of bunch signal | Deformation by finite size and capacitance
Bandwidth | Broadband, but minima | Highpass, but fcut < 1 GHz
Signal strength | Large Large longitudinal and transverse coverage possible | Small Size <3cm, to prevent signal deformation
Mechanics | Complex | Simple
Installation | Inside quadrupole possible improving accuracy | Compact insertion
Directivity | YES | No
From . S. Vilkins, D. Nölle (DESY)
FLASH BPM inside quadrupole
 Tune etc.
 Beam measurements
electronics

## Slide 76
button
Beam envelope
Image envelope
Estimation of finite Beam Size Effect for Button BPM
Ideal 2-dim model:
Due to the non-linearity, the beam size enters in the position reading.
Finite beam size:
Calculation of signal response
at different location
‘Averaging’ of image position
 Cannot be corrected !
Remark: For most LINACs: Linearity is less important, because beam has to be centered
Position correction as feed-forward for next macro-pulse.
Image center
beam center

## Slide 77
Button BPM at Synchrotron Light Sources
Due to synchrotron radiation, the button insulation might be destroyed
buttons only in vertical plane possible  increased non-linearity
PEP-realization: N. Kurita et al., PAC 1995
HERA-e
realization

## Slide 78
Close Orbit Feedback: Results
Orbit feedback:
Example: 12 beam positions at GSI-SIS during ramping from 8.6 to 500 MeV/u for Ar18+
Procedure:
1. Position from all 12 BPMs
2. Calculation of corrector setting on fast (FPGA-based) electronics
3. Submission to corrector magnets
4. New position measurement
 regulation time down to 10 ms
feedback off
time [ms]
400
300
100
0
position [mm]
20
10
0
-10
200
feedback on
time [ms]
0
100
300
400
position [mm]
20
10
0
-10
200
Role of thumb:
Movement related to tune i.e. ‘natural oscillations by periodic focusing’
To determine the ‘sine-like’ oscillation 4 BPMs per oscillation are required
 4 BPMs per tune value (but detailed investigation required to determine the # of BPMs)
position x(t) at 12 BPMs
 conclusion

## Slide 79
β-Function Measurement from Bunch-by-Bunch BPM Data
Excitation of coherent betatron oscillations:
 Time-dependent position reading results the phase advance between BPMs
The phase advance is:
 - function from
Remark: Determination of -function with 3 BPMs:
See e.g.: R. Tomas et al., Phys. Rev. Acc. Beams 20, 054801 (2017)
A. Wegscheider et al., Phys. Rev. Acc. Beams 20, 111002 (2017)
Conclusion
See lecture ‘Imperfections and Corrections’ by Volker Ziemann

## Slide 80
‘Beta-beating’ from Bunch-by-Bunch BPM Data
Result concerning ‘beta-beating’:
Model doesn’t fit reality completely
e.g. caused by misalignments
Corrections executed
Increase of the luminosity
Remark:
Measurement accuracy depends on
BPM accuracy
Numerical evaluation method
Example: ‘Beta-beating’ at BPM with measured meas & calculated model
for each BPM at BNL for RHIC (proton-proton or ions circular collider with 3.8 km length)
From X. Shen et al.,
Phys. Rev. Acc. Beams 16, 111001 (2013)
See lecture ‘Imperfections and Corrections’ by Volker Ziemann
Conclusion

## Slide 81
Single bunch position averaged over 1000 bunches  closed orbit with ms time steps.
It differs from ideal orbit by misalignments of the beam or components.
Example: GSI-synchrotron at two BPM locations, 1000 turn average during acceleration:
Closed orbit:
Beam position
averaged over many turns
(i.e. betatron oscillations).
The result is the basic tool
for alignment & stabilization
horizontal position at 2 BPMs
vertical position at 2 BPMs
time t = 0 ... 600 ms
0
5
-5
-10
-5
0
2
4
-2
time t = 0 ... 600 ms
0
6
-6
-12
position [mm]
position [mm]
position [mm]
along synchrotron
time t = 0 ... 600 ms
signal strength
horizontal
vertical
Close Orbit Measurement with BPMs

## Slide 82
Example for Fast Current Transformer
For bunch beams e.g. during accel. in a synchrotron
typical bandwidth of 2 kHz < f < 1 GHz
 10 ns < tbunch < 1 μs is well suited
Example: GSI Fast Current Transformer FCT:
From
Company Bergoz
 200 mm
FCT
injection
extraction
synchrotron
Numerous application e.g.:
Transmission optimization
Bunch shape measurement
Input for synchronization
of ‘beam phase’
Inner / outer radius | 70 / 90 mm
Permeability | µr  105 for f < 100 kHz µr  1/f above
Windings | 10
Sensitivity | 4 V/A for R = 50 
Droop time droop = L/R | 0.2 ms
Rise time rise= | 1 ns
Bandwidth | 2 kHz ... 500 MHz
Inner / outer radius | 70 / 90 mm
Permeability | µr  105 for f < 100 kHz µr  1/f above
Windings | 10
Sensitivity | 4 V/A for R = 50 
Droop time droop = L/R | 0.2 ms
 | 1 ns
Bandwidth | 2 kHz ... 500 MHz

## Slide 83
However: Not only noise contributes but additionally beam movement by betatron oscillation
 broadband processing i.e. turn-by-turn readout for tune determination.
Comparison: Filtered Signal ↔ single Turn
Example: GSI Synchr.: U73+, Einj = 11.5 MeV/u Eout = 250 MeV/u within 0.5 s, 109 ions
Position resolution < 30 μm (BPM diameter d=180 mm)
average over 1000 turns corresponding to 1 ms
or 1 kHz bandwidth
Turn-by-turn data have much larger variation
1000 turn average for closed orbit
Variation < 10 μm (sufficient for application)
Single turn e.g. for tune
Variation  150 μm

## Slide 84
Fast Current Transformer FCT (also called Passive Transformer)
Simplified electrical circuit of a passively loaded transformer:
beam
A voltages is measured: U = R · Isec = R /N · Ibeam ≡ S · Ibeam
with S sensitivity [V/A] to determine beam current Ibeam
equivalent to transfer function or transfer impedance Z
Equivalent circuit for analysis of sensitivity and bandwidth
(disregarding the loss resistivity RL)

## Slide 85
2π flow
=R/L
2πfhigh
=1/RCS
Response of the Passive Transformer: Rise and Droop Time
beam current
signal
baseline
Baseline: Ubase  1 - exp(-t/droop)
positive & negative areas are equal
Bunch train:
time
time
time
time
beam current
output volt.
test
pulse
beam bunch
Time domain description:
Droop time: τdroop= 1/( 2πflow ) = L/R
Rise time: τrise = 1/( 2πfhigh ) = RCS (ideal without cables)
Rise time: τrise = 1/(2π fhigh ) (with cables)
RL: loss resistivity, R: for measuring.

## Slide 86
Basic Idea of Beam Loss Monitors
Basic idea for Beam Loss Monitors BLM:
A loss beam particle must collide with the vacuum chamber or other insertions
Interaction leads to some shower particle:
e−, , protons, neutrons, excited nuclei, fragmented nuclei
Detection of these secondaries by an appropriate detector outside of beam pipe
Relative cheap detector installed at many locations
Remark: Due to grazing angle a thin vacuum chamber might be a ‘thick target’
BLM detector
front-end
electronics
digitalization
& fast analysis
interlock
display
lost beam particle
Secondary products:
Electromagnetic or hadronic shower products
Charged particles
Neutrons or 
vacuum pipe
beam

## Slide 87
Scintillators as Beam Loss Monitors
Plastics or liquids are used:
Detection of charged particles
by electronic stopping
Detection of neutrons
by elastic collisions n on p in plastics
and fast p electronic stopping.
Scintillator + photo-multiplier:
counting (large PMT amplification)
or analog voltage ADC (low PMT amplification)
Radiation hardness:
plastics 1 Mrad = 104 Gy
liquid 10 Mrad = 105 Gy
Scintillator
2x2x5 cm3
Photo-multiplier
inside
HV base

## Slide 88
Ionization Chamber as Beam Loss Monitors
Energy loss of charged particles in gases → electron-ion pairs → current meas.
W is average energy for creation
for one e- -ion pair:
Gas | Ionization Pot. [eV] | W-Value [eV]
Ar | 15.7 | 26.4
N2 | 15.5 | 34.8
O2 | 12.5 | 30.8
Air |  | 33.8
shower particle
Sealed tube Filled with Ar or N2 gas:
Creation of Ar+-e− pairs,
average energy W = 32 eV/pair
measurement of this current
Slow time response
due to  10 μs drift time of Ar+.
Per definition: Direct measurement of dose !
15 cm
Notes: 88

## Slide 89
Ionization Chamber as BLM: TEVATRON and CERN Type
15 cm
38 cm
Parameter | TEVATRON, RHIC | CERN type
Size | 15cm,  6 cm | 50 cm,  9 cm
Gas | Ar at 1.1 bar | N2 at 1.1 bar
# of electrodes | 3 | 61
Voltage | 1000 V | 1500 V
Reaction time | 3 µs | 0.3 µs
TEVATRON, RHIC type
CERN type
4000 BLMs at LHC each  6m

## Slide 90
Cherenkov Light Detectors as Beam Loss Monitors
Advantage:
Detection of fast electrons only
not sensitive to  & synch. photons
No saturation effects
Prompt light emission
Usage: Mainly at FELs for
short and intense pulses
Cherenkov detectors:
Passage of a charged particle v faster than propagation of light v > cmedium = c /n
Technical: Quartz rod n=1.5 & photomultiplier
Example: Korean XFEL behind undulator
Cherenkov light emission:
For v > cmedium = c /n
light wave-front like a wake
broadband light emission
H. Yang, D.C. Shin, FEL Conf. 2017
beam
11mm
v t

light propagation
beam
120mm

## Slide 91
Comparison of different Types of BLMs
Different detectors are sensitive to various physical processes
very different count rate, but basically proportional to each other
Typical choice of the detector type:
Ionization Chamber:
Advantage:
- Measurement of absolute dose
Disadvantage:
- Low signal (low , eff, no neutron detection),
- Sometimes slow, ion drift time 10 ... 100 µs
 Often used at proton accelerators
Scintillator
2x2x5 cm3
Photo-multiplier
inside
HV base
Scintillator, Cherenkov detector:
Advantage:
- Fast current reading or particle counting
- Can be fabricated in any shape, cheap
Disadvantage:
- Need calibration in many cases
- Might suffer from radiation
 Often used at electron accelerators
38 cm
conclusion