## Slide 1: A Longitudinal Density Monitor for the LHC
ADAM JEFF
DITANET International Conference
November 2011, Sevilla

## Slide 2
Contents
Motivation
Description of the system
Signal Correction methods
Results & comparison with existing instruments
Scheme for even higher dynamic range
Longitudinally profile the entire LHC beam with 50ps time resolution and high dynamic range.
Aim

## Slide 3
Satellite bunches are important…
Current & luminosity normalisation
Machine protection (e.g. satellites in kicker window)
Experiment background
Satellite - main collisions used in the low-luminosity IPs
 | Protons | Lead Ions
Maximum beam energy (2010-11) | 3.5 TeV | 3.5 x 82 TeV / ion
Revolution period | 89 μs | 
RF period | 2.5 ns | 
Minimum bunch spacing | 25 ns | 100 ns
Maximum number of bunches | 2808 | 592
Bunch population (ultimate) | 1.7 x 1011 | 8.2 x 109 charges [108 ions]
9 out of 10 buckets should be empty…
…but they’re not.
What is it for?

## Slide 4
APD
TDC
synchrotron light
LHC turn clock
Electrical pulse
Arrival time
Single photon counting is used to achieve a high dynamic range:
> 105 with 15 minutes integration
Single photon Counting

## Slide 5
Synchrotron light from protons…
… and even lead ions!
Dedicated undulator for diagnostics gives visible light at injection
Undulator
Dipole

## Slide 6
Optical table
Proton beam
Synchrotron light
Dipole
Undulator
Fast camera
LDM
Slow camera
Focus trombone
AGM
Optical layout
TOP VIEW
LDM receives 7% of collected light

## Slide 7
Instrument response

## Slide 8
Deadtime correction is essential
  to restore the true bunch shape
  to measure satellites following the main bunch
Deadtime Correction
Lab test with pulsed LED

## Slide 9
Afterpulses occur due to charge carriers trapped in the silicon
Start at end of deadtime but continue for many μs
Fit by sum of multiple exponentials
Afterpulses are additive
Afterpulse Correction

## Slide 10
Results

## Slide 11
Results

## Slide 12
Main bunch
Satellites
Afterpulsing
Deadtime
Ghost bunches

## Slide 13
Corrected for deadtime

## Slide 14
Corrected for deadtime and afterpulsing
Corrected for deadtime

## Slide 15
Bunch lengths

## Slide 16
Satellite Fraction
Beam halo events
With thanks to Beate Heinemann for MBTS the plot

## Slide 17
High-DR scheme
Free-running APD gets about 10-6 of the light on the extraction mirror
Gated APD gets about 1% of the light on the extraction mirror

## Slide 18
Optical gating
Electro-optic deflector maps longitudinal profile onto transverse plane
First lab test:
Extinction ratio 20:1
Limited by diffraction & driver

## Slide 19
Thanks for your attention

## Slide 21
Spare slides

## Slide 22
Geiger-mode Avalanche photodiode converts photon to electrical pulse
Time to Digital converter records pulse arrival time
cPCI computer makes histogram and corrects for APD dead-time and afterpulsing
Synchrotron light from undulator and dipole
Schematic

## Slide 23
Bunch currents

## Slide 24: Linearity / offset wrt BCT
Fit for nominal bunches only
Statistical errors only

## Slide 25: Emittance dependence
MD with one group of low-emittance bunches and one group with large emittance

## Slide 26: Emittance dependence
Active area of APD is only 50 microns
Samples only part of the beam spot
Slope depends on steering of the LDM
Can be corrected with an optical diffuser