## Slide 1: SM18 Vertical/Horizontal Test Setup
LHC Crab Cavity Engineering Meeting
Session: Cavity Cryostat Testing & Integration in SM18
A. Benoit, O. Brunner, J. Chambrillon, E. Jensen, K. Liao , P. Maesen
BE/RF at CERN

## Slide 2: OUTLINE
CAVITY PREPARATION AND ASSEMBLY AT CERN
- EXISTING ASSEMBLY AREA AND CLEAN ROOMS
- CLEANROOM LAYOUT (Existing facility in SM18)
- CLEANROOM LAYOUT (Updated)
- HIGH PRESSURE RINSING (HPR) AND ULTRA PURE WATER PRODUCTION
RF TEST FACILITES
- NEW RF CRYO INFRASTRUCTURE AT SM18
- RF POWER ZONE
- AVAILABILITY FOR CRAB CAVITY
- RF CONTROL ROOM
DIAGNOSTIC SYSTEMS
- OSCILLATING SUPERLEAK TRANSDUCER AND TEMPERATURE MAPING
- OPTICAL INSPECTION
CONCLUSIONS

## Slide 3: CAVITY PREPARATION AND ASSEMBLY AT CERN

## Slide 4
  Building 252:
      - Low pressure ultra pure water rinsing (max 10 bars)
      - Two clean room ISO 5
    SM18:
    - Assembly area ≈ 120 m2
    - “Baldaquin”: ISO 5 clean room (for tall items such as HIE ISOLDE cavities, main couplers)
    - Two 13 meters long clean rooms (ISO 7, ISO 4)
    Well adapted for crab cavities preparation
    and assembly
EXISTING ASSEMBLY AREA AND CLEAN ROOMS

## Slide 5
CLEANROOM LAYOUT (Existing facility in SM18)
Cleanroom facility
RF facilities
Magnet
Cryogenic
SM18, plan

## Slide 6
CLEANROOM LAYOUT (Actual)
ISO 7
ISO 4
Air Box
Built in 1991.
Goals :
  Increase cleanliness.
  Create a rinsing room.
  Get ride of the Air Box to gain space.
ISO 14644-1 | FS 209
4 | 10
5 | 100
7 | 10 000

## Slide 7: CLEANROOM LAYOUT (Updated)
Extension of the rinsing room
Creation of a preparing room outside of the facility
ISO 5
ISO 4
ISO 4
Optical
Inspection
ISO 5

## Slide 8
Refurbishment of the filtering system with Fan Filter Unit (FFU).
Rinsing room extended to 4.5m high to held the HPR.
Preparing room based on softwall design.
CLEANROOM LAYOUT (Updated)

## Slide 9
HIGH PRESSURE RINSING (HPR)
HPR from SPEC (US)
HP pump (140 bar, 1.1 m3/h)
@ 3 stages to prevent vibration
Rinsing cabinet with top wand
design (needed in SM18)
Cabinet should shall have a pass
trough design => load/unload from
different class cleanrooms
Drying of the cavity by blowing hot N2
Rinsing nozzle to be designed

## Slide 10
ULTRA PURE WATER PRODUCTION
Pre-design of the installation
Production capacity of 1 m3/h
UV lamp (185 nm) to prevent bacteria growing and
decrease Total Organic Carbone (TOC):
  CXHYOZ -> CO2 + H20 -> H2CO3 -> H+ + HCO3-
HCO3- absorbed by
polishing resin.
Need to cool down
the recirculation
water due to the
temperature in
SM18

## Slide 11
RF TEST FACILITES

## Slide 12: NEW RF CRYO INFRASTRUCTURE AT SM18
The new cryo line is under replacement (hardware and process)
available in march 2013
Vertical test stands
2 cryostats for measurement at 4.5K
- 1 long (4m) for spare LHC cavities
- 1 short (2.4m) for HIE-ISOLDE
2 new cryostats for measurement at 2K
- 1 Long (h= 4m) for Crab (400 MHz) and SPL (704 MHz) cavities
- 1 short cryostat (2.5m)
New cryostat for 2K measurement
- Welded connection
- Integrated pumping
- Active screening
- Earth magnetic field compensation

## Slide 13: NEW RF CRYO INFRASTRUCTURE AT SM18
Horizontal test stands
Bunker A:
      - Linac 4 RF structure tests: -> end 2012
        Makes use of the existing 352 MHz power system & rad. safe bunker
      - SPL: 2013
        704 MHz RF power system will replace the 352 MHz system
4.5K and 2K measurement
Hard cryogenic connection
Bunker B:
- LHC
        Spare cryo modules tests, development of LLRF electronics & control tool
        - HIE-ISOLDE cryo modules (end of 2013)
      4.5K and 2K measurement
      Flexible cryogenic connection
CCDTL
Dismantling old lines

## Slide 14
    Very crowded area:
        Cannot be significantly extended (unfortunately)
        Installation of additional power source + associated power distribution system
        to be studied in details…but look rather complicated!
RF POWER ZONE
    Equipped with:
      Demineralized cooling water
      100 kV, 40 A power converter
      400 MHz, 300 kW CW LHC power system
      352 MHz, 1.2 MW pulsed power system (1.6ms/2Hz)
        - Later a 704 MHz pulsed power system will be installed (1.6 ms / 50 Hz)

## Slide 15: AVAILABILITY FOR CRAB CAVITY
Cavity tests:
Possible (we did a first test in November)
Need to improve the cavity support on the insert (under study for the next test)
Cryo-module tests:
    Power source is ok (400 MHz, 300 KW) , difficulty for 400 MHz at 2K
    Bunker B: need of adaption for the cryo connection for 2K (flexible to hard)
    Bunker A: need to install a new power distribution
Complicated (no space), we have to study this in detail, and manage with the other projects
One additional cost is to be expected

## Slide 16
    Common control room for:
      Each vertical cryostats
      Each RF bunkers
        - Complete LHC LLRF system
RF CONTROL ROOM

## Slide 17
RF CONTROL RACK FOR VERTICAL TEST STAND

## Slide 18
TEST CONTROL PANEL PROGRAM
Improvement: automatic control of the PLL, τ measurement
Labview Program by A. Benoit BE/RF
18

## Slide 19: DIAGNOSTIC SYSTEMS

## Slide 20
OSCILLATING SUPERLEAK TRANSDUCER AND TEMPERATURE MAPING
OSTs installed around the 4 rods
Localized hot spot (default) with 2nd sound propagation in He superfluid method
Ref: K. Liao BE/RF
Temperature mapping for elliptical cavity : Using Allen-Bradley resistors temperature sensor (K.Liao)
can be adapted for crab cavity

## Slide 21: OPTICAL INSPECTION
Camera:
3488 x 2616 pix
@ 2.5 ips
Cavity can:
  Slide
  Rotate
Camera can:
  Rotate over 180°
We have to study how can we adapt it for the crab cavity inspection (support, diameter 70 mm)
Ref: J.Chambrillon BE/RF

## Slide 22: CONCLUSIONS
      Preparation and assembly:
        - Clean room upgrade for high performance cavity at SM18
        - New HPR
      Diagnostic systems:
        - OST
        - Temperature mapping
        - Optical inspection
      Low power tests:
        - Four vertical cryostats (2 K & 4.5 K)
        - New RF cryo line
      High power tests:
        - Two bunkers
          - Bunker A (4.5K): LHC, HIE ISOLDE
          - Bunker B (2 K): Linac4, SPL
        - High power area:
          - 352 & 400 MHz high power sources available
Difficult @ 400 MHz or/and 2 K need adaptation for the cryo connections or need to install a new power distribution line over cost