First Measurements with the Tracker Laser Alignment System
The Laser Alignment System of the CMS Tracker
LHC Alignment Workshop, June 2007
R. Adolphi , M. Henke, A. Ostapchouk,
M. Thomas, S. Schael, B.Wittmer

---
Outline
System Overview
Components
  - Optics
  - Electronics
  - Software
Measurements
  - Endcap integration
  - Slice test
Summary

---
Motivation
The HLT requires online track reconstruction which uses powerful pattern recognition algorithms.
Misalignments of the Tracker in the order of 100 mm start interfering with the pattern recognition.
The Laser Alignment System (LAS) is meant to detect movements and deformations of the Tracker support structure with a precision better than 100 mm.
Quick online feedback even without collisions.

---
Working Principle
Send several laser beams through the Tracker
Use the Tracker microstrip sensors to detect the beams
From changes of the measured laser spots infer movements of the modules
Variant 1 (endcap alignment):
Beam passes through several layers of sensors (Silicon is semitransparent to infrared light)
Variant 2 (subdetector alignment):
Beam is split and sent to several sensors (Mirrors have to be mounted on a rigid structure)

---
System Overview
Laser Sources in service cavern
TTC System fires Lasers
Front End Driver (FED) tags laser events
PC Farm recognizes laser events and sends them to alignment process
Service Cavern USC 5
Experimental Cavern UXC 5
CMS
Tracker
Laser Barrack
FED
TTC
PC Farm

---
Beam Layout
2 x 8 beams for internal alignment of each Endcap
8 beams to align subdetectors (Inner/Outer Barrel and Endcaps)

---
Modes of Operation
Dedicated Alignment Runs
  - Tracker standalone
  - Raw data mode
  - Optimized Timing
Laser Evemts during Physics Runs
  - Use orbit gaps (3ms)
  - Zero suppressed mode
  - Timing optimized for particles
3 ms orbit gap

---
Endcap Alignment Silicon Sensors
Optical Components1111111111
Silicon sensors polished on both sides
Hole in backplane metallization
Antireflective coating on backside
No antireflective coating on strips due to effects on interstrip capacitance
transmission
reflection
Sensor Backplane
AR coating
Tranmission improves
Reduced multiple reflections
-> Reduced interference
-> reduced distortion of profiles
Sensor Backplane
Sensor Backplane
Sensor Backplane
transmission
reflection
transmission
reflection
transmission
reflection

---
Beam Splitters
The beam splitters generate a pair of back-to-back beams:
misaligned fibres lead to a kink between the beams
All beam splitters have been measured in the lab for collinearity (with final fibre connected).
This data is used to correct the LAS measurements.

---
Alignment Tubes
Beamsplitters mounted on one end of aluminum tube
6 semitransparent mirrors per tube (3 for TIB, 3 for TOB)
Mirrors are glass plates with single-sided AR-coating
5% reflection per mirror

---
Polarization in Alignment Tubes
Reflection in beam splitter and at mirrors depends on polarization
Polarization mismatch generates 'ghosts'
Cured with l/2 plate

---
Alignment Tube Calibration
Nominal:
Beamsplitter beams collinear, Alignment Tube beams in one plane
Measure deviations previous to installation
Only angles that generate displacements in F are relevant
Use this values later to correct data
x
y
z

---
Electronics Overview
TTCB 'Test Enable' commands are sent
The LAS Trigger Board fires the lasers
A normal trigger follows with a fixed time delay
FED also detects the TTCB command and tags the event header
Laser Pulser
CU
LDD
TTC
Laser Pulser
CU
LDD
Laser Pulser
CU
LDD
Laser Pulser
CU
VME
Trigger Board
OR
LDD
LDD
LDD
LDD
DAQ
Tracker
Modules
TTC

---
Trigger and Laser Board
Trigger Board
  - Detects 'Test Enable' Pattern
  - Generates Trigger for Laser Boards
  - Programmable coarse delay (25 ns steps)
Laser Board
  - Pulses 8 Laser diodes
  - Intensity, width and fine delay programmable
  - Sequence of 100 programmable pulses

---
Increasing S/N with Statistics
The silicon detectors have a limited dynamic range
The intensity has to be adpted for a given layer
Accumulating laser events improves S/N
Single pulse
(S/N = 1.5)
500 pulses
averaged
Effective S/N

---
Software
Detects which modules have optimal intensity
Acquires the required amount of pulses
Determines Laser Spot Positions
Computes Tracker Deformations
Feedback to CMS Geometry DB
First packages have been incorporated in official software framework (CMSSW)

---
Measurements during TEC Integration
The LAS has been operated in all sectors during TEC integration
Alignment parameters are calculated from all laser profiles
Endcaps were in vertical position
Beams at Ring 4 and Ring 6
of each back petal
Ring 6
Ring 4
Beamsplitters on disc 6

---
First Sector
Beam Profiles for Disc 2 and Disc 6
Ring 4
Ring 6
Beam-
Splitters
Laser pulses
RMS of Residuals in first Sector:
<70μm
Reproduceability: 10μm

---
Beam Profiles
Intensity adapted for each disc (20% transmission per layer)
500 Events per intensity to increase S/N
Baseline correction has to be performed
beam
splitter
Ring 4
Disc 9
beam
splitter
Disc 1
Ring 6

---
Non-Gaussian Profiles
Distortion of Beam Profiles due to
Non-perfect beam optics (depends on distance to beam splitter)
-> asymmetric non-Gaussian profiles
Diffraction at microstrips (depends on amount of layers passed and strip pitch)
-> side maxima
Interference inside silicon bulk due to multiple reflections
-> asymmetric non-Gaussian profiles
Position fitting algorithms have to cope with this
Laser beam profile
after 4 layers of sensors
Laser beam profile
next to the beam splitter

---
TEC Alignment
TEC Alignment
2 x 8 Beams
(at ring 4 and ring 6)
9 Discs
Consider translations in x and y and rotations around z
Beams are straight, but can move
Collective movements of discs cannot be distinguished from collective movements of beams!
y
x

---
TEC Alignment Coordinates
TEC Alignment Coordinates
Separate collective movements from individual disc movements
Example for disc displacements along x-axis
Overall TEC movement x0
Overall TEC skew xt
k = 1..9
z

---
TEC Alignment Parameters
TEC Alignment Parameters
φk: Individual disc rotation around z-axis.
xk: Individual x-displacement of each disc.
yk: Individual y-displacement of each disc.
θai: Individual beam displacement in φ on disc 1.
θbi: Individual beam displacement in φ between disc 1 and disc 9.
φ0: Rotation of all discs with respect to the beams.
x0: Displacement in x of all discs with respect to the beams.
y0: Displacement in y of all discs with respect to the beams.
φt: Collective torsion of discs with respect to the beams.
xt: Collective shearing in x-direction of discs with respect to the beams.
yt: Collective shearing in y-direction of discs with respect to the beams.
Only xk, yk, k are used for geometry correction.
Ther is one set of parameters for each TEC.

---
Internal TEC+ Alignment
LAS, Cosmic measurements and metrology survey agree within 60 mm for translations and 80 mrad for rotations
Translation x
Translation y
10-3
Rotation

---
What precision do we achieve?
disc misalignment
Internal TEC Alignment
  - LAS, Cosmic measurements and metrology survey agree within
    60 mm for translations and 80 mrad for rotations => these are upper limits for the LAS precision
  - If straight lines are fitted through the laser spots the residuals have an RMS of 110-120 mm => upper limit for laser spot measurement Dm (contains module mounting precision, disc misalignment and intrinsic measurement precision)
TIB/TOB/TEC alignment
  - expected to be slightly better (Nr. Of measurements is 50% larger)
MC study of measurement precision
8 laser beams in TEC at R0

---
TIB/TOB/TEC Alignment
TIB/TOB/TEC Alignment
8 Alignment tubes
2 x 6 spots per tube on TIB/TOB modules
5 layers per tube in each TEC at ring 4
z
TOB
TIB
TEC

---
TIB/TOB/TEC Coordinates
TIB/TOB/TEC Coordinates
Again collective Tracker movements are indistinguishable from collective laser beam movements
The parameters are defined such that the individual subdetector movements are separated from the tracker movements
x0_TEC+
x0_TEC+ + x0_TOB+ + x0_TIB+ + x0_TIB- + x0_TOB- + x0_TEC- = 0
Similar equations for y,  and shearing of subdetectors
TOB+
TOB-
TEC+
TEC-
TIB+
TIB-

---
Tracker Slice Test
At the Tracker Integration Facility (TIF) a fraction of the tracker is being operated
2 TEC+ sectors (4 LAS beams)
A segment of TIB+/TOB+ covered by 3 Alignment Tubes
A segment of TIB-/TOB- covered by 1 alignment tube
The tracker is being cooled down to -20 oC
(only +side is being read out)
Intermediate cooling steps at +10 oC, 0 oC, -10 oC

---
CMS Tracker
Laser Alignment System
First Data taken with Alignment Tubes
TOB+
TOB-
TIB-
TIB+

---
Data sets taken so far
TEC Integration
  - Each sector was measured with both beams
  - Internal Alignment of TEC+ and TEC-
    (Endcaps were in vertical position)
Slice test warm
  - Internal alignment of 2 TEC+ sectors
  - 3 Alignment tubes for TIB/TOB +
  - 2 Alignment tubes for TIB/TOB -
  - 2 Alignment tubes for TIB/TOB/TEC +
Slice test 10 oC, 0 oC, -10 oC
  - 3 (2) Alignment tubes for TIB/TOB/(TEC) +

---
Alignment Tube Measurement
PRELIMINARY
TIB/TOB/TEC +
Not corrected for beamsplitter kink
Not corrected for AT mirror tilts
Define alignment parameters for slice test configuration and compare with survey

---
First Cooling Step
Data taken with 3 alignment tubes in TIB/TOB and 2 in TEC
Comparison with warm measurement
Observed differences
RMS: 18 mm
Max: 50 mm
No movement observed within measurement precision

---
Cooling down to -10 oC
PRELIMINARY
Movements of the laser beams up to 400 mm are observed when cooling down to -10 oC
Disentangle movements of the beams from real tracker movements
Another cooling step down to
-20 oC is foreseen
15 oC -> 10 oC
15 oC -> 0 oC
15 oC -> -10 oC

---
 Further Plans for the Slice Test
Take LAS at -20 oC
Take Alignment tube data for TEC-
Test and commission final laser pulsers
Study and optimize the timing (and maybe the APV settings)
Determine if alignment parameters improve the tracking
Implement online analysis of laser events (Jan Olzem)
Run the LAS in parallel during physics runs
Generate online feedback for CMS geometry

---
Conclusion
The LAS has been operated in a stable way during Endcap integration and at the slice test
Endcaps have been aligned in vertical position
Achieved precision better than 100 mm
Movements of laser signals when cooling down are significant -> disentangle beam movements from tracker movements
Installation at pit foreseen by end of summer