## Slide 1: The Resistive Cylindrical Chamber, a new detector based on the generalization of the RPC detectors to the quasi-planar field
G. Aielli on behalf of Rome Tor Vergata group

## Slide 2
The Resistive Cylindrical Chamber
Gas pressurization :
Increase the gas target density, with a consequent increase in intrinsic efficiency
  → MRPC time response with thin single gap configuration
  → light eco-friendly CO2 based gas mixtures
Use the detector in hostile environments
The electric field gradient, depending on the polarization allows to
Contribute to the gas discharge quenching
  → new eco-friendly gas components
Increase the charge collection efficiency enhancing the multiplication in the initial part of the gas gap
Study the dependencies and optimize the time resolution
Double gap:
Tracking capability
Improvement in time resolution and efficiency
The RCC detector is a device consisting of two concentric cylinders of resistive material. The detector's stratigraphy is like that of an RPC, but the cylindrical geometry introduces new control parameters on the detector's response, as well as extending its use to hostile environments thanks to the high strength mechanical structure

## Slide 3
[Simulation of the avalanche creation in resitive circular chambers
Oliver Kortner, XVI Workshop on Resistive Plate Chambers and Related Detectors]
Simulation: electric field gradient and primary ionization
Chart: Position (mm)
| 0 | 9.9999999999997868E-3 | 1.9999999999999574E-2 | 3.0000000000000249E-2 | 4.0000000000000036E-2 | 4.9999999999999822E-2 | 5.9999999999999609E-2 | 6.9999999999999396E-2 | 8.0000000000000071E-2 | 8.9999999999999858E-2 | 9.9999999999999645E-2 | 0.10999999999999943 | 0.12000000000000011 | 0.12999999999999989 | 0.13999999999999968 | 0.14999999999999947 | 0.16000000000000014 | 0.16999999999999993 | 0.17999999999999972 | 0.1899999999999995 | 0.20000000000000018 | 0.20999999999999996 | 0.21999999999999975 | 0.22999999999999954 | 0.24000000000000021 | 0.24999999999999023 | 0.25999999999999002 | 0.2699999999999898 | 0.27999999999998959 | 0.28999999999999027 | 0.29999999999999005
D1 | 10873.1352 | 10859.0326 | 10844.9665 | 10830.9368 | 10816.9433 | 10802.986 | 10789.0646 | 10775.1791 | 10761.3292 | 10747.5149 | 10733.7361 | 10719.9925 | 10706.284 | 10692.6106 | 10678.9721 | 10665.3683 | 10651.7991 | 10638.2645 | 10624.7641 | 10611.298 | 10597.866 | 10584.4679 | 10571.1037 | 10557.7732 | 10544.4762 | 10531.2127 | 10517.9826 | 10504.7856 | 10491.6217 | 10478.4908 | 10465.3927
D2 | 10849.5356 | 10837.0792 | 10824.6513 | 10812.252 | 10799.881 | 10787.5383 | 10775.2237 | 10762.9373 | 10750.6788 | 10738.4482 | 10726.2454 | 10714.0703 | 10701.9229 | 10689.8029 | 10677.7104 | 10665.6452 | 10653.6072 | 10641.5964 | 10629.6126 | 10617.6558 | 10605.7258 | 10593.8226 | 10581.9462 | 10570.0963 | 10558.2729 | 10546.476 | 10534.7053 | 10522.961 | 10511.2427 | 10499.5506 | 10487.8844
D3 | 11093.4808 | 11063.5793 | 11033.8384 | 11004.2571 | 10974.834 | 10945.5677 | 10916.4572 | 10887.5011 | 10858.6982 | 10830.0472 | 10801.5471 | 10773.1966 | 10744.9945 | 10716.9397 | 10689.031 | 10661.2673 | 10633.6474 | 10606.1703 | 10578.8348 | 10551.6399 | 10524.5844 | 10497.6673 | 10470.8875 | 10444.244 | 10417.7358 | 10391.3618 | 10365.121 | 10339.0123 | 10313.0349 | 10287.1877 | 10261.4698
Dext=8 mm
Dext=16 mm
Dext=18 mm
Gas-gap 0.3 mm
Gas-gap 1 mm
Primary ionization with pressurized CO2
Position (mm)
Chart: 
| 0 | 2.0000000000000018E-2 | 4.0000000000000036E-2 | 6.0000000000000053E-2 | 8.0000000000000071E-2 | 0.10000000000000009 | 0.12000000000000011 | 0.14000000000000012 | 0.16000000000000014 | 0.18000000000000016 | 0.20000000000000018 | 0.2200000000000002 | 0.24000000000000021 | 0.25999999999999979 | 0.2799999999999998 | 0.29999999999999982 | 0.31999999999999984 | 0.33999999999999986 | 0.35999999999999988 | 0.37999999999999989 | 0.39999999999999991 | 0.41999999999999993 | 0.43999999999999995 | 0.45999999999999996 | 0.48 | 0.5 | 0.52 | 0.54 | 0.56000000000000005 | 0.58000000000000007 | 0.60000000000000009 | 0.62000000000000011 | 0.64000000000000012 | 0.66000000000000014 | 0.68000000000000016 | 0.70000000000000018 | 0.7200000000000002 | 0.74000000000000021 | 0.75999999999999979 | 0.7799999999999998 | 0.79999999999999982 | 0.81999999999999984 | 0.83999999999999986 | 0.85999999999999988 | 0.87999999999999989 | 0.89999999999999991 | 0.91999999999999993 | 0.94 | 0.96 | 0.98 | 1
 | 7398.9104 | 7325.6538 | 7253.8337 | 7183.4081 | 7114.3369 | 7046.5813 | 6980.1041 | 6914.8695 | 6850.843 | 6787.9912 | 6726.2822 | 6665.685 | 6606.17 | 6547.7083 | 6490.2723 | 6433.8351 | 6378.371 | 6323.855 | 6270.263 | 6217.5718 | 6165.7587 | 6114.802 | 6064.6806 | 6015.3743 | 5966.8632 | 5919.1283 | 5872.1511 | 5825.9137 | 5780.3987 | 5735.5894 | 5691.4695 | 5648.0232 | 5605.2351 | 5563.0905 | 5521.5749 | 5480.6744 | 5440.3753 | 5400.6645 | 5361.5293 | 5322.9571 | 5284.936 | 5247.4542 | 5210.5003 | 5174.0632 | 5138.1322 | 5102.6968 | 5067.7468 | 5033.2724 | 4999.2638 | 4965.7117 | 4932.6069
Chart: 
| 0 | 2.0000000000000018E-2 | 4.0000000000000036E-2 | 6.0000000000000053E-2 | 8.0000000000000071E-2 | 0.10000000000000009 | 0.12000000000000011 | 0.14000000000000012 | 0.16000000000000014 | 0.18000000000000016 | 0.20000000000000018 | 0.2200000000000002 | 0.24000000000000021 | 0.25999999999999979 | 0.2799999999999998 | 0.29999999999999982 | 0.31999999999999984 | 0.33999999999999986 | 0.35999999999999988 | 0.37999999999999989 | 0.39999999999999991 | 0.41999999999999993 | 0.43999999999999995 | 0.45999999999999996 | 0.48 | 0.5 | 0.52 | 0.54 | 0.56000000000000005 | 0.58000000000000007 | 0.60000000000000009 | 0.62000000000000011 | 0.64000000000000012 | 0.66000000000000014 | 0.68000000000000016 | 0.70000000000000018 | 0.7200000000000002 | 0.74000000000000021 | 0.75999999999999979 | 0.7799999999999998 | 0.79999999999999982 | 0.81999999999999984 | 0.83999999999999986 | 0.85999999999999988 | 0.87999999999999989 | 0.89999999999999991 | 0.91999999999999993 | 0.94 | 0.96 | 0.98 | 1
 | -7398.9104 | -7325.6538 | -7253.8337 | -7183.4081 | -7114.3369 | -7046.5813 | -6980.1041 | -6914.8695 | -6850.843 | -6787.9912 | -6726.2822 | -6665.685 | -6606.17 | -6547.7083 | -6490.2723 | -6433.8351 | -6378.371 | -6323.855 | -6270.263 | -6217.5718 | -6165.7587 | -6114.802 | -6064.6806 | -6015.3743 | -5966.8632 | -5919.1283 | -5872.1511 | -5825.9137 | -5780.3987 | -5735.5894 | -5691.4695 | -5648.0232 | -5605.2351 | -5563.0905 | -5521.5749 | -5480.6744 | -5440.3753 | -5400.6645 | -5361.5293 | -5322.9571 | -5284.936 | -5247.4542 | -5210.5003 | -5174.0632 | -5138.1322 | -5102.6968 | -5067.7468 | -5033.2724 | -4999.2638 | -4965.7117 | -4932.6069
-
-
Lower field
Higher field
weak dependence on thin gas gaps
In the RCC detector the electric field has a cylindrical geometry and is much more approximable to a planar field as the gas-gap is negligible with respect to the radius of the internal electrode. This feature allows to design detectors with different responses according to the experimental needs:
The RCC detector has a double gap structure, which can be exploited for tracking measurements and to increase efficiency and time resolution
it is possible to reproduce the same performances of an RPC in an easily pressurized cylindrical structure;
it is possible to design a detector with a non-negligible field gradient to increase the time resolution and suppress the transition to the streamer regime
Dext=6 mm

## Slide 4
Detector description and reliminary results
(1 mm gas-gap)
Insulator
4 mm
6 mm
Gas-gap 1 mm
Shielding
Graphite
Alluminium
electorde
Bakelite electrode
OUT
Chart: HV (V)
| 6900 | 6700 | 6600 | 6500 | 6400 | 6300 | 6200 | 6100 | 6000 | 5900 | 5800 | 5600 | 5500
negative | 489 | 598 | 570 | 672 | 664 | 640 | 701 | 728 | 770 | 812 | 920 | 932 | 1232
positive | 414 | 471 | 472 | 540 | 550 | 690 | 699 | 747 | 831 | 842 | 1010
+
-
[Development of Resistive Cylindrical Chambers, R. Cardarelli, XVI Workshop on Resistive Plate Chambers and Related Detectors]
[Moderation of the avalanche gas discharge through a quasi-uniform electric field device: the Resistive Cylindrical Chamber, A. Rocchi, VCI2022]
The time resolution was measured using an RPC detector with two coupled 0.2mm gaps as a reference.
The time resolution improve as the applied voltage increases and for high field values ​​it is systematically better in the case of positive polarization, in which multiplication occurs mainly near the cathode (behavior like that of a thinner gap).
In this measure a UFSD pixel detector was used to optimize geometric acceptance at the expense of statistics .
FE-Electronics + Caen V1742
Caen V1742
The response of a 1 mm gas-gap RCC was studied with the detector described in figure. The smaller radius electrode consists of an aluminum cylinder. A significant separation is observed between the working points of the detector in the different polarities. The intrinsic efficiency is lower in the case of positive polarization as if the useful gas-gap decreases.
Gas mixture
94.7% TFE + 5% iC4H10 + 0.3% SF6

## Slide 5
-
+
-
+
Detector description and reliminary results
(0.3 mm gas-gap)
Cylindrical
Pad
SIPM+
Scintillator
Strips 7 ch
Shielding
Gas
Graphite
Bakelite
Insulator
The detector described in figure was characterized with muons of 180 GeV / c.
A small difference can be observed between the working points in different polarities. Also In this configuration the detector has a lower intrinsic efficiency in positive polarity.
The internal detector made it possible to close the angle of acceptance and discriminate the tangential traces
Gas mixture
94.7% TFE + 5% iC4H10 + 0.3% SF6
[Thanks to Alessandro Paoloni for the SIPM detector and electroinics]

## Slide 6: Study with high pressurized CO2
Average profile of normalized signals
Studies on new eco-friendly gas mixtures show that mixtures with a high concentration of CO2 lead to an improvement in the time resolution at expense of the intrinsic efficiency of the detector.
The RCC thanks to the pressurization of the gas-gap allows to increase the density of the gas recovering the loss of efficiency even for thin gas gaps
The effect of the pressure is visible from the ion signal profile:
as the pressure increases, the instantaneous current decreases as the drift speed of the ions decreases and the duration of the signal increases accordingly. The amplitude of the electronic signal is unchanged
σEco = (0.83 ± 0.03) ns
σSTD = (1.09 ± 0.07) ns
[On a new environment-friendly gas mixture for Resistive Plate Chambers,
G. Proto, XVI Workshop on Resistive Plate Chambers and Related Detectors]

## Slide 7: Conclusions
The functionality of the detector has been demonstrated with prototypes of different structures (1 mm gas-gap, 0.3 mm gas-gap)
As expected, the asymmetry in the detector response with respect to the polarity of the applied voltage increases when the thickness of the gap increases with respect to the internal radius
The effect of suppression of the avalanche in positive polarization leads to a reduction of the intrinsic efficiency but to an improvement in the time resolution, as if the useful gas-gap were thinner
This effect can also be studied to suppress the transition to the streamer regime in gas mixtures with low quenching potential
Perspectives
Systematic study of RCC detectors with different radius and gas-gaps
Study of the response with high concentration of CO2