## Slide 1: Thermal transmittance measurements of SIS & Nb3Sn coated samples at cryogenic temperatures
Marc Wenskat – on behalf SRF R&D Team
emphasis on Leon King, Anton Lorf, Cem Saribal
19.09.2024
Image generated by ChatGPT, an AI model by OpenAI.
“Draw me a picture illustrating thermal conduction in a metallic superconductor.”

## Slide 2: Thermal transmittance is crucial for stable operation
We need to dissipate RF losses / Joule heating away into the liquid He
If the heat is not dissipated away, we will have an unwanted thermal feedback increasing Rs
To achieve gradients >50 MV/m we need to study thermal transmittance K
Does overall thermal transmittance K deteriorate after coating of Nb?
  Is the thermal conductivity λ the culprit?
  The thermal interface resistance Rth between layer and substrate?
  What about the Kapitza-Resistance RK of the coated layer?
Nb
S
Rth
λS
lHe
RK
λNb
λS

## Slide 3: How can we measure K across films?
To enforce heat flow through the film, standard measurements used in SRF are not feasible
Remembered earlier works by J. Amrit and C. Antoine (Thank you!)
Amrit, J., and M. X. François. "Heat flow at the niobium-superfluid helium interface: Kapitza resistance and superconducting cavities." Journal of low temperature physics 119 (2000): 27-40.
1/K → This is what we obtain
We know A, control and measure and

## Slide 4: Measurement procedure
2K

## Slide 5: How does it look like?
Cell Ø=106mm
Sample Ø=45mm
T-sensor inside & outside
Manganin heating wire inside
Parasitically installed at cavity insert
Filling by 1m capillary with Ø=1mm to thermally decouple inside
Electrical wires go through there (heater & T-sensor)
Only glued / screwed → high chance of leaks
Notes: PTFE Sealings

## Slide 6: Systematic uncertainties?
Conduction not only through the sample!
Optimized geometry & material to minimize unwanted paths
Leakage will create overestimation of K at low
  At low , heat carried away through leaks can be substantial amount of total thermal transmissivity
  Fixed value will become less significant at higher
  Leaks will seal themselves off, as He boils in channel
X %
2-5 %
0.65 %
≈94 %

## Slide 7: Samples
All samples cut from 1 EXFEL sheet
All samples underwent coarse BCP + 3h@800°C (Baseline)
1 pair was coated both sides with
  15nm/60nm AlN/NbTiN via PEALD (as-deposited)
  +1h@900°C to achieve high Tc [1] (annealed)
1 pair was coated with 200nm Nb3Sn [2] twice (one side each) by TU Darmstadt (DCMS co-sputtering)
All measurements shown are at 2K
1 González Díaz-Palacio, I., Wenskat, M., Deyu, G. K., Hillert, W., Blick, R. H., & Zierold, R. (2023). Thermal annealing of superconducting niobium titanium nitride thin films deposited by plasma-enhanced atomic layer deposition. Journal of Applied Physics, 134(3)
2 Schäfer, N., Karabas, N., Palakkal, J. P., Petzold, S., Major, M., Pietralla, N., & Alff, L. (2020). Kinetically induced low-temperature synthesis of Nb3Sn thin films. Journal of Applied Physics, 128(13).

## Slide 8: Can we trust the results? Yes, they are reasonable
Shows expected behavior with power law dependency
Values are well in agreement with literature
  K0 = (150 – 200) mW cm-2 K-1
Measured as-fabricated (only water-jet cutting)
  K0 = (307±46) mW cm-2 K-1
  K0 = (332±31) mW cm-2 K-1
Higher K0 also in agreement with literature

## Slide 9: SIS does not reduce the thermal transmittance K
As-deposited K0 = (141±13) mW cm-2 K-1
After annealing K0 = (156±10) mW cm-2 K-1
SIS on both sides → many interfaces
Well within the baseline Nb value!

## Slide 10: Nb3Sn coating shows an unexpected behavior
Nb3Sn shows increased K0 after coating one side
Further increase after coating other side as well
It can‘t be thermal conductivity λ of Nb3Sn
What about the Kapitza-Resistance RK?
  Is RK,Nb3Sn < RK,Nb ?
  Is it the surface roughness?
    Surface area would have to increase by x2
    As-fabricated Nb had similar K0
Next steps:
  Coat second sample pair, measure K0
  Measure roughness before & after
  Compare to as-fabricated roughness

## Slide 11: Summary
Thermal transmittance K is an important property of coated cavities
Yet, it remains unknown for most material / material combinations
We started measurements to assess impact of treatments & coatings
  SIS is comparable to bare Nb, despite many interfaces
  Nb3Sn shows significant increase of K – material property or surface roughness?
Next-gen cell with soldered stainless steel capillary is currently fabricated

## Slide 12
Thanks to…
Questions?
Contact:
Marc Wenskat
MSL/DESY
E-Mail: marc.wenskat@desy.de
you for listening
the conveners for the opportunity to present this work
DESY for the measurement opportunities
BMBF for funding
Swartzwelder, J. (Writer), & Reardon, J. (Director). (1994, January 6). Homer the Vigilante (Season 5, Episode 11) [TV series episode]. In D. Mirkin, J. L. Brooks, M. Groening, & S. Simon (Executive Producers), The Simpsons. Gracie Films; Twentieth Century Fox Film Corporation.

## Slide 13: Temperature Scan
Started testing temperature scans while cooling down at QvT
High power to be close to K0
Measurement starts at 2.1K