## Slide 1
Gantry Beam Optics
M. D’ Addazio, E. Felcini, G. Frisella, A. Mereghetti, M.G. Pullia, S. Savazzi, G. Tosetti

## Slide 2: Gantry Key Features
5.75m
3.09m
16.05m
14.05m
Particles up to a rigidity of 6.6 Tm (430 MeV/u carbon ions)
14 m long, 5.75 high (beam line)
4T curved superconducting dipoles
Superconducting spool piece quadrupoles
Downstream scanning magnets

## Slide 3: Optics details
Parallel to point matching
R11=R33=0; R12 and R34 for beam size control (MF*)
Beam sizes
8mm spot (FWHM)
12mm spot (FWHM)
With const. = 1.4π (7π=5)
Particles with the same initial divergence (x’, y’) end up in the same position independently of the initial position (x, y).

## Slide 4: Optics details
MF 1.5
MF 1

## Slide 5: 4 T SC dipole Features
SIG WP2 scope
Design, construction and test of a curved SC demonstrator magnet for ion gantries
Main demonstrator magnet parameters:
Cos-θ coils
Pure dipolar field: 4 T
Bore diameter: 80 mm
Small curvature radius: 1.65 m
Angular sector: 30° (demo) – 45° (prototype)
High field ramp-rate: 0.15 T/s (demo) – 0.4 T/s (prototype)
Fixed quadrupole gradient

## Slide 6: Constraint on the space required by the beam
Every alignment and magnetic error produces a deviation of the beam orbit from the nominal one.
The space occupied by the beam (𝑅tot) defines the required Good Field Region (GFR) for the magnets.
As commonly done in designing accelerator magnets we consider that the GFR is 2/3 of the magnet aperture.
Using monitors and correctors, one can correct the orbit deviation at the isocenter, but a residual orbit deviation
(𝑅𝑐𝑜𝑟𝑟 ) is till present along the line
Bore Ø: 80 mm
The total (radial) space occupied by the beam (𝑅𝑡𝑜𝑡 ) is given by:

## Slide 7: Beam envelope

## Slide 8: Entry Sin-like and Cos-like

## Slide 9: Field errors
Absolute error (construction errors, i.e. a displacement in the length of the filament, and thus an error in the magnetic length)
  dk = 3・10-4 1/m2 for warm quadrupoles
  dk = 2.5・10-4 (1/m , 1/m2) for SC dipoles and quadrupoles
Relative error (describing errors coming from power supply)
  dkr = 3 ・10-4 1/m2 for warm quadrupole
  dkr = 5 ・10-4 (1/m , 1/m2) for SC dipoles and quadrupole

## Slide 10: Misalignment errors
Besides the deformation of the structure (WIP), that is part of separated studies, random alignment errors have been considered to take into account realistic alignment.
Lcq
Les
Lcd
SC 4T dipole
SC quadrupoles
A conservative assumption has been made assuming Gaussian distributions truncated at 3𝜎 with 0.3mm (0.3 mrad) < 𝜎 < 0.6mm (0.6mrad) for the position (angle) error.
Misalignment error in MAD-x are defined at the entrance of the element.
In case the SC quadrupoles are attached to the dipole, the error is not anymore independent: the error is generated in SC dipole and propagated to the quadrupoles

## Slide 11: Misalignment plus Field errors
1000 random misalignments (minor differences with 2000 cases)
95th percentile

## Slide 12: Space required by the beam

## Slide 13
Problems
Optics-based optimization doesn’t mean minimum space occupied
Non-realistic SC dipole representation
Human bias can reduce space of solutions
Steerers strength can not be infinite

## Slide 14
Solution

## Slide 15
SC Magnet
Entrance
Exit
Entrance
Exit
Multipolar Lens
CF Dipole
Optic axis
CF
ML1
ML2
Field Map Representation
Real Representation
Optic Representation
Multipoles representation

## Slide 16: Total Envelope optimization
STEP 3
Reiteration: Position of active elements are varied until the matching value is under
STEP 1
Active elements are matched using parallel to point matching method
STEP 2
If the matching value is under , the monitor and kicker positions are varied to minimize the beam envelope
Optimization Function
Opt. Function
STEP 2bis
Misalignment and Field gaussian errors are defined for all the elements and correct by kickers.
Storing Data
Once the beam envelope is minimized the data are stored in a matlab table

## Slide 17
Different quadrupole gradients optimization
gradient [T/m] |  | K1 [m-2]
0.1 |  | 0.015107
0.092674667 |  | 0.014
0.08 |  | 0.012085
0.05 |  | 0.007553
0 |  | 0
)
5 good solutions
K1 [m-2] | 0.014 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  | 100th perc |  | 95th perc |  |  |  |  | 
case nr | val1_5 | val1 | l0 | l1 | l2 | l_mon2 | l_corr2 | l_corr3 | r_max_1_5 | r_max_1 | r_max_1_5 | r_max_1 | PercKick_x_MF1_5 | PercKick_y_MF1_5 | PercKick_x_MF1 | PercKick_y_MF1
1 | 1.56E-15 | 6.52E-14 | 0.985893453 | 1.310892133 | 2.433214414 | 0.353185305 | 0.811469266 | 1.319900282 | 0.027395321 | 0.028791809 | 0.023753941 | 0.024251372 | 3.7 | 23.3 | 3.066666667 | 23.83333333
2 | 1.16E-11 | 5.03E-13 | 0.995893453 | 1.310892133 | 2.423214414 | 0.354205945 | 0.812261895 | 1.315558725 | 0.026867793 | 0.028283885 | 0.023290335 | 0.023773574 | 3.733333333 | 23.56666667 | 3 | 23.86666667
3 | 1.35E-12 | 2.62E-13 | 0.985893453 | 1.320892133 | 2.423214414 | 0.355526083 | 0.818081121 | 1.316909179 | 0.02928376 | 0.030574292 | 0.025543659 | 0.02599667 | 3.8 | 22.63333333 | 3 | 23.43333333
4 | 1.24E-12 | 6.22E-15 | 0.995893493 | 1.310892068 | 2.423214438 | 0.35553301 | 0.813116669 | 1.316606913 | 0.029243777 | 0.030540408 | 0.025513499 | 0.025970254 | 3.766666667 | 23 | 2.966666667 | 23.8
5 | 4.57E-12 | 5.71E-15 | 1.2971709 | 1.459722912 | 1.973106188 | 0.358059601 | 0.884171343 | 1.115239679 | 0.025084223 | 0.025385609 | 0.020977018 | 0.021488575 | 1.9 | 19.73333333 | 1.333333333 | 20.46666667
6 | 2.76E-13 | 2.15E-11 | 1.326238785 | 1.500642804 | 1.90311841 | 0.455838127 | 0.850605834 | 1.042112235 | 0.026820347 | 0.026041353 | 0.019977887 | 0.020428963 | 1.733333333 | 19.86666667 | 1.366666667 | 20.46666667
7 | 7.19E-14 | 3.23E-14 | 1.130802393 | 1.448387928 | 2.150809679 | 0.398496051 | 0.875446871 | 1.191234339 | 0.026113262 | 0.026657179 | 0.021809925 | 0.022408743 | 2.5 | 20.06666667 | 1.733333333 | 20.33333333
8 | 1.24E-12 | 6.22E-15 | 0.995893493 | 1.310892068 | 2.423214438 | 0.35553301 | 0.813116669 | 1.316606913 | 0.029243777 | 0.030540408 | 0.025513499 | 0.025970254 | 3.766666667 | 23 | 2.966666667 | 23.8
9 | 2.71E-15 | 2.28E-14 | 1.143227037 | 1.634953015 | 1.951819948 | 0.361252515 | 0.966892228 | 1.109120922 | 0.028661673 | 0.029440686 | 0.025287859 | 0.025544178 | 1.9 | 15.73333333 | 1.266666667 | 16.33333333
10 | 4.22E-15 | 2.13E-15 | 1.341457406 | 1.636490983 | 1.752051611 | 0.360561041 | 0.966571589 | 1.018604769 | 0.027348521 | 0.02855659 | 0.024306131 | 0.02473294 | 1.5 | 15.73333333 | 0.9 | 16.3
K1 [m-2] | 0 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  | 100th perc |  | 95th perc |  |  |  |  | 
case nr | val1_5 | val1 | l0 | l1 | l2 | l_mon2 | l_corr2 | l_corr3 | r_max_1_5 | r_max_1 | r_max_1_5 | r_max_1 | PercKick_x_MF1_5 | PercKick_y_MF1_5 | PercKick_x_MF1 | PercKick_y_MF1
1 | 3.74E-12 | 1.70E-10 | 0.986 | 1.301 | 2.443 | 0.388405815 | 0.807236299 | 1.322454972 | 0.029076527 | 0.030038861 | 0.024888834 | 0.025349283 | 1.8 | 26 | 1.4 | 26.43333333
2 | 1.46E-12 | 1.49E-10 | 1.244236195 | 1.470545018 | 2.015218787 | 0.357118897 | 0.890327759 | 1.135580066 | 0.027154574 | 0.027347109 | 0.022906567 | 0.023345713 | 0.866666667 | 20.8 | 0.533333333 | 21.46666667
3 | 5.07E-15 | 1.63E-10 | 1.706579825 | 1.98376723 | 1.039652945 | 0.51587944 | 1.11659297 | 0.68285577 | 0.02530555 | 0.024958366 | 0.020371337 | 0.020648411 | 0 | 11.5 | 0 | 11.93333333
4 | 4.85E-12 | 1.63E-10 | 1.723999487 | 2.140280892 | 0.977 | 1.218283031 | 0.701733488 | 0.65099999 | 0.039702985 | 0.038250269 | 0.033087129 | 0.032276453 | 0 | 18.53333333 | 0.033333333 | 16.13333333
5 | 6.81E-11 | 1.54E-13 | 1.83379856 | 2.140722688 | 0.977 | 0.450191456 | 1.19277069 | 0.651 | 0.025422002 | 0.025303711 | 0.021959564 | 0.022279032 | 0 | 9.633333333 | 0 | 9.233333333

## Slide 18
95th perc optimization
Chart: Raggio max al 95th percentile
| MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | senza multipoli | con multipoli | senza multipoli | con multipoli | senza multipoli | con multipoli | senza multipoli | con multipoli | senza multipoli | con multipoli | 0.1 T/m case 1 | 0.1 T/m case 2 | 0.1 T/m case 3 | 0 T/m case 1 | 0 T/m case 2
σ=0.6mm | 0.019944 | 0.020367 | 0.020114 | 0.020097 | 0.021978 | 0.022544 | 0.021315 | 0.021593 | 0.02445 | 0.024845 | 0.02414 | 0.024362 | 0.020477 | 0.02077 | 0.020723 | 0.020637 | 0.022095 | 0.022389 | 0.021462 | 0.021475
σ=0.5mm | 0.019438 | 0.019851 | 0.019607 | 0.019585 | 0.021702 | 0.0223 | 0.021041 | 0.021346 | 0.024199 | 0.024621 | 0.023901 | 0.024139 | 0.020247 | 0.020525 | 0.020638 | 0.020637 | 0.021861 | 0.022175 | 0.021461 | 0.021461
σ=0.4mm | 0.019136 | 0.019422 | 0.019208 | 0.019115 | 0.021452 | 0.022054 | 0.020811 | 0.021114 | 0.023985 | 0.024385 | 0.023682 | 0.023908 | 0.020049 | 0.020307 | 0.020637 | 0.020636 | 0.021666 | 0.021981 | 0.021461 | 0.021461
σ=0.3mm | 0.018906 | 0.019197 | 0.019162 | 0.019092 | 0.021225 | 0.021876 | 0.020624 | 0.020938 | 0.023796 | 0.024231 | 0.023496 | 0.023745 | 0.019843 | 0.020144 | 0.020637 | 0.020637 | 0.021489 | 0.021815 | 0.021461 | 0.021463

## Slide 19
Kick strenght analisys
Chart: Percentuale di kick > 5mrad
| x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | senza multipoli | con multipoli | senza multipoli | con multipoli | senza multipoli | con multipoli | senza multipoli | con multipoli | senza multipoli | con multipoli | 0.1 T/m case 1 | 0.1 T/m case 2 | 0.1 T/m case 3 | 0 T/m case 1 | 0 T/m case 2
σ=0.6mm | 1.65 | 20.8333 | 1.1 | 21.45 | 0.16667 | 16.85 | 0.016667 | 16.0333 | 2.45 | 20.7167 | 1.7167 | 20.8167 | 0.76667 | 19.6833 | 0.05 | 16.45 | 1.3 | 16.1167 | 0.66667 | 16.85 | 0.2 | 16.3167 | 0.016667 | 14.5667 | 0.016667 | 11.6333 | 0.016667 | 11.85 | 0 | 11.9667 | 0 | 10.1167 | 0.016667 | 9.9833 | 0 | 9.3833 | 0 | 9.9 | 0 | 8.95
σ=0.5mm | 0.53333 | 15.5833 | 0.3 | 16.1833 | 0.033333 | 13.1833 | 0 | 12.8 | 1 | 16.0333 | 0.56667 | 16.1833 | 0.13333 | 15.45 | 0.016667 | 13.4167 | 0.31667 | 12.15 | 0.13333 | 12.3333 | 0.033333 | 12.8667 | 0 | 10.95 | 0 | 8.2167 | 0 | 8.0167 | 0 | 8.7 | 0 | 7.3333 | 0 | 6.9167 | 0 | 6.1833 | 0 | 6.8167 | 0 | 6.1167
σ=0.4mm | 0.083333 | 10.2 | 0.066667 | 10.9333 | 0 | 9.1333 | 0 | 9.0667 | 0.16667 | 10.9667 | 0.083333 | 10.9333 | 0 | 11.3167 | 0 | 9.6 | 0.066667 | 8.1333 | 0.016667 | 7.6833 | 0 | 9.05 | 0 | 7.4833 | 0 | 4.8167 | 0 | 4.8 | 0 | 5.25 | 0 | 4.4167 | 0 | 3.7333 | 0 | 3.4333 | 0 | 3.8833 | 0 | 3.55
σ=0.3mm | 0 | 5.1 | 0 | 5.9667 | 0 | 4.9 | 0 | 4.9833 | 0 | 5.6667 | 0 | 5.8167 | 0 | 6.7333 | 0 | 5.5 | 0 | 4.0167 | 0 | 3.75 | 0 | 4.6667 | 0 | 3.75 | 0 | 1.8 | 0 | 1.9 | 0 | 2.1167 | 0 | 1.75 | 0 | 1.2833 | 0 | 1.15 | 0 | 1.5167 | 0 | 1.15

## Slide 20
Chart: Raggio max al 95th percentile
| MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | senza multipoli | con multipoli | senza multipoli | con multipoli | senza multipoli | con multipoli | senza multipoli | con multipoli | senza multipoli | con multipoli | 0.1 T/m case 1 | 0.1 T/m case 2 | 0.1 T/m case 3 | 0 T/m case 1 | 0 T/m case 2
σ=0.6mm | 0.019944 | 0.020367 | 0.020114 | 0.020097 | 0.021978 | 0.022544 | 0.021315 | 0.021593 | 0.02445 | 0.024845 | 0.02414 | 0.024362 | 0.020477 | 0.02077 | 0.020723 | 0.020637 | 0.022095 | 0.022389 | 0.021462 | 0.021475
σ=0.5mm | 0.019438 | 0.019851 | 0.019607 | 0.019585 | 0.021702 | 0.0223 | 0.021041 | 0.021346 | 0.024199 | 0.024621 | 0.023901 | 0.024139 | 0.020247 | 0.020525 | 0.020638 | 0.020637 | 0.021861 | 0.022175 | 0.021461 | 0.021461
σ=0.4mm | 0.019136 | 0.019422 | 0.019208 | 0.019115 | 0.021452 | 0.022054 | 0.020811 | 0.021114 | 0.023985 | 0.024385 | 0.023682 | 0.023908 | 0.020049 | 0.020307 | 0.020637 | 0.020636 | 0.021666 | 0.021981 | 0.021461 | 0.021461
σ=0.3mm | 0.018906 | 0.019197 | 0.019162 | 0.019092 | 0.021225 | 0.021876 | 0.020624 | 0.020938 | 0.023796 | 0.024231 | 0.023496 | 0.023745 | 0.019843 | 0.020144 | 0.020637 | 0.020637 | 0.021489 | 0.021815 | 0.021461 | 0.021463
Chart: Percentuale di kick > 5mrad
| x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | x | y | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | MF1.5 | MF1 | senza multipoli | con multipoli | senza multipoli | con multipoli | senza multipoli | con multipoli | senza multipoli | con multipoli | senza multipoli | con multipoli | 0.1 T/m case 1 | 0.1 T/m case 2 | 0.1 T/m case 3 | 0 T/m case 1 | 0 T/m case 2
σ=0.6mm | 1.65 | 20.8333 | 1.1 | 21.45 | 0.16667 | 16.85 | 0.016667 | 16.0333 | 2.45 | 20.7167 | 1.7167 | 20.8167 | 0.76667 | 19.6833 | 0.05 | 16.45 | 1.3 | 16.1167 | 0.66667 | 16.85 | 0.2 | 16.3167 | 0.016667 | 14.5667 | 0.016667 | 11.6333 | 0.016667 | 11.85 | 0 | 11.9667 | 0 | 10.1167 | 0.016667 | 9.9833 | 0 | 9.3833 | 0 | 9.9 | 0 | 8.95
σ=0.5mm | 0.53333 | 15.5833 | 0.3 | 16.1833 | 0.033333 | 13.1833 | 0 | 12.8 | 1 | 16.0333 | 0.56667 | 16.1833 | 0.13333 | 15.45 | 0.016667 | 13.4167 | 0.31667 | 12.15 | 0.13333 | 12.3333 | 0.033333 | 12.8667 | 0 | 10.95 | 0 | 8.2167 | 0 | 8.0167 | 0 | 8.7 | 0 | 7.3333 | 0 | 6.9167 | 0 | 6.1833 | 0 | 6.8167 | 0 | 6.1167
σ=0.4mm | 0.083333 | 10.2 | 0.066667 | 10.9333 | 0 | 9.1333 | 0 | 9.0667 | 0.16667 | 10.9667 | 0.083333 | 10.9333 | 0 | 11.3167 | 0 | 9.6 | 0.066667 | 8.1333 | 0.016667 | 7.6833 | 0 | 9.05 | 0 | 7.4833 | 0 | 4.8167 | 0 | 4.8 | 0 | 5.25 | 0 | 4.4167 | 0 | 3.7333 | 0 | 3.4333 | 0 | 3.8833 | 0 | 3.55
σ=0.3mm | 0 | 5.1 | 0 | 5.9667 | 0 | 4.9 | 0 | 4.9833 | 0 | 5.6667 | 0 | 5.8167 | 0 | 6.7333 | 0 | 5.5 | 0 | 4.0167 | 0 | 3.75 | 0 | 4.6667 | 0 | 3.75 | 0 | 1.8 | 0 | 1.9 | 0 | 2.1167 | 0 | 1.75 | 0 | 1.2833 | 0 | 1.15 | 0 | 1.5167 | 0 | 1.15

## Slide 21: WORK in PROGRESS
Combining errors coming from structure deformations (thanks to Luca P. and Luca D.) (systematic) and the random contributes
0.1 T/m
0 T/m
Without multipoles
With multipoles
MF 1
MF 1.5
MF 1.5
Without multipoles
With multipoles
MF 1.5
MF 1.5
MF 1
MF 1
MF 1

## Slide 22: First results
MAX space occupied by the beam < 21mm
0.1 T/m without multipoles representation
MF 1.5
MF 1

## Slide 23: First results
MF 1.5
MF 1
0.1 T/m without multipoles representation

## Slide 24: WORK in PROGRESS
Simulation of all the configurations to see if there is some incongruence
Effect of losing the beam on the coils in SC dipoles

## Slide 25: Energy deposition studies on a SC magnet
Material:
-Aluminum (Al)
-Insulated cable
-Niobium-Titanium (Nb-Ti)
-Strand
-Copper (Cu)
Particle beam
Carbon ions:
- 400 MeV
- 115 MeV
Proton:
- 230 MeV
- 62 MeV

## Slide 26
 | C 400 | C 115 | p 230 | p 62 | MeV/u
Al | 2,46E+07 | 3,70E+06 | 3,13E+09 | 1,45E+08 | [Particles/s]
 | 2,46E+07 | 3,70E+06 | 3,13E+09 | 1,45E+08 | [Particles]
Cable | 1,18E+07 | 2,26E+06 | 1,10E+09 | 8,45E+07 | [Particles/s]
 | 1,18E+07 | 2,26E+06 | 1,10E+09 | 8,45E+07 | [Particles]
NbTi | 8,33E+06 | 2,02E+06 | 9,62E+08 | 7,44E+07 | [Particles/s]
 | 8,33E+06 | 2,02E+06 | 9,62E+08 | 7,44E+07 | [Particles]
Strand | 6,65E+06 | 1,69E+06 | 6,01E+08 | 6,13E+07 | [Particles/s]
 | 6,65E+06 | 1,69E+06 | 6,01E+08 | 6,13E+07 | [Particles]
Cu | 5,84E+06 | 1,56E+06 | 4,46E+08 | 5,12E+07 | [Particles/s]
 | 5,84E+06 | 1,56E+06 | 4,46E+08 | 5,12E+07 | [Particles]
Quencing 10 mWatt/cm3
Particles/s

## Slide 27: Overlook
Taking a conservative approach, the material selected for the coil composition consists of a strand comprising 36% Nb-Ti and 64% Cu.
The energy deposition study shows that the most dangerous particles are those with minimal energy. Pay attention: particles with maximum energy levels are also unsafe.
Energy deposition studies for the straight magnet geometry consider all the above discussed factors. Currently, simulations not only include front-face impacts but also various scenarios involving rotating beams

## Slide 28
THANK YOU