## Slide 1: Signal Induction
DRD1 Gaseous Detector School
Werner Riegler, CERN, werner.riegler@cern.ch
Nov. 29 2024

## Slide 2
Qind1
Qind2
Qind3
y
Qind0
q
V1=0
V2=0
V3=0
V0=0
x(t)
I2(t)
I3(t)
I1(t)
I0(t)
q
X X X X X
XX X X
X X X
XXX X XX
X XXX X
E
E - 

x=0 x=D
Outline
Signal induction principles and theorems
Energy Loss, charge deposit, transport properties
Examples: Wire chambers, GEMs, Micromegas
Theorems and examples for detectors with resistive elements

## Slide 3
Lecture Series
https://indico.cern.ch/event/843083/
Signals in particle detectors,
CERN Academic Training Lectures 2019
Signals, Noise and Signal processing in Particle Detectors,
CERN Academic Training Lectures 2024
https://indico.cern.ch/event/1458895/
Particle Detection with Drift Chambers,
Springer, 2008 2nd Edition

## Slide 4
Potential for a given charge distribution
Charge distribution with boundary condition, Poisson equation:
In regions of space without charge and given boundary, Laplace equation:
Gauss Law:
Electrostatics
x
y
⍴(x)
x
y
⍴(x)
A

## Slide 5
x
y
⍴(x)
A
Uniqueness
Let’s assume we have two solutions of the same Poisson equation
The difference between these two solutions must satisfy the Laplace equation
In general it holds that
Using the fact that and applying Gauss’ theorem we have
In case on the surface we have on the surface and the left side vanishes. That means that the right hand side must also vanish and we have
Defining the potential on the entire surface therefore uniquely defines the electric field in the volume. The same holds in case we define on the entire surface. Evidently the uniqueness theorem also holds in case we define on a fraction of the surface and on the rest of the surface.
x
y
⍴(x)
A

## Slide 6
Conductors
In conducting materials, charge will move as long as there is an electric field present.
A static situation is only achieved once the electric field inside the conductor is zero and therefore also the charge density inside the conductor is zero. All the charge is sitting on the surface of the conductor.
The electric field on the surface of the conductor must be perpendicular to the conductor, since a component parallel to the conductor would again move the charge.
Moving a test charge across the surface of the conductor one crosses the field lines perpendicular and no work is performed. A conductor surface is therefore and equipotential surface.
Since the electric field inside the conductor is zero and the field-lines are perpendicular to the surface, we can use Gauss law to relate the surface charge density to the field on the surface:
 The charge density (C/m2) at a given point on the conductor surface is equal to ε0 times the electric field in this point.
E
Conductor
A
E
E=0
σ
A
σ

## Slide 7
Induced charge on metal electrodes
V1
V2
V0
V3
𝝋(x)
⍴(x)
Poisson equation:
The charge on the electrode is given by the integral of the electric field over the surface of the electrode

## Slide 8
The Principle of Signal Induction on Metal Electrodes by Moving Charges

## Slide 9
A point charge q at a distance y’ above a grounded metal plate ‘induces’ a surface charge.
Induced charge on an infinite plane
y’
+q
-
-
-
-
-
-
-
-
-

## Slide 10
In order to find the charge induced on an electrode we therefore have to
Solve the Poisson equation with boundary condition that 𝜑=0 on the conductor surface.
Calculate the electric field E on the surface of the conductor
Integrate σ=e0E over the electrode surface
q
y’
+
-
-
-
-
-
-
-
-
-
Induced charge on an infinite plane

## Slide 11
The solution for the field of a point charge in front of a metal plate is equal to the solution of the charge together with a (negative) mirror charge at y=-y’.
+q
y’
-
-
-
-
-
-
-
-
-
=
-q
(x’, y’)
E
The potential for the above problem is therefore:
(x’, -y’)
y
x
+q
Induced charge on an infinite plane

## Slide 12
We therefore find a surface charge density of
And therefore a total induced charge of
y’
+q
-
-
-
-
-
-
-
-
-
Qind
Induced charge on an infinite plane

## Slide 13
q
q
Moving the point charge closer to the metal plate, the surface charge distribution becomes more peaked
The total induced charge is however always equal to –q,
The charge is just rearranged on the surface and no current is flowing between the plate and ground..
I(t)=0
y
x
Charge at position x’=0, y’, z’=0
Induced charge on an infinite plane

## Slide 14
q
If we segment the grounded metal plate and if we ground the individual strips, the surface charge density doesn’t change with respect to the continuous metal plate.
Qind
The charge induced on the individual strips is now depending on the position (x’, y’) of the charge.
y
x
-w/2
-w/2
(x’, y’)
Induced charge on strip electrode

## Slide 15
q
If the charge is moving with a velocity v towards the strip we have
V
Qind(t)
y
x
-w/2
-w/2
Iind(t)
While the charge is moving there is a current flowing between the strip and ground.
 The movement of the charge induces a current.
y0
Induced current on strip electrode

## Slide 16
⍴(x)
⍴(x)
Reciprocity theorem
This is simply a result of the fact that the Coulomb force depends only on the relative distance of the charges and not on absolute position in space.
The expressions can be interpreted as the work needed to move one charge distribution in the electric field of the other charge distribution, actio = reactio
Sounds like a trivial statement, but has very practical consequences.
Two arbitrary charge distributions

## Slide 17
Reciprocity theorem
y
x
y
x
Q1
(x1,y1,z1)
Q2
(x2,y2,z2)
Example: two point charges
 Correct but not really useful

## Slide 18
Reciprocity
y
x
y
x
Grounded metal electrode of arbitrary shape
Point charge q at (x’, y’, z’)
 Knowing the potential for the electrode at voltage V0 in absence of any external charges we can directly calculate the charge induced on the grounded electrode in presence of a point charge !!
q
V0
V=0
(x’, y’, z’)

## Slide 19
r
q
R1
Qind
Induced charge on a grounded sphere
r
R1
V0

## Slide 20
Qind1
Qind2
Qind3
x
y
Qind0
q
V1=0
V2=0
V3=0
V0=0
Theorem, induced charge
The charge induced on a grounded conducting electrode by a point charge q at position x can be calculated the following way:
Remove the point charge, put the electrode in question to potential Vw while keeping all other electrodes at ground potential.
This defines the weighting potential 𝜓n(x) of this electrode and the induced charge is
We therefore do not have to solve the Poisson equation for a point charge but we just have to solve the Laplace equation for the given boundary conditions on the electrodes.
For detectors with long electrode, like wire chambers, RPCs, silicon strip detectors, we only have to solve the 2D Laplace equation instead of the 3D Poisson equation.
Specifically for numeric field calculations this is much easier and numerically stable.
x
y
Qind0
V1=Vw
V2=0
V3=0
V0=0
𝜓1(x)

## Slide 21
Qind1
Qind2
Qind3
x
y
Qind0
q
V1=0
V2=0
V3=0
V0=0
Theorem, sum of induced charges
x
y
Qind0
V1=Vw
V2=0
V3=0
V0=0
𝜓1(x)
The sum of all induced charges is defined by the sum of all weighting potentials
This sum is related to the situation where all electrodes are put to voltage Vw. This will however result in a constant potential Vw in the entire volume between the electrodes, so we have
The sum of all charges induced by a charge q is equal to –q, in case the geometry is such that one electrode encloses all others.

## Slide 22
x
Theorem, induced current
The current induced on a grounded conducting electrode by a point charge q moving along a trajectory x(t) can be calculated the following way:
This weighting field En(x) is given by
 Ramo-Shockley theorem
Since the sum of all induced charges is constant and equal to –q at any time, the sum of all induced currents is zero at any time, in case there is one electrode enclosing all the others.
x
y
Qind0
V1=Vw
V2=0
V3=0
V0=0
𝜓1(x), E1(x)
Qind1
Qind2
Qind3
y
Qind0
q
V1=0
V2=0
V3=0
V0=0
x(t)
I2(t)
I3(t)
I1(t)
I0(t)
q

## Slide 23
Ramo Shockley theorem (Reciprocity theorem)

## Slide 24
Signal polarity
I(t)
The definition of the induced current
defines that the current arrow points away from the electrode.
A positive current refers to a reduction of charge on the electrode.
The signal is positive if:
Positive charge is moving from electrode to ground or
Negative charge is moving from ground to the electrode
The signal is negative if:
Negative charge is moving from electrode to ground or
Positive charge is moving from ground to the electrode
+q
-
-
-
-
-
-q
+
+
+
+
+
I(t)
+q
-
-
-
-
-
-q
+
+
+
+
+
Positive Signal
Negative Signal

## Slide 25
Qind1
Qind2
Qind3
x
y
Qind0
q
V0=0
Theorem, total induced charge
The total charge flowing between the electrode and ground when the charge q is moving from point x0=x(t0) to x1=x(t1)
This charge is independent of the path between x1 and x2.
If a pair of charges q, -q is produced at position x0 and q is moving to x1 while –q is moving to x2, the charge flowing between electrode n and ground is
x0
I2(t)
I3(t)
I1(t)
I0(t)
x1
q
Qind1
Qind2
Qind3
x
y
Qind0
q
V0=0
x0
I2(t)
I3(t)
I1(t)
I0(t)
x1
q
x2
-q
-q

## Slide 26
Theorem, total induced charge
If a pair of charges q, -q is produced at position x0 and q is moving to x1 while –q is moving to x2, the charge flowing between electrode n and ground is
In case q moves to the surface of electrode n (where 𝜓n(x) = Vw) and –q to the surface of any other electrode (where 𝜓n(x) is 0), to total charge that has been flowing between the electrode and ground is equal to q.
If charges are created in pairs it holds that the total charge flowing between an electrode and ground is equal to the total charge that has arrived at this electrode, once ALL charges have arrived at the electrodes.
x
y
Qind0
q
V0=0
x0
I2(t)
I3(t)
I1(t)
I0(t)
x1
q
x2
-q
-q

## Slide 27
Parallel plate geometry
+q
(x, y)
x
y
Q1ind
Q2ind
I1ind(t)
I2ind(t)
d
0
x
y
Vw
𝜓1, E1
2
1
d
0
y
x
Vw
𝜓2, E2
2
1
d
0
Weighting potential, electrode1 Weighting potential, electrode 2 Induced charge
Sum of the weighting potentials Sum of induced charges

## Slide 28
Parallel plate geometry
+q
v
x
y
Q1ind
Q2ind
I1ind(t)
I2ind(t)
d
0
x
y
Vw
𝜓1, E1
2
1
d
0
y
x
Vw
𝜓2, E2
2
1
d
0
y0
Induced current Weighting field, electrode1 Weighting field, electrode 2
I1(t)
t
Total induced charge on electrode 1

## Slide 29
The Heaviside step function is defined as
In case we have a constant current signal with amplitude I0 starting at t=0 and lasting until t=t1 we have
In case we have a constant current signal with amplitude I0 starting at t=t1 and lasting until t=t2 we have
Heaviside step function
t=0
t=t1
I0
t=0
1
t=0
t=t1
I0
t=t2

## Slide 30
Parallel plate geometry
+q
v1
x
y
Q1ind(t)
Q2ind(t)
I1ind(t)
I2ind(t)
d
0
y0
Two charges +q, -q moving from y0 to the electrodes with velocities v1 and v2, arriving at the electrodes at times t1 and t2
Induced currents
Total induced charges
In all physics processes, pairs of charges with opposite sign are produced at the same position, which results in the fact that the total induced charge is equal to the charge that has arrived at the electrode, once ALL charges have arrived at the electrodes.
-q
v2
I1(t)
t
0

## Slide 31
Parallel plate geometry
+λ
v1
x
y
Q1ind(t)
Q2ind(t)
I1ind(t)
I2ind(t)
d
0
y0
Charge particles passing sensors leave a trail of positive and negative charges (electrons/Ions, electrons/holes). Assuming a uniform distribution along the track we have two ‘line charges’ + λ, - λ (C/cm), moving with velocities v1 and v2, with the last charges arriving at the electrode at t1 and t2
The induced current due to the movement of these charges is the sum of two ‘triangles’, Q= λd
The total induced charge on electrode 1 is
-λ
v2
I1(t)
t

## Slide 32
Charge Deposit

## Slide 33
A charged particle passing through a piece of material will produce a trail of ionization.
Since the individual interactions with the atomic electrons are independent, the number of primary interactions follows a Poisson distribution.
The probability f(E) for transferring an energy E to the atomic electron in an interaction is given by the Rutherford crossection at large energy transfers and by the atomic atomic shell structure at low energy transfers.
Scattering on free electrons
Excitation and ionization
L()
Fluctuations of the Energy Loss
X X X X X
XX X X
X X X
XXX X XX
X XXX X
E
E - 

x=0 x=D
f(E)
Convolution

## Slide 34
Charge deposit in gases
XXXX
The solid lines represent measurements from
The points represent results from a PAI model
Average distance between clusters
Argon 0.41mm
Neon 0.83mm
Isobutane 0.12mm
For single particle detection, internal gain is used in gas detectors
1.5cm Argon MP approx. 67/cm

## Slide 35
Transport of electrons and ions in gases

## Slide 36
Electrons are completely ‘randomized’ in each collision.
The actual drift velocity v along the electric field is quite different from the average velocity u of the electrons i.e.
about 100 times smaller.
The velocities v and u are determined by the atomic crossection ( ) and the fractional energy loss () per collision (N is the gas density i.e. number of gas atoms/m3, m is the electron mass.):
Because ( )und () show a strong dependence on the electron energy in the typical electric fields, the electron drift velocity v shows a strong and complex variation with the applied electric field.
v is depending on E/N: doubling the electric field and doubling the gas pressure at the same time results in the same electric field.

Ramsauer Effect
Transport of Electrons in Gases: Drift-velocity

## Slide 37
Typical electron drift velocities are v=20-140um/ns (20 000-140 000m/s). The microscopic velocity u is about ca. 100mal larger.
Transport of Electrons and Ions in Gases
Typical ion drift velocities are around 1000 times smaller than electron drift velocities.
7.5 kV/cm @ 1 bar
Argon Ions at 2kV/cm  0.03um/ns
1000x slower than electrons

## Slide 38
GEM Detector

## Slide 39
v1~50um/ns
v2~ 100um/ns
Gas Electron Multiplier GEM
Single electrons moving through a GEM hole are multiplied in the strong electric field.
Ions are moving to the top side of the GEM or back into the transfer gap.
Electrons are exiting the holes and move to the next amplification stage.
In the last gap the ‘Induction gap’ the electrons are moving to the readout electrodes and induce the signal.
The geometry is equivalent to a parallel plate chamber with only electrons moving through the entire gap.
The signal from a single electron starting in the conversion gap is a ‘box’ with a duration equivalent to the electron transit time in the induction gap approx. 10ns.
v1
x
y
Q1ind(t)
Q2ind(t)
I1ind(t)
I2ind(t)
d
0
y0
-q
I1(t)
t
10ns

## Slide 40
v1~50um/ns
v2~ 100um/ns
Gas Electron Multiplier GEM
The conversion gap above the top GEM is the volume where charge particles deposit e+e- pairs.
Typical thickness is 3-15mm (2.5m for the ALICE TPC !)
For a continuous charge deposit the signal has trapezoid shape where the length is dominated by the drift time in the conversion gap.
I1(t)
t
60ns
10ns
10ns
10ns
50ns

## Slide 41
Parallel field avalanches

## Slide 42
Gas avalanche multiplication
E
Ions
Electrons
At sufficiently high electric fields (100kV/cm) the electrons gain energy in excess of the ionization energy  secondary ionzation etc. etc.  exponential increase  avalanche  Townsend coefficient ⍺
The current induced by these moving electrons is
The ions move from the point of creation in opposite direction from the point of creation and the induced current is
The total charge induced by the electrons and the ions is
x
I1ind(t)
electrons
Ions
Gain 103: ad=7
Gain 104: ad=9.2

## Slide 43
MICROMEGA
3mm
0.1mm
50um/ns
200um/ns
Electrons movement in the induction gap takes about 0.1mm/v1=0.5ns.
Collecting all electrons from the drift gap takes e.g. 3mm/v1=60ns.
The MICROMEGA electron signal has a length of about 60ns.
Ion movement – e.g. Argon Ions take 130ns for 50kV/cm and 100um gap, so the total length of the ions component is around 180ns.
When using ‘fast’ electronics one does not integrate the full charge  ballistic deficit.
MicroMeshGAS detectors
50-100um, 50kV/cm
3-15mm
Single primary electron signal
Single primary electron signal

## Slide 44
Wire Chambers

## Slide 45
Wire Chamber Signals
Wire with radius (10-25m) in a tube of radius b (1-3cm):
Electric field close to a thin wire (100-300kV/cm). E.g. V0=1000V, a=10m, b=10mm, E(a)=150kV/cm
Electric field is sufficient to accelerate electrons to energies which are sufficient to produce secondary ionization  electron avalanche  signal.
b
b
a
Wire

## Slide 46
Wire Chamber Signals
The electrons are produced very close to the wire, so as a first approximation we can assume that the signal is only due to Ntot ions moving from the wire surface to the tube wall:
Ions move with a velocity proportional to the electric field.
Weighting field of the wire: Remove charge and set wire to Vw while grounding the tube wall.
The induced current is therefore
b
b
a
E
Ions
Electrons

## Slide 47
Wire Chamber Signals
b
a
ATLAS muon drift tubes
The signal from a single primary electron has 1/(t+t0) shape with a very long tail …
Typically only a small fraction (e.g. 10%) of the total avalanche charge is induced during the electronics integration time.
Ballistic deficit. In Micropattern detectors one can integrate all the charge  10 times lower gas gain for the same signal.
3.5us
35us
3.5ms

## Slide 48
ATLAS muon system drift tubes
Tubes of 3cm diameter are assembled into chambers.
The first electrons arriving at the wire determine the distance of the track from the wire  drift tube.
The last arriving electrons are originating from r=1.5cm, so the last electrons always arrive at the same time.
The long signal tail needs dedicated electronics filtering to ensure limited deadtime.
80um position resolution over a few thousand m2 area with only 330 000 channels !

## Slide 49
Detector with Resistive Elements

## Slide 50
Quasi-static approximation of Maxwell’s equations
Assuming a conductivity sigma of the material we have a current according to
Maxwell’s equations for this situation
The current je(x, t) is an ‘externally impressed’ current, which is related to the ’externally impressed’ charge density ⍴e by
If we assume that this impressed current is only changing slowly we can neglect Faraday’s law and approximate
and we can then write the electric field as the gradient of a potential, an by taking the divergence of the last equation …

## Slide 51
Performing the Fourier Transform of the quasi-static equation
we find
So we can write this equation as
This is the Poisson equation with an effective permittivity !
We can therefore find the time dependent solutions for a medium with a given conductivity by solving the electrostatic Poisson equation in the Fourier domain !
Knowing the electrostatic solution for a given permittivity ε(x) we just have to replace ε(x) by ε(x)+σ(x)/iω and perform the inverse Fourier transform !
Quasi-static approximation of Maxwell’s equations

## Slide 52
Extension of the Ramo Shockley theorem

## Slide 53
Theorem, induced current
Applying the delta voltage pulse to the electrode in question we find the potential 𝜓n(x, t) and the field En(x, t) from which the induced current can be calculated the following way:
 Ramo-Shockley theorem extension for conducting media
Note that En is not physical potential, since the delta function gives it a dimension of V/cm s.
In case the material is an insulator there is no time dependence of the weighting field and we recuperate Ramo’s theorem.
x
y
Qind0
V1(t)=Vwẟ(t)
V2=0
V3=0
V0=0
𝜓1(x, t), E1(x, t)
ε(x, s)
σ(x, s)
Qind1
Qind2
Qind3
x
y
Qind0
q
V0=0
x(t)
I2(t)
I3(t)
I1(t)
I0(t)
q
ε(x, s)
σ(x, s)

## Slide 54
q
x(t)
Iind2(t)
Iind3(t)
Iind1(t)
q
ε(x, s)
σ(x, s)
Vind1(t)
q
x(t)
q
ε(x, s)
σ(x, s)
Vind2(t)
Vind3(t)
Equivalent circuit, Impedance elements
z11(s)
z22(s)
z33(s)
z13(s)
z23(s)
z12(s)
In case the electrodes are not insulated but connected with discrete linear impedance components we can consider them as part of the medium and we therefore just have to add these elements in the equivalent circuit.

## Slide 55
https://cds.cern.ch/record/2890572
Djunes Janssens

## Slide 56
Djunes Janssens

## Slide 57: Micromegas toy-model example
V
V
  The time-dependent weighting potential is comprised of a static prompt and a dynamic delayed component:
Static weighting potential of a readout strip
Dynamic weighting potential of a readout strip with resistive layer
Djunes Janssens

## Slide 58
Resistive strip bulk Micromegas
  Calculated dynamic weighting potential
While the previous slide the embedded readout electrodes were beneath a thin resistive layer, the ATLAS-type MM, used in the NSW upgrade, instead features perpendicular resistive strips over a dielectric foil.
Here we use a 2D strip readout to capture the delayed component coming from the resistive strips.
T. Alexopoulos, et al., Nucl. Instrum. Meth. A 640 (2011) 110.
M. Byszewski and J. Wotschack, JINST 7 C02060 (2011).
Djunes Janssens

## Slide 59
Resistive strip bulk Micromegas
  APV25
  Electrode 5
  Electrode 6
  After having calculated the signals induced on the strip electrodes, the electronics with which the detector is read out needs to be taken into account.
Djunes Janssens

## Slide 60
Resistive strip bulk Micromegas
-
-
For the comparison we look at the average induced current response of neighboring strips. This averaging is performed over muon events positioned between the leading and the next-to-leading strip.
Djunes Janssens

## Slide 61
Djunes Janssens

## Slide 62: 6-gap MRPC
HV layer
  The dynamic weighting potential was calculated using COMSOL and then imported into Garfield++ for the induced signal calculations. Given a graphite layers with O(100 kΩ/□), the signal induced by electrons remains unaffected.
10 kΩ/□
10 kΩ/□
Consistent with G. Battistoni et al., NIM in Physics Research 202 (1982) 459.
Djunes Janssens

## Slide 64: Signal formation in a MicroCAT detector
  The MicroCAT’s two-dimensional interpolating readout structure allows for a reduced number of electronic readout channels without loss of spatial resolution.
  This resistive readout concept has recently enjoyed renewed interest with the development of a DC-Coupled LGAD device: arXiv:2204.07226 [physics.ins-det].
F. Bartol et al. J. Phys. III France, 6 (1996), p. 337
A. Sarvestani et al., Nucl. Instrum. Meth. A 410 (1998) 238–258
H. Wagner et al. Nucl. Instrum. Meth. A 482 (2002) 334–346
Weighting potential map for one readout node
Djunes Janssens

## Slide 65
MicroCAT resistive position interpolation readout
  During production, the resistivity can fluctuate on the surface of the resistive layer. This could make the timing response or reconstruction capability of your detector non-uniform over the active area.
R. Arcidiacono et al., arXiv:2211.13809v1 [physics.ins-det]
  Correction map
  Map of non-uniform resistivity in one readout cell
Djunes Janssens

## Slide 66
Summary
The movement of charges induces signals on metal electrodes.
The signal shapes are therefore determined by the dynamics of the charges inside the detector i.e. by the movement of electrons and Ions in gaseous detectors.
These signals can be calculated by means of weighting potentials and weighting fields that are solutions of the Laplace equation.
The formalism can be extended to detectors containing resistive elements by use of time dependent weighting fields.
Signals in GEM detectors are mainly due to the movement of the electrons in the induction gap.
Signals in Micromegas detectors are characterized by a short electron ‘spike’ and a longer ‘flat’ ion tail.
Signals in wire chambers are mainly due to the movement of ions and are characterize by a long hyperbolic 1/(t+t0) tail.