## Slide 1: HTS Cavity Copper Cold Test S11
4 MAR 2024
Greg P. Le Sage

## Slide 2: A copper spool was made to attach the cavity to the TM01 mode converter
Small steps on the faces ensure good electrical contact.

## Slide 3: S11 measurement was performed
Maximum number of samples
Narrow band calibration
Other modes did not move with cavity probing, so they are external and irrelevant.

## Slide 4: Compare measurement to simulation
MATLAB Fit of S11 Data 28 FEB 2024: MEASUREMENT
Beta = 0.4321 SE 0.000069394
Q0 = 13,896.293 SE 3.3289
f0 = 11.04188 GHz SE 0.0000000642
QE = 32,162.916 calculated
QL = 9,703.712 calculated
HFSS Eigenmode SIMULATION
Copper Cavity with PML QL = 12,058.9 at 11.4278 GHz, Df = 0.0003498
PEC Cavity with PML QE = 56,148.1 at 11.4282 GHz, Df = 3.67716e-5
Copper Cavity with Short Q0 = 15,385.5 at 11.40088 GHz, Df = 0.008132
HFSS Beta = Q0/QE = 0.2740
|QL – (1/Q0 + 1/QE)-1| = 17.475
Recall S21 Q0 measurement was 12,800

## Slide 5: HFSS modal solution (Thanks to the cluster!)
Discrete frequency sweep
Use same MATLAB fit
Beta = 0.3127 SE 0.000018240
Q0 = 15,311.27 SE 1.7853e-10
f0 = 11.4016 SE 3.0106e-08
QE = 48,968.74 calculated
QL = 11,664.18 calculated
Beta calc = 0.3127