## Slide 1: Update on the BPM study for Run 2c
Laurence Stant
13.03.2024 / Daresbury, UK

## Slide 2: Plan for AWAKE BPMs
STATUS
We want to improve the proton BPM resolution from 70 um to around 20 um.
We need to add 12 electron BPMs plus 3 spares for the Run 2c layout.
The TRIUMF agreement for producing and maintaining eBPMs has expired and software/firmware support is not available.
PROPOSAL
Develop a new integrated pBPM frontend/backend based on RF system-on-chip technology being developed for other CERN BI projects (such as HiLumi LHC BPMs).
Develop new eBPM electronics similar to that for the pBPM and therefore bring in-house to CERN.
16 proton BPMs
~12 Existing electron BPMs
12 New electron BPMs
EXPERIMENTAL AREA
TT41
TT43

## Slide 3: Plan for AWAKE BPMs
STATUS
We want to improve the proton BPM resolution from 70 um to around 20 um.
We need to add 12 electron BPMs plus 3 spares for the Run 2c layout.
The TRIUMF agreement for producing and maintaining eBPMs has expired and software/firmware support is not available.
PROPOSAL
Develop a new integrated pBPM frontend/backend based on RF system-on-chip technology being developed for other CERN BI projects (such as HiLumi LHC BPMs).
Develop new eBPM electronics similar to that for the pBPM and therefore bring in-house to CERN.
16 proton BPMs
~12 Existing electron BPMs
12 New electron BPMs
Consider pBPMs as four identical groups of four, plus the exp. area
EXPERIMENTAL AREA
TT41
TT43

## Slide 4: Proton BPM Layout - Existing
Each of the 16 transfer line pBPMs uses a local frontend below the beam.
Position information is sent over a coaxial cable to a shared backend located near the AWAKE control room.
The five pBPMs located in the experimental area are connected to frontends grouped in a rack in the adjacent corridor.
BPM
Rack
Frontend
Frontend
Frontend
Frontend
Backend
Tunnel
Near Control Room
Tech. Network
X, Y data @ 10 Mbps
64 m
32 m
32 m

## Slide 5: SoC Platform – Fronted and Backend
Development shared with LHC HiLumi project.
Xilinx/AMD Zynq RFSoC featuring:
Ultrascale+ FPGA fabric
8x 5 GSPS 14-bit ADC channels
8x 10 GSPS 14-bit DAC channels
High performance ARM processing cores
Flexible timing and integration options
Procurement
ZCU208 development board
Researching system-on-module for production, with CERN-built carrier.

## Slide 6: Time Multiplexing
Because of the very high sampling rate of the RFSoC, we don’t need to time-stretch pulses (via lowpass filtering) to achieve adequate signal to noise ratios.
This means we can time multiplex multiple pickups using cable delays and reduce the required number of ADC channels.
Using this technique, we can also remove systematic errors (position offsets) arising from any difference between channel gains for the same axis:
  Cable drift
  Component response variation
  ADC drift
Time
Voltage
Y Combiner
Delay
Delay

## Slide 7: Proton BPM Layout - Proposed
New RFSoC frontends cover four pBPMs each.
Position data is calculated onboard as before, but is now available directly to the technical network through FESA (the plan is for one FESA device per BPM).
Ethernet connection runs over fibre to adapters in the eBPM racks which connect to a network switch.
Experimental area frontends located in eBPM racks as space will be made with the new eBPMS.
Frontend
Y Combiner
X Combiner
Tunnel
Fibre to Tech. Network
Corridor beside AWAKE

## Slide 8: Beam Simulation via CST and Python
Simulate pilot and 3E11 proton beams in CST and create fits for intensity and position from pickups.
Simulate each step of the analogue signal path from pickup to ADC.

## Slide 9: Beam Simulation via CST and Python
Simulate pilot and 3E11 proton beams in CST and create fits for intensity and position from pickups.
Simulate each step of the analogue signal path from pickup to ADC.
Sample 1000 pulses with random quantisation noise from ADC based on datasheet values and lab measurements of sinusoids (8.5 to 9.5 ENOB).
Calculate position using on RMS of samples with high order sensitivity fit and obtain standard deviation as “resolution”.

## Slide 10: Beam Simulation via CST and Python
Simulate pilot and 3E11 proton beams in CST and create fits for intensity and position from pickups.
Simulate each step of the analogue signal path from pickup to ADC.
Sample 1000 pulses with random quantisation noise from ADC based on datasheet values and lab measurements of sinusoids (8.5 to 9.5 ENOB).
Calculate position using on RMS of samples with high order sensitivity fit and obtain standard deviation as “resolution”.
Vary bandwidth of frontend from 800 MHz to 100 MHz.
For all simulations, set input attenuation such that ADC is 70% full scale at ¼ aperture beam offset.

## Slide 11: Simulation Results – Science Beam
We can improve resolution by reducing bandwidth. This stretches the pulse and gives more samples for the RMS. The stretched pulse is weaker but still no amp needed for 1E11 beam.
Error increases with beam offset due to increasing pickup sensitivity coefficient. Typical beam offset within 5 mm.

## Slide 12: Simulation Results – Pilot
For the pilot beam we would require amplification to achieve the same resolution. For low bandwidths this becomes significant.
If no amplification is used the resolution will degrade. We envisage ~200 MHz bandwidth so this would reduce to ~100 µm.

## Slide 13: Lab Results from Pulse Generator
Simulated waveforms were exported to a pulse generator and measured by the RFSoC, then run through the same processing as before.
Noise/resolution was shown to be limited by pulse generator, but we must be close to the ADC effective resolution so I would not bet on any improvement over these numbers.
Filters should improve the resolution as attenuation was removed to match the amplitude with the 800 Mhz signal.
Beam Offset | 800 MHz | Filter 1 (300 MHz) | Filter 2 (220 MHz) | Filter 3 (180 MHz)
0 mm | 15 µm | 17 µm | -- | 18 µm
5 mm | 16 µm | -- | -- | --
10 mm | 21 µm | -- | -- | --
15 mm | 26 µm | 24 µm | 28 µm | 28 µm
σ of 1000 acquisitions
This agrees well with 9.5 ENOB simulation

## Slide 14: Next steps for pBPM
Perform more lab testing to check filter behaviour.
Measure delay effect on pulse crosstalk and hence accuracy (beam offset).
Measure some real LEP button signals on SPS with the RFSoC hardware.
Attach a complete combiner setup to an AWAKE pBPM plane and measure during this run. This would be simple on an experimental area pBPM in the grouped rack as network access is adjacent.

## Slide 15: AWAKE Electron BPMs
About to begin the study into these.
Planning to share as much of the pBPM design as possible.
Should achieve similar performance to the pBPM, especially as we would not use combiners and so can match amplitudes of each channel to better utilise the ADCs.
Reuse of TRIUMF frontends is possible, but this would require updating the design against components obsoleted especially after the global chip shortage.
We would then need to onboard production, test and maintenance for a board which we will not use anywhere else!

## Slide 16: Conclusion
Promising results from initial pBPM studies with RFSoC frontend.
We should achieve around 15 µm resolution on the pBPM with this setup.
New eBPM configuration will now be studied, with the objective of a shared design and resolution matching the existing system.

## Slide 18: ADC ENOBs
Lab measurements of ENOB with sinusoids show there is a knee just above 70% ADC full-scale, after which the signal is distorted.
However for our RMS detector we are not sensitive to this distortion.
So, we should expect an ENOB between 9 and 10.
This value is used when simulating the quantisation noise in the virtual ADC.
0.35
70% ADC FS

## Slide 19: Rack space
New BPMs should be 2U plus 1U gap.
Four pBPMs per 3U.
For exp. area (5 pBPMs) this means 6U.
Retain 14 eBPMs, add 12.
Two eBPMs per 3U.
So this is 39U.
Should not require RF distribution kit.
So total for both racks = 39U + 6U = 45U.
Racks are 35U. Space in right for rest.