## Slide 1: Characterization of pixelated high-Z sensor for X and gamma-ray detection
Petr Smolyanski & Benedikt Bergmann
Institute of Experimental and Applied Physics, Czech Technical University in Prague
Petr.smolyanskiy@utef.cvut.cz
Benedikt.bergmann@utef.cvut.cz

## Slide 2
Radiation imaging detector

## Slide 3: Timepix3 and Timepix4
Timepix3:
256 x 256 pixels with 55 µm pitch (1.98 cm2)
Data-driven readout (Max. count rate 40 Mpix/s)
Per-pixel dead time of 475 ns
Measurement of energy and time (Δt = 1.56 ns)
Minimal detectable energy per pixel 2.4-3.5 keV
  Sensor layer (Silicon, GaAs, CdTe, ...) flip-chip bump bonded to the ASIC
Timepix3 with chipboard.
Timepix4:
512 x 448 pixels with 55 µm pitch
Data-driven readout (Max. count rate 640 Mpix/s)
Measurement of energy and time (Δt = 200 ps)
Minimal detectable energy per pixel ~2 keV (?)

## Slide 4: Selection of the sensor material
Physics of interaction with matter
  Detection efficiency vs photon energy
  Compton vs photoabsorption
Semiconductor properties, in particular for high-Z
  Charge carrier drift  shaping times of amplifiers
  Charge trapping  signal loss and deformation of energy peaks in spectra
  Polarization effects  time/radiation dependence of energy measurement
  Internal flourescence & their range  escape peaks or image smearing (imaging application)

## Slide 5: Characterization of high-Z sensors in charged particle beams

## Slide 6: Measurement principles: Experiment design
Simultaneous measurement of time and energy in each pixel together with the tracking allows a determination of
Charge carrier drift times  mobility µ
Charge losses during drift across the sensor thickness  CCE
Measurement of the lateral spread  Models for diffusion and repulsion
as a function of z

## Slide 7: Grazing angle irradiation: Sensitivity to charge carrier properties
e- and h+ drift described by
Mobility of e-/h+
Charge carrier drift motion:
Electric field parametrization:
Si: ;
CdTe:
Bias voltage; Depletion voltage; Sensor thickness
+
-
-
-
-
-
-
-
-
-
-
+
+
+
+
+
+
+
-
-
+
+
+
-
+
t = 0 ns
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
t ~ 15 ns
t ~ 25 ns
-
-
-
-
z
Bergmann et al. Eur. Phys. J. C (2017) 77: 421. https://doi.org/10.1140/epjc/s10052-017-4993-4
Bergmann et al., Eur. Phys. J. C (2019) 79: 165. https://doi.org/10.1140/epjc/s10052-019-6673-z
Small pixel effect:
Measureable signal if charge carriers are in close vicinity to the pixel electrodes
Ramo-Theorem:
Qind = [Φ(z(ti)) - Φ(z(t0))] x Q0
 Look-up table: z(tmeas.,Emeas.)

## Slide 8
https://link.springer.com/article/10.1140/epjc/s10052-019-6673-z

## Slide 9: 40 GeV/c pion beam: 2 mm thick CdTe (Acrorad) sensor
Drift time as a function of interaction depth
;
Negligible signal loss at bias above 300 V

## Slide 10: 40 GeV/c pion beam: 3D track reconstruction – 2 mm CdTe
cosmic µ
e-/+ like event
pair production (!?)
fragmentation
z-resolution: σz ~ 60 µm
Improved determination of track directions
Improved separation of different particle classes
Bergmann et al., Eur. Phys. J. C (2019) 79: 165. https://doi.org/10.1140/epjc/s10052-019-6673-z
Trajectory reconstruction precision <200 µm at 1 m distance
Notes: Here you can see the methodology applied to real-world data from a measurement with 120 GeV/c pions at SPS. The sensor used here was a 2 mm thick CdTe.
Analysing the cosmic muon track (measured in my office) which is fitted with a 3D line, we can estimate that the extrapolation error at a distance of 1 m amounts to below 200 µm.
The presented track reco does not only improve particle tracking and identification, it also provides means to utilize Timepix3 as a single-layer Compton camera (and polarimeter). This is currently used by several groups mainly in the medical sector and for environment monitoring.
Compton camera in particular in combination with high-Z sensors

## Slide 11
https://iopscience.iop.org/article/10.1088/1402-4896/ad10dd/meta
CZT from Kromek

## Slide 12: 5 mm CZT: Drift time and drift velocity vs electric field
μe = (944.8 ± 1.3) cm2/V/s
Slope is the drift velocity
Bias up to 1.8 kV

## Slide 13: 5 mm CZT: Determination of µτe and the τe
Iraddiate of the sensor with 60 keV gammas from 241Am (interacting shallowly in the sensor) from the backside at different bias voltage
Fit the 60 keV peak position vs bias in each pixel
μe = (944.8 ± 1.3) cm2/V/s
 τe = (1.38 ± 0.31) µs
<µτe> = 1.3 x 10-3 cm2/V

## Slide 14: 5 mm CZT: Charge collection efficiency and gamma spectrum measurement
Due to charge trapping energy the energy for photons aborbed at the common electrode (backside) is underestimated below 130 keV.
241Am source

## Slide 15: 3D reconstruction and gamma-spectroscopy: Application as a single-layer Compton Camera
5 mm CZT Timepix3 (110 µm pitch) irradiated with a 22Na source at a distance of 10.5 cm at different lateral positions (separated by 5 cm)
D. Turecek et al 2020 JINST 15 C01014

## Slide 16: 3D reconstruction: Single-layer Compton Camera - Experiment
5 mm CZT Timepix3 (110 µm pitch) irradiated with a 22Na at different lateral displacement.
Source position 1
Source position 2
10.5 cm
5 cm
P Smolyanskiy et al 2024 Phys. Scr. 99 015301
Measured detection efficiencies:
Energy (keV) | εoverall (%) | Εphoto (%)
511 | 24.4 ± 1.2 | 7.00 ± 0.50
1275 | 6.80 ± 0.4 | 0.53 ± 0.03

## Slide 18: Timepix3 with 500 µm thick GaAs Drift time analysis in negative polarity
B. Bergmann et al 2020 JINST 15 C03013
Maximal Drift time differences ~2 ns
 GaAs as sensor material is fast enough to fully utilize Timepix3 time resolution

## Slide 19: Timepix3 with 500 µm thick GaAs Drift time analysis in negative polarity
Ruch-Kino model:
Drift time vs
For
B. Bergmann et al 2020 JINST 15 C03013

## Slide 20: Timepix3 with 500 µm thick GaAs µτe measurement in negative polarity
Irradiation of the sensor with 15.8 keV XRF from Zr (shallow interaction depth)
Measure the peak position as a function of the bias

## Slide 21: Timepix3 with 500 µm thick GaAs Drift time analysis in negative polarity
Ruch-Kino model:
Drift time vs
For
B. Bergmann et al 2020 JINST 15 C03013

## Slide 22: Timepix3 with 500 µm thick GaAs µτe measurement in negative polarity
Irradiation of the sensor with 15.8 keV XRF from Zr (shallow interaction depth)
Measure the peak position as a function of the bias

## Slide 23: Timepix3 with 500 µm thick GaAs µτh measurement in positive polarity
Hole lifetime is too short to see the signal created far from the pixel side
 Lifetime determination through comparison with simulation (drift time, track length)

## Slide 24: Implementation into the simulation and comparison with measurement
Implementation of determined charge carrier properties into Allpix2
 Satisfactory agreement of energy and cluster size distribution

## Slide 25: Some words on silicon ...

## Slide 26: Timepix3 with a silicon sensor as Compton imager and polarimeter
Astrophysical gamma source emit in the energy range from 0.1 – 1 MeV
For polarimetry Compton scattering and subsequent photoabsorpion are needed
Silicon has better ration of Compton/Photo
Silicon has better spectroscopic performance

## Slide 27
120 GeV/c pion tracks accompanied by δ-rays:
A cosmic µ measured in the Prague laboratory:
z-resolution:
σz ~ 30 µm
Bergmann et al. Eur. Phys. J. C (2017) 77: 421. https://doi.org/10.1140/epjc/s10052-017-4993-4
Test measurements: 3D track reconstruction in silicon

## Slide 28: Timepix3 Compton polarimeter - principle
Differential cross section for Compton scattering off an electron is described by the Klein-Nishina formula:
Detected Compton-scatter pair
Ф is the angle between the electric field vector of the incoming photon and the scattering plane
(modulation)
Degree of linear polarization
T. Michel, J. Durst, J. Jakubek, “X-ray polarimetry by means of Compton scattering in the sensor of a hybrid photon counting pixel detector,” NIMA 603, Issue 3, pp 384-392 (2009) https://doi.org/10.1016/j.nima.2009.02.032
Notes: T. Michel, J. Durst, J. Jakubek, “X-ray polarimetry by means of Compton scattering in the sensor of a hybrid photon counting pixel detector,” NIMA 603, Issue 3, pp 384-392 (2009)
https://doi.org/10.1016/j.nima.2009.02.032

## Slide 29: Timepix3 Compton polarimeter
U=75 kV I=60 A
Detector was rotated around target-detector line to show that the cosine-squared curve is fixed in laboratory frame
Timepix3 with a 1 mm thick silicon sensor
B. Bergmann et al. “Results and Perspectives of Timepix Detectors in Space—From Radiation Monitoring in Low Earth Orbit to Astroparticle Physics.” Instruments 2024, 8, 17. https://doi.org/10.3390/instruments8010017

## Slide 30: Timepix3 Compton polarimeter: Figure of merit – MDP99%
Minimum detectable polarization at 99% confidence MDP99% can be described by (background neglected):
Modulation for 100% polarized incoming radiation = 92% (from simulation)
 To measure a polarization of 10% at a 99% confidence at least 200,000 Compton pairs need to be detected

## Slide 31: Conclusion
Pixelated high-Z sensor are available for hybridization with Timepix3/Timepix4
The sensor materials have been characterized to allow for proper modelling of the response in simulation
Choice of the sensor material strongly depends on application

## Slide 32: Thank you very much for your attention!
The research activities can be partly funded from the Czech Science Foundation Junior Star Grant with No. GM23-04869M.