## Slide 1
November 21, 2024
Ultra-Low Resource Time-to-Digital Converter for PET TOF Systems Achieving Sub-10 ps Resolution Using LUT-Based Counter on FPGA
Daehee Lee1, Ian Hwang2, Sun Il Kwon1
Department of Biomedical Engineering, UC Davis, California, USA
Baskin School of Engineering, UC Santa Cruz, California, USA
Applications Session
6th International Workshop on New Photon-Detectors

## Slide 2
1. Time-to-Digital Converter
LHC cern
IBM Quantum System One
LiDAR
Total-body TOF PET, uExplorer
Medical Field
Physics
Industry
Computing
Time-to-Digital Converters (TDCs)
are electronic circuits that digitize continuous time values into discrete digital values.
1

## Slide 3
1. Time-to-Digital Converter
TDC applications
LHC cern
IBM Quantum System One
LiDAR
Total-body TOF PET, uExplorer
Medical Field
Physics
Industry
Computing
Time-to-Digital Converters (TDCs)
are electronic circuits that digitize continuous time values into discrete digital values.
1

## Slide 4
1. Time-to-Digital Converter
NON-TOF PET
TOF PET
Coincidence Time Resolution (CTR)
Source
TLeft
TRight
LOR
Noise
TRight
TLeft
Scintillator
FWHM
D
*LOR: Line of Response
Sensor
*D: Diameter
Siemens Biograph mCT (PET/CT), 2019
Time-of-flight Positron Emission Tomography
2

## Slide 5
1. Time-to-Digital Converter
Application-specific Integrated Circuit (ASIC) vs. Field-programmable Gate Array (FPGA)
Feature | ASIC | FPGA
Purpose | Fixed function | Reconfigurable
Performance | High | Moderate
Power | Low | High
Flexibility | Low | High
Development cost | High | Low
Unit cost | Low | High
Time to market | Long | Short
Logic density | High | Moderate
Rui Machado et al., (2019)
FPGA-based TDC
Coarse Counter
Phased Clocks
Tapped Delay Line
Differential
Pulse shrinking
3
+

## Slide 6
1. Time-to-Digital Converter
Tapped Delay Line (TDL) with CARRY4
CLB N
TIMEIN
MUX
1
1 0
O0
CO0
MUX
1
1 0
O1
CO1
MUX
1
1 0
O2
CO2
MUX
1
1 0
O3
CO3
D Q
D Q
D Q
D Q
C0
C1
C2
C3
DFFs
CARRY4
CLB 0
TDL
CLB N-1
A CLB is crucial for fast signal propagation and TDL implementation.
C0 - C1 - C2 - C3 - C4 …
Thermometer codes
1 - 0 - 0 - 0 - 0 …
Clock
4

## Slide 7
1. Time-to-Digital Converter
C0 - C1 - C2 - C3 - C4 …
Thermometer codes
1 - 0 - 0 - 0 - 0 …
1 - 1 - 0 - 0 - 0 …
4
Tapped Delay Line (TDL) with CARRY4
CLB N
TIMEIN
MUX
1
1 0
O0
CO0
MUX
1
1 0
O1
CO1
MUX
1
1 0
O2
CO2
MUX
1
1 0
O3
CO3
D Q
D Q
D Q
D Q
C0
C1
C2
C3
DFFs
CARRY4
CLB 0
TDL
CLB N-1
Clock
A CLB is crucial for fast signal propagation and TDL implementation.

## Slide 8
1. Time-to-Digital Converter
C0 - C1 - C2 - C3 - C4 …
Thermometer codes
1 - 0 - 0 - 0 - 0 …
1 - 1 - 0 - 0 - 0 …
1 - 1 - 1 - 0 - 0 …
4
Tapped Delay Line (TDL) with CARRY4
CLB N
TIMEIN
MUX
1
1 0
O0
CO0
MUX
1
1 0
O1
CO1
MUX
1
1 0
O2
CO2
MUX
1
1 0
O3
CO3
D Q
D Q
D Q
D Q
C0
C1
C2
C3
DFFs
CARRY4
CLB 0
TDL
CLB N-1
Clock
A CLB is crucial for fast signal propagation and TDL implementation.

## Slide 9
1. Time-to-Digital Converter
Average timing resolution with CARRY4 is
~10 ps (Virtex-7 series, Xilinx).
C0 - C1 - C2 - C3 - C4 …
1 - 0 - 0 - 0 - 0 …
1 - 1 - 0 - 0 - 0 …
1 - 1 - 1 - 0 - 0 …
1 - 1 - 1 - 1 - 0 …
1 - 1 - 1 - 1 - 1 …
Thermometer codes
4
Tapped Delay Line (TDL) with CARRY4
CLB N
TIMEIN
MUX
1
1 0
O0
CO0
MUX
1
1 0
O1
CO1
MUX
1
1 0
O2
CO2
MUX
1
1 0
O3
CO3
D Q
D Q
D Q
D Q
C0
C1
C2
C3
DFFs
CARRY4
CLB 0
TDL
CLB N-1
Clock
A CLB is crucial for fast signal propagation and TDL implementation.

## Slide 10
1. Time-to-Digital Converter
Average timing resolution with CARRY4 is
~10 ps (Virtex-7 series, Xilinx).
C0 - C1 - C2 - C3 - C4 …
1 - 0 - 0 - 0 - 0 …
1 - 1 - 0 - 0 - 0 …
1 - 1 - 1 - 0 - 0 …
1 - 1 - 1 - 1 - 0 …
1 - 1 - 1 - 1 - 1 …
Thermometer codes
4
Tapped Delay Line (TDL) with CARRY4
CLB N
TIMEIN
MUX
1
1 0
O0
CO0
MUX
1
1 0
O1
CO1
MUX
1
1 0
O2
CO2
MUX
1
1 0
O3
CO3
D Q
D Q
D Q
D Q
C0
C1
C2
C3
DFFs
CARRY4
CLB 0
TDL
CLB N-1
Clock
Average timing resolution with CARRY8 is
~5 ps (Ultrascale series, Xilinx).
A CLB is crucial for fast signal propagation and TDL implementation.

## Slide 11
1. Time-to-Digital Converter
6
FPGA-based TDC
Coarse Counter
Phased Clocks
Tapped Delay Line
Differential
Pulse shrinking
Jun Yeon Won et al, 2016
+
Tapped Delay Line (TDL) with CARRY4
FPGA resources are limited.
 The smaller the resource usage, the more TDC implementations can be accommodated.
Ultra bins are unavoidable due to clock region changes.

## Slide 12
1. Time-to-Digital Converter
Ideal FWHM: 198 ps
Measured FWHM: 205 ps
5
Simulation
Average bin size 5 ps
Std 5 ps
Tapped Delay Line (TDL) with CARRY4
Simulation results
CARRY8-based TDC
CARRY4-based TDC

## Slide 13
2. Proposed LUT-based TDC
Proposed LUT-based TDC
7
FPGA-based TDC
Coarse Counter
Phased Clocks
Tapped Delay Line
Differential
Pulse shrinking
LUT-based Counter
Not Synchronized to the Clock
#1
#2
Clock Cycle
…
#1
#2
#5
#4
#3
#6
#7
…
LUT-based Counter
Clock-based Counter
Low resource is required for LUT-based counter
But the average bin size is ~250 ps
~250 ps

## Slide 14
2. Proposed LUT-based TDC
Proposed LUT-based TDC
7
FPGA-based TDC
Coarse Counter
Phased Clocks
Tapped Delay Line
Differential
Pulse shrinking
LUT-based Counter
Not Synchronized to the Clock
#1
#2
Clock Cycle
…
#1
#2
#5
#4
#3
#6
#7
…
LUT-based Counter
Clock-based Counter
Low resource is required for LUT-based counter
But the average bin size is ~250 ps  10 ps
~250 ps

## Slide 15
8
TimeIN
Time
Logic out
Clock
Time
IN logic
TimeIN
CLK
Logic
out
Schematic of the proposed LUT-based TDC
The input logic generates a pulse when the TimeIN is enqueued
and terminates the pulse at the next clock edge.
Time
2. Proposed LUT-based TDC

## Slide 16
8
TimeIN
Time
Logic out
Enable 1
Enable 2
Enable 3
Enable N
Enable N-1
Clock
Time
…
IN logic
AND
Enable
1
…
Delay
TimeIN
CLK
Logic
out
Enable
2
Enable
3
Enable
N
Time
Time
Schematic of the proposed LUT-based TDC
Time
Enable signals are gradually shorten
from Enable 1 to Enable N.
2. Proposed LUT-based TDC

## Slide 17
8
TimeIN
Time
Logic out
Enable 1
Enable 2
Enable 3
Enable N
Enable N-1
Clock
Time
…
IN logic
AND
Enable
1
…
Delay
TimeIN
CLK
Logic
out
CNT 1
CNT 2
CNT 3
CNT N
Enable
2
Enable
3
Enable
N
LUT-based Counter
2-bit: 0-1-2-3
LUT-based Counter Value
 #3
 #3
 #2
 #1
Sum(All LUT-based Counters)
1
2
3
1
2
3
1
2
1
CNT 1
CNT 2
CNT 3
CNT N-1
…
Bin size
Schematic of the proposed LUT-based TDC
2. Proposed LUT-based TDC

## Slide 18
8
TimeIN
Time
Logic out
Enable 1
Enable 2
Enable 3
Enable N
Enable N-1
Clock
Time
IN logic
AND
Enable
1
…
Delay
TimeIN
CLK
Logic
out
CNT 1
CNT 2
CNT 3
CNT N
Enable
2
Enable
3
Enable
N
LUT-based Counter
2-bit: 0-1-2-3
LUT-based Counter Value
 #2
 #2
 #1
Sum(All LUT-based Counters)
1
2
1
2
1
CNT 1
CNT 2
CNT 3
CNT N-1
…
Bin size
Schematic of the proposed LUT-based TDC
2. Proposed LUT-based TDC

## Slide 19
9
Schematic of the proposed LUT-based TDC
2. Proposed LUT-based TDC
IN logic
AND
…
CLK
CNT 1
CNT 2
CNT 3
CNT N
∑
IN logic
AND
…
Delay
TimeIN
CLK
Logic
out
CNT 1
CNT 2
CNT 3
CNT N
∑
…
Ch 1
Ch N
MEMORY
Chart: Summation results
 | 1 | 2 | 3 | 4 | 6 | 8 | 10 | 12 | 14 | 16 | 18 | 21 | 24 | 27 | 30 | 33 | 36 | 39 | 43 | 47 | 51

## Slide 20
Single channel test
3. Experiment Results
Average Bin Size: 9.4 ps
STD of Bins: 10 ps
Total active bins: 212
2000 ps
10
80%
90%
16ps
10ps

## Slide 21
TDC ch2
FGPA board
CLK
TDC ch1
TimeIN Ch1
TimeIN Ch2
𝛥t
𝛥t
Random input
Multi-channel Test
3. Experiment Results
12
-14 ps
-2 ps
9 ps
11 ps
Linearity Test Results

## Slide 22
TDC ch1
TimeIN Ch1
Random input
Temperature Test
Average bin size is stable.
 No temperature correction circuit, which requires a huge resource, is needed.
3. Experiment Results
FGPA board
CLK
13
Average bin sizes depending on temperature changes

## Slide 23
Location Test
TDC ch1
TDC ch1
FGPA board
CLK
TDC ch1
Random input
TimeIN Ch1
3. Experiment Results
14
Average bin size is stable.
 No location correction circuit, which requires a huge resource, is needed.
Average bin sizes depending on location changes

## Slide 24
TDC ch2
FGPA board
CLK
TDC ch1
TimeIN Ch1
TimeIN Ch2
Calculated FWHM
Osc. (sampling 40ps, interpolation) 245 ps
LUT-based TDC (10ps/10ps) 240 ps
CARRY8-based TDC (5ps/5ps) 241 ps
3. Experiment Results
Experimental Setup
22Na
15
NUV-MT (FBK)
LYSO
3 x 3 x 20 mm3
PZ-circuit
NUV-MT (FBK)
LYSO
3 x 3 x 20 mm3
PZ-circuit
Ave. / STD

## Slide 25
4. Conclusion
Utilization |  | CARRY8-based TDC | LUT-based TDC | Differences
 | CARRY8 | 102 | 2 | -100 (▼98%)
 | CLB LUTs | 837 | 404 | -432(▼52%)
 | CLB Registers | 1110 | 807 | -303(▼27%)
Resource use comparison
16
Power consumption comparison
Power consumption | CARRY8-based TDC | LUT-based TDC | Differences
Ch | 0.109 W | 0.006W | -0.103W(▼94%)
Reduced resource usage and lower power consumption enable the implementation of more TDC channels within a limited area.

## Slide 26
4. Conclusion
LUT-based Time-to-Digital Converter
is successfully implemented with dramatically reduced resources
17
Average bin size 9.4 ps  Sub-10 ps Bin Size
is internally free from location and temperature variation
The clock-asynchronized LUT-based TDC operates in a random sequence, minimizing the impact of temperature and location variations on performance.
As a result, there is no need to implement correction circuits, which would otherwise require a significantly large amount of resources.
Next step
The LUT-based TDC still exhibits large bins (>40ps).
 It will be investigated to make bin size more even

## Slide 27
Q & A
Thank you for your attention.
17

## Slide 28
3. Experiment Results
Differential Nonlinearity (DNL) and Integral Nonlinearity (INL)
11
Single channel test

## Slide 29
17
3. Experiment Results