## Slide 1: RFQ2 Bdg. 152 Measurements
L. Millar
SY-RF-MKS
RFQ2 Meeting – 22/04/2024

## Slide 2: Survey Measurements

## Slide 3: Survey Measurements
#1 | RUN3 supporting UV (b2002) | Radial X [µm] | Beam direction Y [mm] | Vertical Z [µm] | Roll angle [mrad] | Radial RMS [µm] | Vertical RMS [µm] | Scale factor BF estimation
 | T1_US | 1 | 0 | 78 | 0.69 | 17 | 16 | 1.000125
 | T1_DS | -41 | 1020.01 | -64 |  | 17 | 16 | 
 | T2_US | 23 | 1020.087 | -62 | 0.19 | 12 | 8 | 1.000079
 | T2_DS | 56 | 2046.087 | -61 |  | 12 | 8 | 
 | T3_US | -3 | 2046.265 | -52 | 0.24 | 23 | 19 | 1.000107
 | T3_DS | -21 | 3058.055 | 72 |  | 23 | 19 | 
#2 | RUN4 supporting UV (b152) | Radial X [µm] | Beam direction Y [mm] | Vertical Z [µm] | Roll angle [mrad] | Radial RMS [µm] | Vertical RMS [µm] | Scale factor BF estimation
 | T1_US | 2 | 0 | 56 | 0.49 | 10 | 10 | 0.999993
 | T1_DS | -35 | 1020.12 | -34 |  | 10 | 10 | 
 | T2_US | 27 | 1020.113 | -24 | 0.06 | 17 | 10 | 1.000010
 | T2_DS | 52 | 2046.179 | -31 |  | 17 | 10 | 
 | T3_US | 5 | 2046.218 | -23 | 0.10 | 23 | 13 | 0.999998
 | T3_DS | -31 | 3058.146 | 41 |  | 23 | 13 | 
#3 | RUN4 supporting AP (b152) | Radial X [µm] | Beam direction Y [mm] | Vertical Z [µm] | Roll angle [mrad] | Radial RMS [µm] | Vertical RMS [µm] | Scale factor BF estimation
 | T1_US | 2 | 0 | 54 | 0.51 | 17 | 16 | 1.000011
 | T1_DS | -42 | 1020.118 | -35 |  | 17 | 16 | 
 | T2_US | 22 | 1020.109 | -26 | 0.09 | 13 | 7 | 1.000000
 | T2_DS | 57 | 2046.178 | -36 |  | 13 | 7 | 
 | T3_US | -10 | 2046.236 | -16 | 0.19 | 26 | 20 | 0.999996
 | T3_DS | -24 | 3058.153 | 55 |  | 26 | 20 | 

## Slide 4: Survey Measurements

## Slide 5: Survey Measurements

## Slide 6: Frequency Evolution

## Slide 7: Notes
Cavity remeasured in air on 19/04/2024. Unfortunately, no RF measurements were performed before on 18/04/2024 immediately prior to venting.

## Slide 8: Frequency Correction
The following correction is applied to measurements in air:
The following is applied to measurements with the cavity under vacuum (εr =1):

## Slide 9: Frequency Correction
Thermal expansion coefficient for copper, αT, taken as 16.6E-6.
εr is calculated from the empirical relation (see notes for references and notation):
For a given temperature, TC, in °C, PS is obtained from the modified Arden Buck eq. as:
Notes: RH is the relative humidity (%).
P is the pressure of the air (mmHg) (assumed to be 760).
PS is the pressure of saturated water vapour (mmHg) at the temperature, T, in Kelvin.
TC is the temp in degrees Celsius.
REFS:
F. Cverna. ASM Ready Reference: Thermal Properties of Metals. ASM International, 2002. url: https://api.semanticscholar.org/CorpusID:107873170.
J. W. Broxon. ‘The Dielectric Constant of Air at High Pressures’. In: Physical Review 37 (1931), pp. 1338–1344. url: https : / / api . semanticscholar . org / CorpusID : 121254373.
A. G. Galka et al. ‘Microwave Cavity Sensor for Measurements of Air Humidity under Reduced Pressure’. In: Sensors 23.3 (2023). issn: 1424-8220. doi: 10 . 3390 /s23031498. url: https://www.mdpi.com/1424-8220/23/3/1498.
M. Zarnik and D. Belaviˇc. ‘An Experimental and Numerical Study of the Humidity Effect on the Stability of a Capacitive Ceramic Pressure Sensor’. In: Radioengineering 21 (Apr. 2012), pp. 201–206.
B. R. I. LLC. Model CR-1A Hygrometer with Autofill Operating Manual. http://www.hygrometers.com/wp- content/uploads/CR- 1A- users- manual- 2009- 12.pdf. Accessed: 2023-03-01.

## Slide 10: Frequency Evolution
Very strange result.

## Slide 11: Frequency Evolution Table
Ten of the last measurements shown (those on the previous frequency plot). Full spreadsheet updated regularly and available on CERNbox (also contains notes for each measurement, e.g., the status of the cavity, its support, vacuum system): https://cernbox.cern.ch/s/GH50jWEpJYe4MzL

## Slide 12: Frequency Evolution Plot
Raw data.
Corrected for cavity temperature.
Corrected for εr (applies to measurements in air).

## Slide 13: Temp/Humidity Evolution Plot

## Slide 14: Frequency Check
RFQ1 was measured at 22.8°C. Their target was 352.103 MHz at 26 °C.
In short, air accounts for ~100 kHz shift at out working frequency.
Page 177 of RFQ1 report.

## Slide 15: Q Evolution

## Slide 16: Q Evolution
Q0 = 6202
Qext = 4123
β = 1.5
Q0 = 6258
Qext = 4154
β = 1.51
Q0 = 6912
Qext = 4195
β = 1.65
No coupler installed; antenna-to-antenna measurements.
Q0 = 6952
Qext = 4168
β = 1.67
No change after venting.

## Slide 17: Antenna Measurements

## Slide 18: Antenna Signal Comparison
Changes from venting are smaller than those observed previously.

## Slide 19: Antenna Signal Comparison
Changes from venting are relatively small (<1%). Much smaller than those observed when moving the RFQ from building 2002 to building 152.
Calculated field change after venting the RFQ in building 152.
Chart: Longitudinal Position
q1 | 0.28824 | 0.35754 | 0.28824 | -0.77982
q2 | 0.62363 | 0.23052 | -0.17254 | 0.19591
q3 | 0.89044 | 0.17284 | 0.069101 | -0.10356
q4 | 0.21899 | 0.39221 | 0.24206 | 0.31133