## Slide 1: Fast readout system for beam signal digitalization in the FCC era
Thibaut Lefevre, Tom Levens, Stefano Mazzoni, Andreas Schlögelhofer
22.05.2025 / FCC Week 2025

## Slide 2: Table of contents
High bandwidth considerations
Fast readout using radio over fibre and time stretch
Test results from lab setup
Current LHC installation
Conclusion
Andreas Schlögelhofer
22.05.2025 / FCC Week 2025

## Slide 3: How to maintain a high bandwidth?
  Short propagation distance for high frequency signals leads to the requirement of short transmission lines (~ 2 dB/m @ 20 GHz)
    radiation hardness of high frequency components
  Maintaining high resolution at high bandwidth
    limited by ADC performance
    comes at a high price (8 GHz, 20 GSa/s, 10-bit, EUR ~100k)
What about resolution?
?
Can we avoid long distance RF transmission?
?
Can we overcome this limit?
22.05.2025 / FCC Week 2025
Andreas Schlögelhofer

## Slide 4: How to maintain a high bandwidth?
  Short propagation distance for high frequency signals leads to the requirement of short transmission lines (~ 2 dB/m @ 20 GHz)
    radiation hardness of high frequency components
  Maintaining high resolution at high bandwidth
    limited by ADC performance
    comes at a high price (8 GHz, 20 GSa/s, 10-bit, EUR ~100k)
What about resolution?
Radio-over-fibre Encoding and transporting RF signal using an optical carrier
?
Can we overcome this limit?
22.05.2025 / FCC Week 2025
Andreas Schlögelhofer

## Slide 5: How to maintain a high bandwidth?
  Short propagation distance for high frequency signals leads to the requirement of short transmission lines (~ 2 dB/m @ 20 GHz)
    radiation hardness of high frequency components
  Maintaining high resolution at high bandwidth
    limited by ADC performance
    comes at a high price (8 GHz, 20 GSa/s, 10-bit, EUR ~100k)
What about resolution?
Radio-over-fibre Encoding and transporting RF signal using an optical carrier
Time stretch Lowering ADC requirements by stretching time
22.05.2025 / FCC Week 2025
Andreas Schlögelhofer

## Slide 6: Radio-over-fibre with electro-optical modulators
Modulation due to Pockels effect
    linear variation of refractive index in response to an applied electric field
Electro-optic material
    Lithium niobate (LiNbO3)
    Gallium arsenide (GaAs)
    …
Interference-based modulation of light
    laser light split into two arms, modulated, and recombined
    designed for continuous wave lasers
    large variety of off-the-shelf products for telecommunication
22.05.2025 / FCC Week 2025
Andreas Schlögelhofer

## Slide 7: Spectral encoding
Use a chirped laser pulse instead of a continuous wave laser
    increase power density of the laser
Narrow optical spectrum
    keep reasonable performance of Mach-Zehnder interferometer
Encode the signals on the laser spectrum
    possibility to use laser spectrum also for decoding
22.05.2025 / FCC Week 2025
Andreas Schlögelhofer

## Slide 8: Time stretch through spectral encoding
Stretch the spectrum in time using dispersion in optical fibres
    moving away from real-time sampling
    opens up the possibility for lower bandwidth, continuous acquisition (RFSoC)
real time
stretched time
time stretch
linear dispersion in optical fibre
~ km
radio-over-fibre
22.05.2025 / FCC Week 2025
Andreas Schlögelhofer

## Slide 9: Test setup
9

## Slide 10: Test setup
Photodetector
DC bias voltage
EO modulator
stretched time
Read-out electronics
Laser and electronics
Modulator and signal source
RF signal
DUT 1
20 km
DUT 2
Femtosecond
laser
Chirped volume
bragg grating
1 ns
PM fibre
pulse generator
Stretching
5 km
5 ns
22.05.2025 / FCC Week 2025
Andreas Schlögelhofer

## Slide 11: Time conversion
22.05.2025 / FCC Week 2025
Andreas Schlögelhofer
time
spectrum
Δλ
Δλ
Δt
Δt
t
λ
-λ
t
Establish relationship between wavelength and time

## Slide 12: Time conversion
FCC-ee:
Bunch length (4σ): 25 to 200 ps
Bunch spacing: 25 ns
Factor 6 to 36 shorter bunch lengths in FCC-ee provide margin for much longer stretching  stretching factor in the order of 100 feasible
background
signal
 stretching factor 4.68 with test setup
Time delay of laser pulse with respect to input signal
Laser pulse spacing: 12.5 ns LHC:
Bunch length (4σ): 920 to 1200 ps
Bunch spacing: 25 ns
?
22.05.2025 / FCC Week 2025
Andreas Schlögelhofer
Stretched time, ns

## Slide 13: Stretching and bandwidth
background
signal
1.0 ns
14.7 ps (1 σ) ≈ 14.4 GHz
14.4 GHz / 4.68 ≈ 3.1 GHz  limited by the bandwidth of a 3 GHz oscilloscope
FCC-ee:
Stretching factor in the order of 100 realistic with a 25 ns bunch spacing.
500 MHz ADC would be sufficient for a 50 GHz equivalent signal digitalization
?
Current commercial Mach-Zehnder modulators have an analog bandwidth of up to 50 GHz
22.05.2025 / FCC Week 2025
Andreas Schlögelhofer

## Slide 14: LHC installation
14

## Slide 15: LHC installation
Photodetector
DC bias voltage
30 GHz EO-modulator
stretched time
Read-out electronics
BQK
Gallery (UA47/U0-001)
Surface (SX4 R-002)
Tunnel (2438/U0-2438)
+12.5 ns
BQK
RF signal
Attenuation (> 30 dB)
DUT 1
DCM 20 km
DUT 2
Femtosecond
laser
Chirped volume
bragg grating
DCM 5 km
PM fiber
1 ns (FWHM)
0.6 ns (FWHM)
1 ns (FWHM)
0.6 ns (FWHM)
4.7 ns (FWHM)
2.8 ns (FWHM)
200 fs (FWHM)
SM fiber
22.05.2025 / FCC Week 2025
Andreas Schlögelhofer

## Slide 16
Photodetector
DC bias voltage
30 GHz EO-modulator
stretched time
Read-out electronics
BQK
Gallery (UA47/U0-001)
Tunnel (2438/U0-2438)
+12.5 ns
BQK
RF signal
Attenuation (> 30 dB)
DUT 1
DCM 20 km
DUT 2
Femtosecond
laser
Chirped volume
bragg grating
DCM 5 km
PM fiber
1 ns (FWHM)
0.6 ns (FWHM)
1 ns (FWHM)
0.6 ns (FWHM)
4.7 ns (FWHM)
2.8 ns (FWHM)
200 fs (FWHM)
SM fiber
Main part of the setup on the surface is accessible 24/7
22.05.2025 / FCC Week 2025
Andreas Schlögelhofer

## Slide 17
DC bias voltage
30 GHz EO-modulator
BQK
Tunnel (2438/U0-2438)
+12.5 ns
BQK
RF signal
Attenuation (> 30 dB)
PM fiber
0.6 ns (FWHM)
1 ns (FWHM)
0.6 ns (FWHM)
Enclosure
Strip-line (BQK)
22.05.2025 / FCC Week 2025
Andreas Schlögelhofer

## Slide 18
DC bias voltage
30 GHz EO-modulator
BQK
Tunnel (2438/U0-2438)
+12.5 ns
BQK
RF signal
Attenuation (> 30 dB)
PM fiber
0.6 ns (FWHM)
1 ns (FWHM)
0.6 ns (FWHM)
  Being tested with beam at the moment!
LiNbO3
GaAs
Modulators
Splitter / Combiner
Attenuators
22.05.2025 / FCC Week 2025
Andreas Schlögelhofer

## Slide 19: Setup considerations
Laser pulses need to be synchronized to particle beams
    e.g. for protons in the LHC: 40.078 885 MHz to 40.078 970 MHz (0.45 Hz/s during ramp)
    Not a problem in FCC-ee main ring
Radiation hardness of various parts still to be assessed
    GaAs vs. LiNbO3 modulators
    Polarization maintaining fibers
Vertical resolution being tested with beam
    various detection schemes possible, in case the current setup underperforms in resolution: 1) optical background subtraction using a balanced photodetector 2) IQ modulator to encode difference and sum optically
22.05.2025 / FCC Week 2025
Andreas Schlögelhofer

## Slide 20: Conclusion
High-bandwidth, high-resolution acquisition is challenging
  demonstrated high linearity of time stretching
    demonstrated an analog bandwidth of 14.4 GHz, which can be easily adapted with different stretching stages
    current modulator technology limits bandwidth up to 50 GHz
    time stretching can make efficient use of the ‘void’ created through bunch spacing
Versatile setup
    most critical components accessible 24/7
    modifications for various detection schemes possible (balanced detection, IQ modulation, …)
  Big scale projects
    long transmission lines part of the design
    fibres have a much smaller footprint than high-frequency cables
First proton beam tests ongoing right now
22.05.2025 / FCC Week 2025
Andreas Schlögelhofer

## Slide 21: Thank you for your attention.
21