## Slide 1: Bead-Pull Measurement and Re-tuning of SNS RFQ02
Haitao Ren
On behalf of SNS RF Structure Team
Sep. 11, 2024
13th Continuous Wave and High Average Power RF Workshop, Knoxville, 2024

## Slide 2: SNS RFQs
 | RFQ01 | RFQ02 | RFQ03
Operation begins | 2002 | 2018 | 2023
RF Frequency | 402.5 MHz | 402.5 MHz | 402.5 MHz
Length | 3.7 m 4 segments | 3.7 m 4 segments | 3.7 m 4 segments
Material | Glidcop + Copper | Copper | Copper
Design transmission | 90% | 90% | 90%
Beam dynamic design | Original design | Same as RFQ01 | Same as RFQ01
Vane-vane volt | 83 kV | 83 kV | 83 kV
Kilpatrick | 1.85 | 1.85 | 1.85
Dipole mode suppression | π-mode stabilizers | 4-rod stabilizers | 4-rod stabilizers
Coupler | 8 coaxial couplers (later 2 coaxial couplers) | 2 coaxial couplers | 2 coaxial couplers
Tuners | 80 fixed slug tuners | 64 fixed slug tuners | 64 fixed slug tuners
Vacuum pumping | 6 cryo-pumps | 6 cryo-pumps + 4 turbo pumps | 10 turbo pumps
10 turbo pumps

## Slide 3: Original SNS RFQ01 (Berkeley RFQ)
RF frequency: 402.5 MHz
Design beam transmission: 90%
Peak surface field: 83 kV (1.85 Kilpatrick)
Dipole mode suppression: π-mode stabilizers
Resonance control: water temperature (two chillers: one for vanes and one for walls)
Glidcop/OFC braze joints

## Slide 4: RFQ02 (manufactured by RI)
Same beam dynamics design
Same RF input (2 coaxial couplers)
Same resonance control scheme
Design changes
  Solid copper RF structure
  (vs. Glidcop/OFC brazing)
  Improved vacuum pumping
  Improved water cooling
  4 rods for dipole stabilizer
  (vs. π-mode stabilizers)
  RF seal scheme
Broken piece of C-seal
Downstream end wall

## Slide 5: SNS New RFQ03 (manufactured by RI)
Design is the same as RFQ02 except RF seal
SNS Front End with the New RFQ03 (April 2023)
RFQ03 high power conditioning completed in Jan. 2023

## Slide 6: RFQ02 Needs Re-tuning
RFQ field tilt was noticed in early 2020
Required Field
Measured Field

## Slide 7: How Bead-pull Works
(z)
For Bead-pull measurement, RFQ H field is a function of the S21 phase
Bead-pull position
Small Perturbation Theory
  Relative magnetic field can be given by
(z) is bead-pull measured data

## Slide 8: Bead-pull Setup
Bead-pull plate with motor and pulleys
Laptop with MATLAB control program
Control box with Arduino
RFQ02
Network Analyzer

## Slide 9: RFQ Tuning Process
Untuned RFQ
CST Simulations
Quad field mode
Dipole field mode
Nonlinear check
Tuning Program
Algorithm
Excel tuning, or
Matlab tuning
Bead-pull Measurement
- Measured F
- Measured field profile
Run Tuning Program with Measured Data
- Output: tuners depth
Tune RFQ with Adjustable Tuners
RFQ tuned?
- Delta field <1%
- Delta F <50 KHz
Tuning Complete
Cut and Install Slug Tuners
- Section by section, or
- Quadrant by quadrant
RFQ tuned with all slug tuners?
No
Yes
No, next iteration
Yes

## Slide 10: RFQ02 Model and Tuning Algorithm
Q1
Q2
Q3
Q4
Q1
Q2
Q3
Q4
L+Δ
L-Δ
L
L
L
RFQ CST Model
Quad field simulation
Dipole field simulation
4 tuners move in or out at the same time
2 tuners in Q1 and Q3 move towards the same direction
(1 in , 1 out)
L
L
L
CST simulations:
- Field response / tuner
- Frequency response / tuner
Responses are scaled and added up for all tuners
Nonlinear responses to tuners are considered
To minimize
- Delta Frequency
- Delta fields

## Slide 11: Tuning Program – Excel (Quad Field Tuning)
Tuned Frequency
H field after tuning
H field before tuning
Field error after tuning
Excel Solver
Both Field & Frequency are tuned
Tune Quad field to flat
All 4 tuners (at the same longitudinal position Z) move in/out by same length together

## Slide 12: Tuning Program – Excel (Dipole Field Tuning)
Dipole tuning almost no impact on the frequency and quad field profile!
Only field tuning, No Frequency tuning
Tune Dipole Field to zero
2 tuners in opposite quadrants
  - One moves in
  - One moves out by the same length
Q1
Q2
Q3
Q4

## Slide 13: Tuning Program Improvements
Air/Vacuum Effect
  Relative permittivity: Air > Vacuum
  Frequency shift (CST) : = + 0.1187 MHz
Atmospheric Pressure Effect
  Frequency shift (CST): = - 0.0121 MHz
Nonlinear response of longer tuners
Noise check
  Noise level is small (< +/- 0.27%)
  Impact can be ignored
Chart: 
| 0 | 2.5 | 5 | 7.5 | 10 | 12.5 | 15 | 17.5 | 20
F_t1_CST | 400.386 | 400.4648 | 400.5369 | 400.6115 | 400.6746 | 400.7269 | 400.7782 | 400.8181 | 400.8526
F_t1_scale | 400.386 | 400.4581 | 400.5303 | 400.6024 | 400.6746 | 400.7467 | 400.8189 | 400.891 | 400.9632
Example: Tuner T1 nonlinear response
Chart: Noise Check

## Slide 14: Bead-pull Measurements – Before Tuning (Initial)
Chart: Measured Dipole Field - Before Tuning
Chart: SNS RFQ2 Measured H Field_before Tuning
Chart: SNS RFQ02 Measured Quad Field_before Tuning
Quad H = (H1 + H2 + H3 + H4)/4
D13 = (H1 - H3)/2; D14 = (H2 – H4)/2
Quad field tilt ~16.5% at entrance, ~9.0% at exit
Dipole field D24 ~7.0% (Max) and D13 ~2.5% (Max)
Target
Target
~16.5%
~9.0%
~6.9%
~7.0%
~2.0%
~2.5%

## Slide 15: Tuning Progress
Chart: SNS RFQ02 Measured H Field_6th Iteration
Chart: SNS RFQ02 Measured Quad Field_6th Iteration
Chart: SNS RFQ02 Measured Dipole Field_6th Iteration
Dipole field < 0.5%

## Slide 16: Frequency Tuning
Chart: RFQ02 Frequency Tuning
| 0 | 1 | 2 | 3 | 4 | 5 | 6
F_tuning | 402.5 | 402.39 | 402.405 | 402.385 | 402.395 | 402.395 | 402.39
F_target | 402.3934 | 402.3934 | 402.3934 | 402.3934 | 402.3934 | 402.3934 | 402.3934
20 kHz

## Slide 17: Final Measurement after 20 Slug Tuners Installed
Frequency: 402.395 MHz (Target 402.393 MHz)
  Measured F under vacuum: 402.525 MHz (7/27/2023)
Dipole field
  D13 (between Q1 and Q3) ~ 0.43% (Max)
  D24 (between Q2 and Q4) ~ 0.61% (Max)

## Slide 18: Tuner Depth Changes
Chart: Q2 tuners
| t1 | t2 | t3 | t4 | t5 | t6 | t7 | t8 | t9 | t10 | t11 | t12 | t13 | t14 | t15 | t16
L_Initial | 0.11909 | 0.15809 | 0.20709 | 0.23622 | 0.13809 | 0.20409
L_6th iter | -0.24902 | -0.029346 | 0.26438 | 0.28219 | 0.13841 | 0.14784
Chart: Q3 tuners
| t1 | t2 | t3 | t4 | t5 | t6 | t7 | t8 | t9 | t10 | t11 | t12 | t13 | t14 | t15 | t16
L_Initial | 0.35009 | 0.12909 | 0.056087 | 0.20609
L_6th iter | -0.2747 | 0.13012 | 0.13977 | 0.25586
Chart: Q4 tuners
| t1 | t2 | t3 | t4 | t5 | t6 | t7 | t8 | t9 | t10 | t11 | t12 | t13 | t14 | t15 | t16
L_Initial | 0.47009 | 0.38009 | 0.24009 | 0.21409 | 0.20009 | 0.22009
L_6th iter | 0.20516 | 0.56752 | 0.17009 | 0.18052 | 0.19977 | 0.27634
Module 1
Module 2
Module 3
Module 4
Module 1
Module 2
Module 3
Module 4
Module 1
Module 2
Module 3
Module 4
Q1
Q2
Q3
Q4
Chart: Q1 tuners
| t1 | t2 | t3 | t4 | t5 | t6 | t7 | t8 | t9 | t10 | t11 | t12 | t13 | t14 | t15 | t16
L_Initial | 0.26009 | 0.31309 | 0.37509 | 0.22209
L_6th iter | 0.25184 | 0.29934 | 0.30381 | 0.17231
Module 1
Module 2
Module 3
Module 4
RFQ field tilt mainly caused by detuning in Module 1

## Slide 19: Summary
Bead-pull system was developed for SNS RFQ
RFQ02 bead-pull retuning was successfully complete
Both field flatness and frequency met requirements after retuning
Thank you for your attention!
Questions or Comments?