## Slide 1: RFQ2 Discussion
L. Millar
SY-RF-MKS
RFQ2 Meeting – 16/05/2024

## Slide 2: Recap
Building 2002:
The RFQ was tuned.
Leak test (Ridge 2 was removed, vacuum port attachments added). The following gaskets were replaced: Tuners T2A2, T7, T4 A4, Antennas T1A1, T17A9, T20A12.
RFQ frequency checked with the antennas, a decrease of 33.41 kHz was observed (corresponds to ~5° in temperature or moving a single tuner by 1.76mm).
Building 152:
RFQ transported to building 152 and placed on the final support. No further frequency change since building 2002.
RFQ was moved to high-power test location. Pumps, the final ridge 2, and the window were installed. The cavity was placed under vacuum.
Measurements performed with WG-Coax coupler indicate that Q0 is higher, f0 is even lower, and the field profile is out of spec compared to when it was in bdg. 2002.
Survey measurements were performed (results shared previously).

## Slide 3: Since last time…
Building 152:
Simulations performed to determine the deformation which can be expected due to vacuum.
Flanges on vacuum ports removed – No change.
Endplate removed – No change.
Additional support installed in high-power test position.
RFQ moved to test area floor.
RFQ moved back to high-power test position.

## Slide 4: Effect of Vacuum

## Slide 5: Results of Deformation study
Preliminary mechanical simulations performed. There is a lot of information, to summarise:
Simulations indicate that when placed under vacuum a ~1µm level decrease in the distance between opposing vanes is present, peaking at ~3.8µm locally in the central segment (by design, the average distance between opposing vanes is 6.6mm).

## Slide 6: Frequency Evolution

## Slide 7: Frequency Evolution
Installing the blue support on the flange opposite the power coupler seems to have reduced the frequency by ~7 kHz.
Notes: Full spreadsheet updated regularly and available on CERNbox (also contains notes for each measurement, e.g., the status of the cavity, its support, vacuum system): https://cernbox.cern.ch/s/GH50jWEpJYe4MzL

## Slide 8: Field Profile Evolution (from bdg. 2002 to bdg. 152)

## Slide 9: Antenna Change - 2002 to 152
Change from building 2002 to building 152 (point 1 to 8 on slide 7). Note, gaskets were replaced on A1, A2, A4, A9, A12 (plotted with x). Throughout the manipulations, Q and f0 remained virtually unchanged.

## Slide 10: Field Profile - 2002 to 152
Building 2002, after final tuning.
Building 152, after transport and addition of final ridge 2, window, etc.

## Slide 11: Field Profile - 2002 to 152

## Slide 12: Field Profile Evolution (25/04/24 – 26/04/24)

## Slide 13: Antenna Change - adding sleeves on 25/04/24
Addition of sleeves on the top of the cavity 25/04 – 26/04 (point 12 to 14 on slide 7).
Conclusion: Changes from venting are small (~1% or less), much less than those observed when moving the RFQ from building 2002 to building 152. The added sleeves have a negligible effect.

## Slide 14: Calculated Field Profile – with sleeves on 25/04/24

## Slide 15: Field Profile Evolution (07/05/24 – 13/05/24)

## Slide 16: Antenna Change - 07/05/24 with support in test position
Effect of installation the blue central support (point 18 to 19 on slide 7).
Installed while cavity was still in high-power test area.

## Slide 17: Calculated Field Profile - 07/05/24 with support in test position

## Slide 18: Antenna Change - 07/05/24 with support on floor
Effect of moving the cavity to the floor (point 19 to 20 on slide 7).

## Slide 19: Calculated Field Profile - 07/05/24 with support on floor

## Slide 20: Antenna Change - 13/05/24 with support in test position
Effect of moving the cavity back to the test position (point 20 to 21 on slide 7).

## Slide 21: Calculated Field Profile - 13/05/24 with support in test position

## Slide 22: Summary (25/04/24 – 13/05/24)