## Slide 1: Introduction to CERN’s Beam Gas Ionisation (BGI) Profile Monitor
Visit to CNAO, 14th May 2024
James Storey, Clara Fleisig (CERN SY-BI-XEI)
Experimental Areas, Electron Beams & Ionisation Profile Monitors (XEI) Section,
Beam Instrumentation (BI) Group,
Accelerator Systems (SY) Department, CERN.

## Slide 2: Outline
[James]
Very brief introduction to Beam Instrumentation (BI) at CERN
Introduction to the Beam Gas Ionisation (BGI) beam monitor
[Clara]
3. Preliminary investigation of BGI for CNAO

## Slide 3: CERN Accelerator Complex

## Slide 4: Beam Instrumentation (BI) at CERN
Example beam observables:
Beam position;
Beam intensity & current;
Transverse & longitudinal beam size;
Emittance;
Luminosity;
Particle identification...
Instruments used to measure beam observables – the “eyes” of the accelerator.

## Slide 5: BI @ CERN in a nutshell
Interact with the electromagnetic field of the beam
  Used for measurements of beam position, intensity & current
  + (Almost) no effect on the beam & no effect on the monitor
  - Can’t measure beam profile (size)
How to measure the beam parameters?
Example – Beam position monitor (BPM)
Ref. https://videos.cern.ch/record/2697274
Electrostatic pickup consists of metallic (button) electrodes situated on opposite sides
As beam passes electric charges induced on electrodes, with more induced on the side closer to the beam
Horizontal (top) and vertical (bot.) position vs. LHC BPM position
LHC electrostatic button electrode

## Slide 6: BI @ CERN in a nutshell
Interact with the beam itself (insert material into the beam path)
  Used for measurements of beam profile (size) & loss
  + Allows to measure beam profile (size)
  - Destructive to the beam & damaging to the instrument
Example – Beam observation TV (BTV)
Ref. https://videos.cern.ch/record/1750709
Beam interacts with phosphor screen inserted into beam path
Emitted light passed though viewport & detected by camera
Direct measurement of beam profiles in the transverse (x,y) planes
How to measure the beam parameters?

## Slide 7: BI @ CERN in a nutshell
Interact with the beam itself (insert material into the beam path)
  Used for measurements of beam profile (size) & loss
  + Allows to measure beam profile (size)
  - Destructive to the beam & damaging to the instrument
Example – Beam Wire Scanner (BWS)
Ref. https://videos.cern.ch/record/1750709
How to measure the beam parameters?
Thin carbon wire (34μm) passes through the beam at up to 20m/s
Correlate wire position with secondary shower
Damage to the wire

## Slide 8: BI @ CERN in a nutshell
Use the synchrotron light emitted by the beam
  E.g. Transverse & longitudinal beam profile monitors
  + Completely non-destructive & continuous measurement
  - For proton accelerators can only be applied for energies above 100 GeV
  - Spatial resolution limited by diffraction & depth of field effects
How to measure the beam parameters?
Beam traverses superconducting undulator and dipole magnets – extract & detect visible range light produced by the undulator (450 GeV to 1.5 TeV) and dipole (>1.5 TeV)
Example – LHC synchrotron light monitor (BSRT)
LHC Beam 1 (left) and Beam 2 (right)

## Slide 9: Wishlist for an ideal beam profile monitor
Beam doesn’t damage the instrument & the instrument doesn’t perturb the beam
Continuous measurement throughout the acceleration cycle;
  Ideally every revolution of the accelerator, e.g. PS = 2.3 μs per turn.
Independent measurement with high accuracy & precision;
  Beam size = 270 μm in the LHC at 6.5 TeV.
Bunch-by-Bunch (BbB) and Turn-by-Turn (TbT) measurements
  e.g. LHC beam made up of 2556 bunches separated by 25 ns.
( Very promising ) solution:
Rest gas ionisation beam profile monitor based on Timepix3 hybrid pixel detectors.

## Slide 10: Rest gas ionisation beam profile monitor
Basic principle of operation:
Beam ionizes rest gas particles in the beam pipe vacuum
Transport ionization electrons with E-field and B-field
Image ionization electrons with a Timepix3 hybrid pixel detector
Original proposal for nondestructive beam profile monitor (1967)

## Slide 11: Beam Gas Ionisation (BGI) profile monitor for the CERN Proton Synchrotron (PS)
Requirements:
  Non-destructive;
  Continuous ( 1000 profiles / s );
  Bunch-by-Bunch & Turn-by-Turn;
  Measure beam size with 1% accuracy.
Typical use case e.g. LHC type beam (BCMS)
  Intensity = 13x1010 protons per bunch
  Number of bunches at extraction = 48
  Beam size = ~1mm
Development is part of the LHC Injector Upgrade (LIU) programme to deliver the high brightness beams required for the High Luminosity LHC.

## Slide 12: Instrument operating constraints
Operate directly inside the ultra-high vacuum of the accelerator beam pipe
  Vacuum pressure = 10-9 – 10-10 mbar
  Degassing < 5x10-6 mbar l s-1
  Criteria for the gas composition – concentration of:
    Gas species 18 – 44u < ( 1 / 100th ) of water (un-baked vacuum system)
    Gas species 44 – 100u < ( 1 / 1000th ) of water
Operate during the acceleration cycle
  EMI from the beam
  Beam impedance and beam induced heating
Operate in a radiation environment
  10 kGy/yr at beam pipe
  1 kGy/yr at 40cm

## Slide 13: Instrument overview
Measures beam profile in horizontal plane ( x, s ) – second instrument design for vertical plane measurement
ion trap
cathode (-20kV)
B-field (0.2T)
E-field (200kV/m)
vacuum flange
Faraday cage for electronics

## Slide 14: Instrument overview
Power, control & data cables
Ionisation electron detector

## Slide 15: Hybrid pixel detector
Amp.
Disc.
THR
Digital
Amp.
THR
Depletion
region
P+
P+
Aluminium
solder
bump
solder
bump
+
-
N+
N-bulk
sensor bias voltage
charged particle
sensor
readout
chip
Timepix3
Hybrid Pixel Detector (HPD): Pixelated sensor bump bonded to a pixelated readout chip.
Sensor: 2-dimensional array of PN-diodes processed (for example) in high-resistivity silicon.
Readout chip: array of readout channels designed in CMOS technology.

## Slide 16: Timepix3 response to charge
Amplifier out
Discriminator
640 MHz clock
40 MHz clock
Time-Over-Threshold
(ToT) clock
Threshold (THR)
Amplifier out > threshold
→ Event
time resolution = 1.5625 ns
Max.
80M events / s
Each event consist of:
Pixel position → Where
Time of Arrival (ToA) → When
Time-Over-Threshold (ToT) → Energy
minimum threshold = 500 e- (1.8 keV)

## Slide 17: Ionisation electron detector based on Timepix3 HPD
Sensor
Non-metalized, p-in-n, 100m deep
256 x 256 array of PN-diodes
Pixel size = 55um x 55um
Sensor area = 14mm x 14mm
Timepix3 readout chip
Each sensor pixel is connected to an individual readout channel (pixel)
Amp.
Disc.
THR
Digital
Timepix3
Amp.
THR
Depletion
region
P+
P+
solder
bump
solder
bump
N+
N-bulk
Cathode = -20kV
proton beam
100μm
10keV
E-drift field
Ionisation electron detector requirements
Detect 10keV electrons (penetration depth in silicon = 1.5μm)
Detect each electron with time resolution < 25ns & spatial resolution < 100μm
Meet outgassing requirements for installation in the UHV of the PS beam pipe
Operate during the acceleration cycle

## Slide 18: Transport of ionisation electron
Field cage transports ionisation electron from the point of creation to the detection plane, while preserving the initial transverse position of the electron:
Ideal
Complication 1 –
Bunch space charge
Complication 2 –
Secondary electrons

## Slide 19: Beam profile distortion due to bunch space charge
Ref. Sapinski et al., IBIC2012, TUPB61
Ref. https://accelconf.web.cern.ch/IBIC2012/papers/tupb61.pdf
Common interest among several accelerator labs (GSI, J-PARC, CERN, ESS, FNAL) to develop a common IPM simulation tool
Virtual-IPM (D.Vilsmeirer GSI):
Ionisation process
Electron / ion tracking in the EM-fields due to:
  Field cage
  Dipole magnet
  Bunch space charge
ipmsim.gitlab.io/Virtual-IPM
Solution = B-field parallel to the drift E-field
Strength of LHC BGI B-field (0.2T) insufficient for E > 4 TeV
Simulation tool to aid instrument design

## Slide 20: Suppression of secondary electrons
1. Wire grid – LHC BGI (2007)
2. Ion trap – PS BGI
Secondary electrons repelled by wire grid, but secondary electrons generated by ions bombarding the wires
Ions directed through an opening on the cathode – no path for secondary electrons back to the detector (”ion trap”)
Ref. https://accelconf.web.cern.ch/ipac2018/papers/wepal019.pdf

## Slide 21: Ionisation electron detector – Prototype
Ceramic carrier board
2 metal layers, Al2O3 substrate
4 x Timepix3 HPD’s attached with Staystick 672 and wire bonded
Sensor bias wire glued (Mk.1) / wire bonded to Al pad (Mk.2)
Flexible cables
Connects ceramic board to electrical feedthroughs on vacuum flange
Two metals layers with a Liquid Crystal Polymer (LCP) substrate
Qualified for installation in the PS accelerator beam pipe
Ref. Swann Levasseur CERN-THESIS-2019-354

## Slide 22: BGI based on Timepix3 UHV modules
Timepix3 UHV modules:
Timepix3 attached to aluminum base with Staystick (672)
Ceramic thin-film PCB to fan out TPX3 control & data signals to Panasonic narrow pitch connectors (A4S)
Connect to DC power (mini-Cu bus bars) via pogo-pins
Timepix3 DC power:
FEAST DC/DC on air side of vacuum flange ( can’t place in UHV )
Network of mini-Cu bus bars, embedded in ceramic, to minimise voltage drop
Final assembly

## Slide 23: Example beam images
TPX3 HPD #1
Beam direction
Example 1
Example 2
TPX3 HPD #2
TPX3 HPD #3
TPX3 HPD #4

## Slide 24: Selecting ionisation electrons
Before filtering
After filtering
Signal – ionisation electrons:
Mostly single pixel events
Energy < 10keV
Background – shower of secondary particles due to beam loss:
Multi-pixel events
Energy > 26keV
Signal selection:
Cluster finding to identify particle events
Size & energy criteria to select ionisation electrons
Ref. CERN-THESIS-2019-354 & CERN-THESIS-2020-236

## Slide 25: Preparation of LHC beam in the PS
Timepix3 data-driven readout enables "live" display of the beam throughout the cycle
1.5 seconds in real time: slowed down here for viewing purpose
Each frame is 10 ms of data
Not filtered: background particles are interesting to look at!
LHC type beam, single bunch ( I = 20x1010 p )

## Slide 26: Beam profile measurement
Compute beam profile by summing counts in each column

## Slide 27: Single bunch measurements
Single bunch in the PS at the start of beam commissioning
Beam size evolution
Beam position evolution
Continuous & non-destructive measurement of of beam size and position throughout the PS cycle (1.2s)

## Slide 28: Time structure of ionisation electrons
180ns
2.3μs
Turn 1
Turn 2
Turn 3
Turn 4
Turn 5
Turn 6
Turn 7
Turn 8
Turn 9
Time for 1 revolution (turn) of the PS = 2.3 μs.
→ Longitudinal bunch length consistent with Wall Current Monitor (WCM).
Ionisation counts from a single bunch circulating in the PS

## Slide 29: Beam position vs. turn number after injection

## Slide 30: Beam position vs. turn number after injection
Measure 0.217 ± 0.001 oscillations/turn → Consistent with PS fractional tune!
doi:10.18429/JACoW-IBIC2019-TUBO04

## Slide 31: Beam size vs. turn number after injection
Beating of the beam size due to injection mismatch into the PS
→ causes irreversible beam size growth, which ultimately limits the LHC luminosity.
doi:10.18429/JACoW-IBIC2019-TUBO04

## Slide 32: Examples of operational measurements
Following example measurements taken by PS operators & physicists in the CCC (*)
(*) Thanks to Marcel Coly, Matt Fraser & Alexander Huschauer
Operator defines the time to ”Start” the measurement in the cycle and the “Integration time” for each beam profile measurement, in units of seconds or turn revolutions
Up to 1024 profiles can be published per PS cycle (typically 1.2s)
Pb54+ beam

## Slide 33: Example 1: Beamlets in Multi-Turn Extraction beam
First time that is possible to see the time evolution of the beamlets
Extraction method developed to minimise particle loss during extraction
Beam is split into several beamlets in the transverse phase space
Each beamlet is then extracted one by one

## Slide 34: Example 2: Unexpected “bumps”
Horizontal beam size
Observe unexpected sudden change of the horizontal beam size at 1450ms
Explanation: sudden change of the beam’s longitudinal momentum spread, which couples to the horizontal plane through dispersion
New beam diagnostic insights.

## Slide 35: Example 3: Injection beam losses
Part of a cycle with a stable transverse beam size
“Glitch” visible at 1320ms in the cycle
Can distinguish between ionisation electrons and charged particle interaction using ToT data
“Glitch” is actually beam loss, which was subsequently traced to an injection kicker magnet that was active ( but no beam was injected )
Losses were not visible on other nearby beam loss monitors
Very sensitive beam loss monitor.

## Slide 36
Backup

## Slide 37: Ionisation electron signal
How to do turn-by-turn measurements?
Inject (very small) quantity of argon gas to locally bump pressure to 2x10-8 mbar
Argon ionisation cross section = ~5 x hydrogen (residual gas)
→ 1000 ionisation electrons per bunch per turn
detector length
protons(ions)
per bunch
vacuum
pressure
Residual gas ionisation electron signal:
Vacuum pressure = 1x10-9 mbar
Intensity = 60x1010 protons / bunch
→ 10 ionisation electrons per bunch per turn
How many ionisation electrons are needed?
1000 electrons → 2% precision (stat. error) on the beam size of a Gaussian distribution
Assume ~1000 electrons are needed for a meaningful beam profile
Residual gas insufficient for turn-by-turn measurements
ionisation electrons
ionisation electrons
- detected
electron detection efficiency
ionisation cross section

## Slide 38: Operation of the Timepix3 inside the beam pipe during the acceleration cycle
Acquisition for 100 ms during acceleration cycle with high voltage off (no ionisation electrons)
Time-Over-Threshold
( ∝ √Energy )
Timepix3 successfully operated inside the beam pipe during the acceleration cycle.

## Slide 39: Detection of rest gas ionisation electrons
proton
beam
sensor-to-sensor bias wires
ceramic carrier-to-sensor bias wires
honeycomb structured RF shield

## Slide 40: Detection of rest gas ionisation electrons
Ref. http://dx.doi.org/10.18429/JACoW-IBIC2017-WE2AB5

## Slide 41: Detection of rest gas ionisation electrons
10keV
Successful detection of rest-gas ionisation electrons!
Ref. doi:10.18429/JACoW-IBIC2017-WE2AB5

## Slide 42: Installation in the CERN PS
0.2T self-compensating triplet dipole magnet (Dominique Bodart TE-MSC) & instrument vacuum chamber
Installation at PS SS82
Vacuum pump down:
1 x 10-8 mbar after 24 hours
2 x 10-10 mbar steady state
Instrument prior to installation

## Slide 43: From prototype to operational instruments
BGI-Horizontal prototype
2017-2018
BGI-Horizontal operational
2021
BGI-Vertical operational
2021
Proved feasibility to operate TPX3 in the beam pipe
New modular in-vacuum electronics to improve Timepix3 cooling & production yield