## Slide 1: Medical-BGI Preliminary Investigations
May 14th, 2024
CNAO-CERN Collaboration Meeting
Clara Fleisig

## Slide 2: Tentative Focus on HEBT
Closer to end-point of beam
Lower vacuum improves signal per second (reduces required integration time)
A single BGI can validate beam quality for all three treatment rooms
Assumed Beam Parameters
Spills of 1 to 10 seconds
10e-6 to 10e-7 mbar vacuum
60-250 MeV proton beams and 120 to 400 MeV carbon-ion beams
Transverse profile with a maximum of 12 mm at full-width half-maximum
Transverse displacement of a few mm
*Measurements from Claudio Viviani with the SFP

## Slide 3: Timepix for the HEBT
Given beam parameter recommend field of view (FOV) > 40 mm
  Timepix3 width is 14 mm
  Timepix4 width is 30 mm
Recommend using two Timepix4 chips per instrument
Larger FOV reduces complications of multiple chips (threshold calibrations and accounting for overlap)
Longer chip length increases signal per spill by 75%
Newest version of chip means more likely to be supported for longer
Increased time resolution of ~200 ps (Timepix3 time resolution is 1.56 ns)

## Slide 4: Looking at Mechanics
Investigating use of a standard CF-flange
Size of BGI instruments vary depending on model
Rectangular Flange Model (four Timepix3s)
Model (without magnet) | Length along Beam Pipe | Height | Width | Depth
Rectangular flange | ~1 m | 215 mm | 283 mm | 435 mm
CF flange | > 350 mm | 130 mm | 310 mm | 310 mm
Circular Flange Model (one Timepix4)

## Slide 5: High Voltage Considerations
Increasing high voltage safety changes geometry of design
  Minimum 3 kV/mm in vacuum
  Minimum 1 kV/mm along insulator surfaces
  Minimum 4 kV/mm within insulators
Increasing high voltage increases signal
  Not all electrons have sufficient energy to penetrate silicon
  Higher voltage increases electron energies which increases the number of detected electrons
~9 keV Ionisation Electrons
~90 keV Ionisation Electrons

## Slide 6: Electric Field Uniformity
Clara Fleisig | Medical-BGI
Increasing electrode size increases electric field uniformity
More uniform electric fields reduce beam size error

## Slide 7: Electric Field Uniformity
Clara Fleisig | Medical-BGI
Increasing electrode size increases electric field uniformity
More uniform electric fields reduce beam size error

## Slide 8: Calculating the Expected Signal (Ionised Electrons)
Ionisation Cross-section (source):
Z: atomic number of beam particle
m: mass of beam particle
C: speed of light
: beam speed over the speed of light
: constant (dependent on rest gas)
C: constant (dependent on rest gas)
Number of Ionised Electrons:
L: detector length
N: protons per spill
P: pressure
K: Boltzmann constant
T: temperature
Need minimum ~1200 electrons per profile (corresponds to 2% beam size uncertainty)

## Slide 9: # of Expected Ionised Electrons
Assumed parameters
Detector length of 14.1 mm
23℃ gas temperature
Proton Beam
Best: 60 MeV, 1e09 protons/spill
Worst: 230 MeV, 1e08 protons/spill
Carbon-ion Beam
Best: 120 MeV, 1e08 ions/spill
Worst: 420 MeV, 1e07 ions/spill

## Slide 10: Investigating Magnetic Field
Simulation parameters
3 mm diameter proton beam
240 MeV energy
100,000 counts
150 mT Magnetic Field
0 mT Magnetic Field
Conclusion
|B| > 30 mT

## Slide 11: Next Steps
Narrow scope of the project
Investigate ion-based BGIs to reduce the need for a magnetic field
Converge on final mechanical design, high voltage, gas injection system, and magnet
Share conceptual design and get feedback on it,
Clara Fleisig | Medical-BGI

## Slide 12: Appendix
Clara Fleisig | Medical-BGI

## Slide 13: Beam Size Uncertainty
Beam size uncertainty comes from:
Systematic Uncertainty Sources | Other Uncertainty Sources
Electric field uniformity and strength Magnetic field uniformity and strength | Number of ionised electrons Noise (e.g. beamloss)
Clara Fleisig | Medical-BGI

## Slide 14: Complications from Non-Uniform Fields (neglecting space charge)
Solution to Lorentz equation for non-uniform electric field in “prime” frame of reference, where , from Bittencourt et al.
’
Rotate to BGI’s frame of reference to consider non-uniform magnetic field
ExB Drift
Clara Fleisig | Medical-BGI

## Slide 15: Understanding Profile Shift (neglecting space charge, omitting cyclotron motion)
ExB Drift
Accelerating drift
Clara Fleisig | Medical-BGI