## Slide 1: MBHDP301b 11T hybrid magnet Intermediate results - Thermal cycle n°1 & powering tests
S. Mugnier on behalf of Mechanical Measurement Lab of EN-MME
EDMS #2711698

## Slide 2: Instrumentation overview
SP301
12 instrumented collars
8 Bullet gauges
SP302
12 instrumented collars
8 Bullet gauges
12 Tie rods (End cage)
Instrumented SS Shell

## Slide 3: Instrumentation overview
SP301 without end-cage

## Slide 4: Instrumentation overview

## Slide 5: SP301 collars compression stress
Chart: MBHDP301b - Collars MBHSP301 - Compression Stress
| After SP301 NEW Collaring [RT / 17-01-2024] | After SS shell welding [RT / 29-02-2024] | Inside Cluster D cryostat [RT / 29-04-2024] | Cooldown n°1 [1,9K / 07-05-2024]
CCS1 | -98.6612 | -82.5864 | -81.9975 | -32.873
CCS2 | -106.3842 | -93.0062 | -92.1375 | -24.179
CCS3 | -79.8642 | -75.1043 | -74.3925 | -34.7374
CCS4 | -90.6633 | -83.1675 | -82.68 | -32.303
CM1 | -51.9422 | -47.2193 | -46.41 | 7.3899
CM2 | -78.7195 | -73.0763 | -72.345 | -40.06
CM3 | -49.1459 | -43.7775 | -42.8025 | -10.087
CM4 | -79.2675 | -71.37 | -70.2975 | -19.53
CNCS1 | -102.1254 | -83.7642 | -82.7775 | -22.1794
CNCS2 | -103.1404 | -81.51 | -80.8275 | -51.9322
CNCS3 | -73.0957 | -58.0027 | -57.2325 | -4.3277
CNCS4 | -95.4164 | -76.3815 | -75.66 | -26.1883

## Slide 6: SP302 collars compression stress
Chart: MBHDP301b - Collars MBHSP302 - Compression Stress
| After SS shell welding [RT / 29-02-2024] | Connection plate installation [RT / 16-04-2024] | Inside Cluster D cryostat [RT / 29-04-2024] | Cooldown n°1 [1,9K / 07-05-2024]
CCS1 | -40.7326 | -40.2675 | -40.6575 | -9.119
CCS2 | -67.9624 | -67.5675 | -67.86 | -30.8734
CCS3 | -36.0253 | -35.7825 | -35.9775 | 0.73281
CCS4 | -50.5898 | -50.4075 | -50.505 | -5.0425
CM1 | -70.6895 | -69.9075 | -70.59 | -43.4497
CM2 | -60.9736 | -60.84 | -61.1325 | -15.3014
CM3 | -83.8208 | -83.5575 | -83.85 | -38.5144
CM4 | -70.3433 | -70.005 | -70.785 | -25.2512
CNCS1 | -59.826 | -59.28 | -59.7675 | -16.3785
CNCS2 | -22.7955 | -21.84 | -23.01 | -4.5067
CNCS3 | -45.6398 | -45.045 | -45.9225 | -13.6689
CNCS4 | -30.4229 | -29.835 | -30.42 | 4.8445

## Slide 7: SP301 End plates longitudinal stress
Chart: MBHDP301b - End Plates bullets MBHSP302 - Longitudinal Stress
| Bullet gauges loading [RT / 28-02-2024] | After SS shell welding [RT / 29-02-2024] | Inside Cluster D cryostat [RT / 29-04-2024] | Cooldown n°1 (Before powering) [1,9K / 07-05-2024]
BCS1 | -32.06 | -32.76 | -31.59 | -76.53
BCS2 | -28.8 | -29.913 | -27.3 | -66.69
BCS3 | -29.8 | -31.317 | -28.08 | -63.83
BCS4 | -29.92 | -32.9745 | -29.64 | -60.17
BNCS1 | -36.32 | -38.532 | -38.61 | -72.58
BNCS2 | -37.5 | -39.8775 | -37.245 | -97.53
BNCS3 | -34.06 | -37.44 | -38.805 | -70.42
BNCS4 | -35.89 | -38.2395 | -39.585 | -77.12
Chart: MBHDP301b - End Plates bullets MBHSP301 - Longitudinal Stress
| Bullet gauges loading [RT / 28-02-2024] | After SS shell welding [RT / 29-02-2024] | Inside Cluster D cryostat [RT / 29-04-2024] | Cooldown n°1 (Before powering) [1,9K / 07-05-2024]
BCS5 | -31.31 | -30.8295 | -22.23 | -61.22
BCS6 | -30.5 | -31.941 | -28.665 | -31.23
BCS7 | -31.55 | -29.6205 | -28.86 | -96.87
BCS8 | -33.58 | -34.476 | -30.81 | -48.36
BNCS5 | -31.17 | -34.2615 | -33.54 | -18.73
BNCS6 | -31.63 | -35.607 | -34.905 | -64
BNCS7 | -31.35 | -30.3225 | -28.08 | -63.66
BNCS8 | -33.81 | -34.008 | -33.735 | -87.06

## Slide 8: SP302 End cage tie rods stress
Chart: Tie rods MBHSP302 (End cage) - Connection Side Longitudinal Stress
| After SS shell welding [RT / 29-02-2024] | Inside Cluster D cryostat [RT / 29-04-2024] | Cooldown n°1 (Before powering) [1,9K / 07-05-2024]
CS_Rods_213_U | 92.001 | 90.09 | 0
CS_Rods_213_M | 116.2395 | 113.685 | 136
CS_Rods_213_D | 99.0795 | 96.525 | 78.11
CS_Rods_212_U | 0 | 0 | 0
CS_Rods_212_M | 74.568 | 73.515 | 56.25
CS_Rods_212_D | 67.704 | 67.08 | 55.2
CS Average (Mpa) | 89.9184 | 88.179 | 81.39

## Slide 9: SP302 End cage tie rods stress
Chart: Tie rods MBHSP302 (End cage) - Non Connection Side Longitudinal Stress
| After SS shell welding [RT / 29-02-2024] | Inside Cluster D cryostat [RT / 29-04-2024] | Cooldown n°1 (Before powering) [1,9K / 07-05-2024]
NCS_Rods_213_U | 76.401 | 76.44 | 32.4
NCS_Rods_213_M | 97.9095 | 95.55 | 78.45
NCS_Rods_213_D | 48.3795 | 45.435 | 37.15
NCS_Rods_212_U | 56.8425 | 55.965 | 72.91
NCS_Rods_212_M | 63.8235 | 61.23 | 58.06
NCS_Rods_212_D | 62.244 | 58.695 | 35.05
NCS Average (Mpa) | 67.6 | 65.5525 | 52.3367

## Slide 10: MBHSP301 aperture

## Slide 11: SP301 Powering test
SM18 Event | 13/05/2024 at 13h21
Imax | 10,947 kA
Aperture powered | SP301
Magnet temperature | 1,9 K

## Slide 12: SP301 collars compression stress

## Slide 13: SP301 End plates bullets longitudinal stress & force

## Slide 14: MBHSP302 aperture

## Slide 15: SP302 Powering test
SM18 Event | 14/05/2024 at 16h30
Imax | 12,017 kA
Aperture powered | SP302
Magnet temperature | 1,9 K

## Slide 16: SP302 collars compression stress

## Slide 17: SP302 End plates bullets longitudinal stress & force
NCS group
CS group

## Slide 18: SP302 End cage tie rods (relative values)