## Slide 1
Medical applications in CLEAR
P. Korysko*, on behalf of the CLEAR team.
CLEAR Review 2024
May 29th 2024
*pierre.korysko@cern.ch

## Slide 2
Outline
2
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
Context: VHEE beams and the FLASH effect.
Tools and Methods used.
Selected Medical Applications performed at CLEAR in 2022/2023.
Medical Applications planned at CLEAR in 2024/2025.
Conclusions.

## Slide 3
Very High Energy Electrons

## Slide 4
Very High Energy Electrons
4
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
The potential use of very high-energy electron (VHEE) beams (50-250 MeV) for Radio Therapy (RT) recently gained interest, since electrons at these energies can travel deep into the patient.
Potential advantages of VHEE RT:
  Depth – dose profile for electrons better than X-rays.
  Charged particles can be focused and steered (not possible with X-rays) .
  Electron beams rather unsensitive to tissue inhomogeneities .
  Electron accelerators comparatively more compact, simpler and cheaper than proton/ion machines .
This last advantage is now especially true given the recent advancements on high-gradient acceleration (e.g. X-band CLIC technology) .
Ultra-high dose rate (above 100 Gy/s) radiation delivery, also known as FLASH RT, showed normal tissue sparing capabilities, without compromising tumor control. Most electron linacs can easily reach high beam currents needed for FLASH treatment of large fields.
More and more existing electron linac facilities are now being intensively used to investigate VHEE/FLASH RT.
VHEE/FLASH RT studies at the CLEAR facility (CERN)

## Slide 5
The FLASH Effect

## Slide 6
The FLASH Effect
6
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
The Flash effect is a biological effect that destroys cancerous cells while sparing healthy surrounding tissues.
Observed for the first time in 2014: mice tumors were irradiated with short pulses (≤500 ms) at Ultra High Dose Rate, UHDR (≥40 Gy/s).
The FLASH effect has been seen with protons, gamma and low energy electrons.
Very High Energy Electrons (VHEE) would be used to treat deep seated tumors.
The FLASH effect is extensively studied including in CLEAR.

## Slide 7
Treatment of a first patient with FLASH-radiotherapy
7
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
In 2019, 15 Gy delivered in 90 ms, using a 5.6-MeV electron linac, to a 75-years old patient with a multi-resistant cutaneous lymphoma:
  On healthy tissues: no decrease of the thickness of the epidermis and no disruption at the basal membrane with limited increase of the vascularization.
  On Tumor: Tumor response was rapid, complete, and durable with a short follow-up of 5 months.
Conclusions: This first FLASH-RT treatment was feasible and safe with a favorable outcome both on normal skin and the tumor.

## Slide 8
CERN/CHUV/THERYQ Collaboration
8
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
CHUV, CERN and now THERYQ are actively collaborating on the realization of a clinical facility for the treatment of large, deep-seated tumors by a VHEE beam in FLASH conditions.
The target is a facility that:
  Uses 100 MeV-range electrons and optimized dose delivery.
  Is compact enough to fit on a typical hospital campus.
The collaboration demonstrated that the facility is feasible and are now finalizing the details of its technical implementation aiming at first clinical tests in 1-2 years.
If successful, the facility may open the way for many future VHEE/FLASH facilities.
https://home.cern/news/news/knowledge-sharing/cern-chuv-and-theryq-join-forces-world-first-cancer-radiotherapy

## Slide 9
CLEAR Tools and Methods

## Slide 10
In order to facilitate the precise control of samples for multiple irradiations, the CLEAR-Robot (C-Robot) was designed and built by members of the CLEAR Operation Team.
It consists of 3 linear stages, 6 limit switches, a 3D-printed grabber, two water tanks and an Arduino board.
It has a precision in position in 3 axis of 50 µm.
It is fully remotely controllable from the CERN Technical Network.
Thanks to a mounted camera, it can also measure the beam sizes and transverse positions at the longitudinal position of the sample.
It is an open-source project: pictures, 3D renders, drawings and all the codes for the Arduino and the Graphical User Interface can be found on: https://pkorysko.web.cern.ch/C-Robot.html
Used for 100% of Medical Applications in CLEAR in 2023.
The C-Robot
10
Pierre Korysko - CLEAR Review 2024 - May 29, 2024

## Slide 11
11
The C-Robot 2.0
A new robot was built for the new CLEAR beamline:
  Mirrored, to adapt to the new in-air test area.
  51 available slots.
  With temperature probes, mounted camera, optical filters.
  This robot was sent to PITZ (DESY Berlin) for them to build a copy.
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
I. Najmudin

## Slide 12
12
The C-Robot in action with beam
Pierre Korysko - CLEAR Review 2024 - May 29, 2024

## Slide 13
Selected Medical Applications performed at CLEAR in 2023: Dosimetry Studies

## Slide 14
Graphical User Interface
Experiment setup w/ beam
e- beam
Camera
e- beam
Storage area
What can the C-Robot do?
14
Pierre Korysko - CLEAR Review 2024 - May 29, 2024

## Slide 15
Camera
e- beam
Real-Time Dosimetry
15
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
A real-time dosimetry measurement is done using the charge and the beam size measurement with the scintillating screen (in air or in water).
The samples are then irradiated at the same exact location.
A similar method is being developed using a thin scintillating screen in air in front of the water phantom for real-time dose measurement using charge density methods.
Gaussian Beam
Uniform Beam
e- beam
Camera

## Slide 16
Eppendorf tube with sample to irradiate
Beam water tank
Heated
storage
water tank
3D printed holder with 2 films: one before and one after the sample
Laser cut Radiochromic films to measure the delivered dose
Before irradiation
After irradiation
Dose range: 1 – 100 Gy
Experimental Setup & Dosimetry for VHEE at UHDR irradiations
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
16
V. Rieker
This is the passive standard way to measure the delivered dose, but real-time dosimetry is crucial for clinical treatments.

## Slide 17
Optical Fibre Dosimetry
Goal : Measure in real time the doses delivered by VHEE at UHDR and CDR with two arrays of optical fibers.
J. Bateman
17
Experiment : Recontruct the transverse profile of the VHEE beam to measure the dose in real time and compare with radiochromic films.
e-
e-
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
Published in Physics in Medicine & Biology

## Slide 18
Scintillator Dosimetry
Goal: Measure the dose at UHDR with a real-time readout and a high spatial resolution thanks to a scintillator and an optical fiber.
Experiment: Measure the responses of the scintillator for different doses and water depths and compare them with the doses measured on films.
A. Hart & C. Giguère
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
18
Published in IEEE Sensors Journal

## Slide 19
Selected Medical Applications performed at CLEAR in 2023: Irradiation Methods

## Slide 20
VHEE Strong Focusing
Goal: Focus the beam on the tumor in order to minimize the dose and damage on the nearby healthy tissues.
Experiment: Measure the beam sizes on a YAG screen in the water phantom (good model of the human body) and perform irradiations on long dosimetry films holders placed at different longitudinal positions.
L. Whitmore
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
20
Published in Nature Scientific Reports

## Slide 21
VHEE Scatterers
Goal: Obtain a flat beam that has a constant transverse distribution at patient’s tumor in order to minimize the dose and damage on the nearby healthy tissues.
Experiment: Measure beam profiles, sizes and intensity on a YAG screen and films after carefully inserting two scatterers with the beam with the C-Robot.
Scatterer 1
Scatterer 2
Screen
e-
X and Y beam profile
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
C. Robertson
21
YAG Screen
e-
Collimator
Paper under review
Now used by CLEAR Operation.

## Slide 22
VHEE GRID
Goal: Study the dose at UHDR for highly non-uniform dose distributions using a GRID Collimator (Spatially-fractionated RT, known for normal tissue sparing).
Experiment: Compare the dose values and profiles with and without the GRID collimator inserted for different water depths, with the YAG screen and films.
M. Bazalova-Carter, N. Clements, N. Esplen & A. Hart
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
e-
22
Published in Physics in Medicine & Biology

## Slide 23
Selected Medical Applications performed at CLEAR in 2023: Looking for the FLASH effect

## Slide 24
VHEE Chemistry Studies
Experiment: Measure and compare the production of Reactive Oxygen Species (ROS) in water at Conventional Dose Rate (CDR) and Ultra High Dose Rate (UHDR).
M-C. Vozenin & H. Kacem
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
24
UHDR=1.2 109 Gy/s CONV=0.15-0.41 Gy/s
Holder with films and Eppendorf tube
C-Robot view when performing irradiations for chemistry studies
Paper under review

## Slide 25
In vivo radiobiology at UHDR
Goal: Compare the impact of 200 MeV VHEE irradiations at UHDR and CDR on Drosophila melanogaster larvae.
Experiment: Deliver 15 to 45 Gy at UHDR and CDR to larvae with VHEE and measure the eclosion rate.
A. Hart & T. Esmangart de Bournonville
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
25
Not yet published
Preliminary results

## Slide 26
Biodosimeter Irradiations
Goal : Measure the response effect of the dose and the dose rate on biodosimeters with VHEE.
Experiment : Irradiate biodosimeters with numerous doses and dose rates: UHDR (Ultra High Dose Rate) and CDR (Conventional Dose Rate) and measure the length defficit.
M-C Vozenin & J. Ollivier
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
26
Paper under review
Preliminary results

## Slide 27
Selected Medical Applications planned at CLEAR in 2024:

## Slide 28
Selected Medical Applications in 2024
28
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
VHEE at UHDR Studies with Liposomes VHEE at UHDR Studies with Biodosimeters VHEE at UHDR Studies with Short Peptides & LCMS VHEE at UHDR Studies with Cells
VHEE at UHDR Studies with Drosophilae
& Gold Nanoparticles Plasmid Studies
VHEE Material Irradiation Studies
Radio-enhancement effect of
Nanoparticles in VHEE beams
Plan Delivery to an Anatomical Phantom
Goal: explore dose and dose rate parameters for both healthy and cancerous cells.

## Slide 29
Directions for VHEE/FLASH Experiments
29
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
Studying the impact of the beam time structure on the FLASH effect (average dose rate, instantaneous dose rate, etc.)
Studying the FLASH effect on both healthy and cancerous cells (first experiment done on cancerous cells in Nov. 2023).
Studying the FLASH effect on several biodosimeters (Zebra Fish Embryos & Drosophila Larvae).

## Slide 30
Comments from the CLEAR Scientfic Board
30
Pierre Korysko - CLEAR Review 2024 - May 29, 2024

## Slide 31
Conclusions
More and more users are studying the FLASH effect, Irradiation Methods and Dosimetry in CLEAR, leading to:
  14 weeks of beam dedicated to medical applications in 2023.
  More than 12 conference proceedings and 9 journal papers (published or being reviewed) for medical applications, see the full list on: https://clear.cern/content/publications
  13 Medical Application Experiments planned for 2024 (so far), see the full list on: https://pkorysko.web.cern.ch/CLEAR/Table/CLEAR_experiments_2024.html
  A new robot, the C-Robot 2.0 was built. 3 similar robots are being built in Germany (PITZ), Australia (Australian Synchrotron) and China (IHEP).
  New beam line with flexible optics, particularly suited for medical applications.
  New collaborations in 2024 with HUG and Gustave Roussy.
Pierre Korysko - CLEAR Review 2024 - May 29, 2024
31

## Slide 32
Thank you