## Slide 1
BTV NACONS
S.Burger
Introduction
Status/Needs
Consolidation
Upgrade: BTV230925
Specific thanks to
B.Moser / S.Girod / M.Lazzaroni / N.Charitonidis
F.Velotti / G.Lerhner / Y.Aguiar / F.Roncarolo / J.Tan
Beam Instrumentation for Fixed Target (BIFT) facilities working group - 1st meeting
2023_02_20

## Slide 2
INTRODUCTION (1)
S.Burger BIFT 2023_03_13
Layout
BA80
TCC2
No BTV
TDC2

## Slide 3
BTV | Line | Detector | Actuator type | Screen type |  | 
 |  |  |  | 1 | 2 | 3
21002 | LSS2 | Camera CCD | Maltese Cross – 4 positions | ALUMINE AF995 | SCINT. 31NSG | Ti 12um
210026 | TT20 | Camera Tube | Maltese Cross – 4 positions | ALUMINE AF995 | QUARTZ SIO2 | ALUMINE AF995
210272 | TT20 | PMT | Maltese Cross – 4 positions | ALUMINE AF995 | QUARTZ SIO2 | Ti 12um
210352 | TT20 | Camera Tube | Maltese Cross – 4 positions | ALUMINE AF995 | QUARTZ SIO2 | ALUMINE AF995
211628 | TDC2 | Camera Tube | Maltese Cross – 4 positions | ALUMINE AF995 | QUARTZ SIO2 | ALUMINE AF995
211743 | TDC2 | Camera Tube | Linear stepper – 2 positions | ALUMINE AF995 | - | -
220415 | TDC2 | Camera Tube | Maltese Cross – 4 positions | ALUMINE AF995 | QUARTZ SIO2 | ALUMINE AF995
230122 | TDC2 | Camera Tube | Maltese Cross – 4 positions | ALUMINE AF995 | SCINT. 31NSG | ALUMINE AF995
240119 | TDC2 | Camera Tube | Maltese Cross – 4 positions | ALUMINE AF995 | QUARTZ SIO2 | ALUMINE AF995
250118 | TDC2 | Camera Tube | Maltese Cross – 4 positions | ALUMINE AF995 | QUARTZ SIO2 | ALUMINE AF995
230925 | TDC2 | Camera Tube | Linear stepper – 2 positions | ALUMINE AF995 | - | -
BTV Hardware
INTRODUCTION (2)
S.Burger BIFT 2023_03_13
TDC2
TT20

## Slide 4
BTV210026
BTV210352
TDC2
TT20

## Slide 5
BTVs used ~regularly (every year at least at startup and debugging) but not much
During 2022, the multimages acquisition mode was tested on the BTVI application
Image check with low intensity beam over 4.8s Logbook images SPS (cern.ch)
no need to deal with saturation
Can rely on Chromox material decay time ~1.23ms (very little overlap of light in ‘slices’ of 20ms)
Intensity profile, beam size and position along spill could be observed
For higher intensity, saturation may appear, and beam size measurement will no longer be possible  Need filter wheel to mitigate this.
STATUS (1)
S.Burger BIFT 2023_03_13

## Slide 6
Dosimetry
 Very harsh environment !
Few kGy < Dose < few 10 KGy
Data from our yearly BTV radiation survey (RPL)
STATUS (2)
S.Burger BIFT 2023_03_13

## Slide 7
R2E expertise requested to better understand the radiation environment, checking:
the possibility to use other camera type than tube based ones
the possibility to put the cameras at safer locations
how shielding could help
 Monitoring of BTV locations in the North Area - TDC2 | Document 2753520 (v.0.5) (cern.ch)
 For most of the BTVs, optical line would not be feasible in this environment + integration not easy or feasible. As of today, no other choice than tube based camera
Environment TDC2 | TD [Gy] | HeH [cm-2]
Commercial camera | Should be x10 to 400 less | Should be x1k to 40k less
Rad tol. | Should be x3 to 100 less | Should be x3k to 1k less
Previous study (EDMS 2761320) has shown that an Iron wall of 40cm can reduce the TID levels by approximately a factor of 20, the HeH being reduced by a factor of 5 only
STATUS (3)
S.Burger BIFT 2023_03_13
Radiation Performances | TD [Gy] | HeH [cm2]
Commercial camera | 100 | 1e8
BeamCam (rad tol / development ongoing) | 250 | 1.5e9

## Slide 8
Summary of the BTV work to be done
For 8x BTVs: 210026, 210352, 216628, 216743, 220119, 230122 & 250118
  Replace screens (new markings)
  Check for rust on the vacuum chamber (vac advise)
     see with vac if a replacement is needed
     if mechanical changes/repairs are to be done, separate bellows from the instrument (interface modules to ease vacuum maintenance)
  New ‘Cloche’ (camera support) / illumination system
  New viewports (standardization)
  New interface box (standardization)
CONSOLIDATION (1)
S.Burger BIFT 2023_03_13

## Slide 9
Budget
For 8x BTVs: 210026, 210352, 216628, 216743, 220119, 230122 & 250118
CONSOLIDATION (2)
S.Burger BIFT 2023_03_13

## Slide 10
  Specifications:
    Screen size: Min 15x5cm
      to be able to see as much as possible of the wobbling (protons) + ion beam (only instrument there)
    Resolution <250um
     New vacuum chamber
        Viewport dimension needed CF250  ~197mm view diam.
        Keep or similar stepper insertion device
    digital camera i.e. optical line
BTV230925 UPGRADE (1)
S.Burger BIFT 2023_03_13

## Slide 11
220Gy
27.6KGy
1.2KGy
68.8KGy
1KGy
Proposal for camera location
BTV230925 UPGRADE (2)
S.Burger BIFT 2023_03_13
Looking for a ‘safe’ location for the camera
Also taking into account the feasibility of the installation of an optical line
 TBC, location could be in the TA801 tunnel

## Slide 12
BTV230925 UPGRADE (3)
S.Burger BIFT 2023_03_13
- Radiation survey
2m
4m
6m
8m
Distance from the door

## Slide 13
 | No shielding | HeH [cm-2]
 | TD [Gy] | *HeH [cm-2]
RPL1 | 120 | 1.2e11
RPL2 | 36 | 3.6e10
RPL3 | 16 | 1.6e10
RPL4 | 8 | 8e9
 With 40cm iron:
All locations are OK regarding TD
All locations are still not safe enough while using commercial digital camera
BTV230925 UPGRADE (4)
S.Burger BIFT 2023_03_13
13
Camera radiation limits
 | TD [Gy] | HeH [cm2]
Commercial camera | 100 | 1e8
BeamCam (rad tol / development ongoing) | 250 | 1.5e9
Iron Shielding 40cm | 
TD [Gy] | HeH [cm-2]
6 | 2.4e10
1.8 | 7.2e9
0.8 | 3.2e9
0.4 | 1.6e9
*HeH values estimation:1Gy  1e9HeH [cm2]
Radiation map in TA801
High-Energy Accelerator and Space Environments: Comparison of Radiation Effects on Electronic Components HMEE405 Document prepared by Arnaud MONTEUUIS Supervised by Ruben Garcia ALIA
/20
/5
2m
4m
6m
8m
Distance from the door
+ extra shielding
+ extra distance
 OK See slide annexe

## Slide 14
BTV.230925
1.8 m
0.7 m
To TA801 door
13.8 m
Beam ht 1.2 m
Relocation of the BTV230925:
Leave some room for the future BSG
Reduce the optical distance
Reduce the number of mirrors (maintenance)

## Slide 15: TDC2/TCC2 frontier
(13.8 m)
Ht 1.9 m
Relocation of the pipes on the wall
Drill hole through the door

## Slide 16: TA 801
BTV camera + shielding
1.9 m
Relocation of the pipes on the wall
Drill hole through the door
Excavate the round shape tunnel for a vertical/flat location for shielding and camera/optics
10/12 m

## Slide 17
Check resolution @25m for BTV230925 optical line
3.4lp/mm (150um) 22/110
 Visibility 20%
2.4lp/mm (200um) 33/110
 Visibility 30%
Optical resolution
Spatial resolution
18mm for 176px  ~100um/px
18mm
Camera lens: F600mm + focal doubler  F1200mm
30%
20%
100%
BTV230925 UPGRADE (5)
S.Burger BIFT 2023_03_13

## Slide 18
Summary
  New BTV location  straight in front of TA801 door
  New vacuum chamber
    separated bellows from the instrument (interface modules to ease vacuum maintenance)
    New screen, viewport
  New optical line (mirrors, hole in door, shielding, etc…)
  Use Digital camera @25m + shielding
  Cabling
BTV230925 UPGRADE (6)
S.Burger BIFT 2023_03_13
18

## Slide 19
Budget
BTV230925 UPGRADE (7)
S.Burger BIFT 2023_03_13
19

## Slide 20
CONCLUSION
S.Burger BIFT 2023_03_13
20
Information collected about radiation was mandatory !
Work on the BTVs was identified during 2022
 Consolidation program + budget is proposed for 8 BTVs
Upgrade was studied for the BTV230925:
  Need of a new tank
  Possibility to use digital system in TA801
    Integration of an optical line and shielding
EVM should be reviewed
(total budget stays of the same order)
THANKS

## Slide 21
ANNEXE

## Slide 22
 New vacuum chamber + new design needed with 250mm long stroke (not existing in BI standard
Need
 This is too demanding …
 Does it mean that we can remove the present BTV and install a new (simpler) setup is air?
 i.e. 10x10px/mm2
BTV230925 Upgrade first tentative specs (Nikos)

## Slide 23
High-Energy Accelerator and Space Environments: Comparison of Radiation Effects on Electronic Components HMEE405 Document prepared by Arnaud MONTEUUIS Supervised by Ruben Garcia ALIA
https://edms.cern.ch/ui/file/2761320/1/Monteuuis_2017_docx_cpdf.pdf

## Slide 24
 | With shielding 40cm |  | With shielding 60cm | 
Distance [m] | TD [Gy] | *HeH [cm-2] | TD [Gy] | *HeH [cm-2]
2 | 120 | 1.20E+11 | 12 | 4.80E+10
4 | 36 | 3.60E+10 | 3.6 | 1.44E+10
6 | 16 | 1.60E+10 | 1.6 | 6.40E+09
8 | 8 | 8.00E+09 | 0.8 | 3.20E+09
10 | 4 | 4.00E+09 | 0.4 | 1.60E+09
12 | 2 | 2.00E+09 | 0.2 | 8.00E+08
14 | 1 | 1.00E+09 | 0.1 | 4.00E+08
16 | 0.5 | 5.00E+08 | 0.05 | 2.00E+08
Estimation distance and shielding needed to be below HeH 2e8 [cm-2]