## Slide 1: WP4-WP9 Crab Cavities Experience from RFD proto CM test at CERN SM18 – M7 bunker
RFD proto
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Thursday, July 4th, 2024 - on Indico: 1426215

## Slide 2
2023 highlights
Focus on instrumentation aspects
Focus on process aspects
NC relief valve
Next steps
Introduction
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Thursday, July 4th, 2024 - on Indico: 1426215

## Slide 3
2023 HIGHLIGHTS
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Thursday, July 4th, 2024 - on Indico: 1426215

## Slide 4: 2023 HIGHLIGHTS
Thursday, July 4th, 2024 - on Indico: 1426215
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Dec’22
Jan’23
Feb
Mar
Apr
May
Jun
Jul
Aug
Sep
Oct
Nov
Dec
Jan’24
RFD proto
SM18 M7 bunker preparation for commissioning @ cold
RFD proto delivered @ CERN
Installation, SVs NCR treatment, commissioning
4.5 K
24.11.2023
2 K
28.11.2023
Instrum installation phase 1 / UK

## Slide 5
INSTRUMENTATION
(feedback CRG-IC/ T. Feniet / N. Vauthier)
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Thursday, July 4th, 2024 - on Indico: 1426215

## Slide 6
Organization
  Instrumentation shipping to optimize: the cryomodule builder should receive only the instrumentation he will install => clarification of the installation responsibilities (+ at which phase)
Instrumentation mounting, training on site
  Very positive and productive, to be re-conducted – CRG-IC, in charge, should have 1 direct contact with the cryomodule builder (no intermediate with EN-MME, to facilitate the communication), ideally the people in charge of the instrum installation on site
  Need for CRG-IC to know the profile of the people in charge on site, to better prepare their training & qualification
Installation on site by STFC
  Quality assurance
    cables labels were removed by STFC without CERN approval
    Pinning strategy not 100% respected by STFC
    Some TTs were not installed by STFC => had to be done by CERN once delivered
  Tuners: is there a need for sockets on TT/EH 861/62, to allow their dismounting without removing the sensors?
3D integration model & drawings
  To be included in the model:
    the sensor wire + its bending radius, the feedthroughs + the connectors, the cabling routing inside & outside the vacuum vessel, the electrical connecting box
  Quality assurance: actual drilling diameter for the orifices of the Acim Jouanin cartridges was not always consistent with the drawings
Instrumentation aspects (1/2)
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Thursday, July 4th, 2024 - on Indico: 1426215

## Slide 7
Necessary pending reviews
  After cold tests: are all the flange heaters necessary? => check their operation (if any), room for simplification
  Operation: TT847a not operational, to be checked with IC
  P&ID:
    TAGs update (TT819a/b to be replaced by 1TT819/2TT819)
    Add thermocouples on the EH821 & 822 foil heaters (TT827 & 828)
    Include all existing and physically installed EH on the P&ID (FPC)
    EH876/877/878/879 (flexible band heaters): they are all equipped with 2 heaters and 2 PT100 => update P&ID accordingly
    RFD series cryomodules: TT842 to be removed, TT814 & 816 to be added
    Safety port: import the DQW drawing into the RFD P&ID (1 flange for the chimney, 1 flange for the burst disc)
FTEs for M7 operation
  Reception + installation + commissioning: 2 FTEs for 2 weeks, continuous
  Dismounting: 2 FTEs for 3 days, continuous
Cryomodules reception at CERN
  Schedule already available?
  Quid storage conditions (for the instrum) after M7 dismounting and before LHC installation? Storage duration and conditions (where)?
Instrumentation aspects (2/2)
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Thursday, July 4th, 2024 - on Indico: 1426215

## Slide 8
CRYO PROCESS
(feedback CRG-ML / R. mauny)
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Thursday, July 4th, 2024 - on Indico: 1426215

## Slide 9
Cryo process aspects – intro & overview
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
2 main documents used to write the logic specification :
  «Functionnal description of operation modes» [HiLumi-LHC-CC-Cryo-N-39; EDMS 2647024]
  Engineering specification HL-LHC CC : cryomodules for crab cavities [LHC-ACF_A-ES-0001; EDMS 2043014]
3 main phases for cool down :
  Cool down from 300 K to 130 K: with controlled inlet temperature ; dTmax = 50 K.
  Cool down from 130 K to 4.5 K: by direct LHe injection into CCCM ; no limitation with dTmax below TTmax = 130 K; cooldown as fast as possible (AFAP).
  During these 2 phases, specific attention needed to CCCM pressure (< 1.5 bar)
  Pumping phase from 4.5 K to 2 K: beam screen (BS) circuit put in service, then pumping.
  ------------------------------------------------------------------------------------------------------------------------
1st RFD proto operation : 21 Nov. – 10 Dec 2023 (Logbook event 117205)
2nd RFD proto operation : 11 April – 14 June 2024 (Logbook event 119101)
Thursday, July 4th, 2024 - on Indico: 1426215

## Slide 10
Cryo process aspects – PFD1- CD 300K-130K – dTmax 50K
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Step 120 : CoolDown 300-130K with mixing chamber
Thursday, July 4th, 2024 - on Indico: 1426215

## Slide 11
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Cryo process aspects – PFD2- CD 130K- 4.5K – AFAP
Step 130 : CoolDown 130 K- 4.5 K - As fast as possible
Thursday, July 4th, 2024 - on Indico: 1426215

## Slide 12
Process aspects – Cool Down 300 K -> 4.5 K
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Cool Down time (300 K -> 130 K – dT 50 K) | 46 h
Cool Down time (130 K -> 4.5 K - AFAP) | 2 h
Flow during CD in cryomodule | ~2.2 – 3 g/s
Flow during CD in TS | ~1.3 g/s
Filling | 6 h
Flow during filling | ~ 5.2 g/s
Flow to maintain cav filled at 4.5 K | ~ 4 g/s
Tot. CD time to 4.5 K | 54 h
Constraint : keep Pressure < 1.4 bar
CD AFAP
CD with mixing chamber – dT 50K
Filling
Thursday, July 4th, 2024 - on Indico: 1426215

## Slide 13
Cryo process aspects – PFD 3 - Nominal operation @ 2 K
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Thursday, July 4th, 2024 - on Indico: 1426215

## Slide 14
Process aspects – Pumping down to 2 K
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Pumping time to 30 mbar | 2.5 h
Min LHe level during pumping step | 60 %
Flow during pumping step | ~5 g/s
Refill time at 30 mbar | 40 min
Flow during refill step | ~8.5 g/s
Total Time for CD 300 K -> Cryo_Ok (without interruption) | 58 h (2.5 days)
1.05 bar
30 mbar
Lhe level decreases during pumping phase
refill
Thursday, July 4th, 2024 - on Indico: 1426215

## Slide 15
Process aspects – Nominal operation @ 2 K
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Flow during nominal operation @ 2 K : 3.5 g/s [Eq. Heat load: 70 W]
Flow through Thermal Screen + Beam screen : 2.4 g/s + 0.3 g/s
Limitation of instabilities during RF conditioning :
LHe level not easy to maintain with dynamic heat load. It can drop suddenly due to geometrical aspects (chimney)
  Setting level controller with reactive PID parameters.
Adaptation of Cryo_Ok logic to gain margin
  Bigger margin with LT threshold (Low level set at 80% instead of 88% before)
  Pressure SetPoint lowered from 30 to 20 mbar
Lhe buffer too thin

## Slide 16
What went well
  cooldown & warm-up
Engaged manpower / resources, time spent for preparation
  3 days for cool down; 3 days for warmup
What could have gone better
  Bigger LHe buffer on top of the cavities
    => to have more operational margin on liquid helium level regulation
Process aspects - conclusion
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Thursday, July 4th, 2024 - on Indico: 1426215

## Slide 17
NC RELIEF VALVE
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Thursday, July 4th, 2024 - on Indico: 1426215

## Slide 18
Valves found leaky on Nov. 10th’24
  Reference: EDMS document 2998467 v0.1
  HSE test outcome: all 4 valves leaky between 160 mbarg and 220 mbarg (for a set pressure of 350 mbarg)
Mitigation for M7 operation in 2023-2024
  Removal of the PRV HeGuard
  Use of 1 single relief valve (Circle Seal) with set pressure of 700 mbarg
Retained setup
NC relief valve
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Thursday, July 4th, 2024 - on Indico: 1426215
References of the Circle Seal relief valves used for the RFD prototype cooldown and warmup in M7

## Slide 19
NEXT STEPS
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Thursday, July 4th, 2024 - on Indico: 1426215

## Slide 20
Complete the tests of RFD prototype in SM18
Review the series P&IDs (+ update instrum list)
Solve every non-conformity
  CRG side: replace the currently installed pressure relief valve and re-install the associated HeGuard
Get prepared for SPS BA6 operation of the RFD prototype
  Installation & Commissioning with cryogenic test facility at SPS BA6
Next steps
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Thursday, July 4th, 2024 - on Indico: 1426215

## Slide 21
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Thursday, July 4th, 2024 - on Indico: 1426215
THANK YOU

## Slide 22
HL-LHC – Crab Cavities RFD Proto CM test @ CERN – TE-CRG
Thursday, July 4th, 2024 - on Indico: 1426215
SPARES

## Slide 23: 10 series cryomodules
RFD
1 from CERN (CH)
DQW
5 from TRIUMF (CA)
4 from STFC (UK)
2 prototypes
DQW *
1 from CERN (CH)
RFD **
1 from STFC (UK)
* Currently operated in SPS BA6 @ CERN
** Being manufactured at STFC
LHC
SPS
DQW: double quarter wave
RFD: RF dipole
Run 4 configuration: RFD in P1, DQW in P5
Later on: possibility to swap them
DQW being tested at cold in SPS BA6
Later on: RFD to be installed / tested at cold in SPS BA6