## Slide 1: Characterization and Performance Evaluation of ITk Pixel Production Modules for the ATLAS Upgrade
Md Arif Abdulla Samy, Craig Buttar, Richard Bates, Dzmitry Maneuski, Liam Cunningham, Ruirui Han, Sneha Naik, Kamil Dulak, Fiona Grant, Calum Gray, Frederick Doherty, Christopher Lickley, Jon Taylor, Vedika S. Kumar

## Slide 2
Large Hadron Collider (LHC)
High-luminosity large hadron collider (HL-LHC) :
integrated luminosity: more than 250 fb-1/year
instantaneous luminosity expected to be 75 x 1033 cm2 s-1
Source:https://home.cern/resources/faqs/high-luminosity-lhc / NIMA 1070 (2025) 169978
50 pile-up events
200 pile-up events
Last update: Oct 2024
We are here
https://atlas.web.cern.ch/Atlas/GROUPS/DATAPREPARATION/PublicPlots/2024/
DataSummary/figs/peakLumiByFill.png
https://hilumilhc.web.cern.ch/sites/default/files/2024-10/HL-LHC_October2024.pdf

## Slide 3
ATLAS update for HL-LHC
Inner Detector (ID)
Max Acceptance |ɳ| = 2.5
92 million
Tracker: 6
Pixel Detector: channels
Semiconductor million channels
Transition
Radiation
Tracker: 350.000 channels
The current innermost detection system of ATLAS is the Inner Detector
Present layout
Inner Tracker (ITk)
Max Acceptance |ɳ| = 4.0
ITk Pixel Detector: 5 billion channels
ITk Strip Detector: 60 million channels
3
HL-LHC
era
Future layout
https://atlas.cern/Updates/Fe ature/High-Luminosity- ATLAS
https://atlas.cern/Discover/Detector
https://atlas.cern/Discover/Detector/Inner-Detector
ID will be replaced by the all-Si tracker ITk
Notes: The ITk Pixel detector has been designed to replace the Inner Detector of ATLAS for operating in the hasher environment of the high luminosity era of the LHC:
  Higher granularity
  Better radiation hardness
  Higher acceptance

## Slide 4
Inner Tracker-ITk
Closest to the interaction point
At least 9 hits/track
Occupancy kept < 1% thanks to the higher granularity
Composed of two parts: ITk Pixel with 5 pixel barrels and rings and ITk Strip with 4 strip barrels and strip disks
Replaceable two innermost layers at half of the lifetime: ITk Pixel
Inner System (IS) with Endcaps and Barrels
For the ITk Strip, see Katie W. Pass’ talk
ATL-PHYS-PUB-2021-024
ATL-PHYS-PUB-2021-024
For Pixel overview- see the talk of Umberto M.
For Strip overview- see the talk of Marta B.

## Slide 5
Pixel Module Sites
For the ITk Strip, see Katie W. Pass’ talk
10 clusters for pixel module assembly (31 assembly sites) and testing across the globe
United Kingdom Pixels is one of the clusters
UK cluster is responsible for assembly endcap quad pixel modules with planar sensors, QA/QC them, load them in half ring and deliver them to CERN
UK cluster has three assembly and testing sites:
University of Glasgow
University of Liverpool
University of Oxford
These three site will assemble and test about 600 modules each
University of Glasgow has finished their pre-production batch with ITkPixV1.1 chip, and started their production batch of pixel modules with ITkPixV2 chip
Quad module half ring
Matheson, John, Jon Taylor, and A. T. L. A. S. ITk. "ATLAS ITk pixel endcap demonstrator “Ring-0” construction and testing." Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment 986 (2021): 164753.
Quad module
from top
Notes: Half ring picture

## Slide 6
Pixel Module workflow
For the ITk Strip, see Katie W. Pass’ talk
Bare module Reception
Bare module comes from hybridization vendor after flip chip and bump bonding
Visual inspection (VI) and IV is done
Assembly
Flexible PCB is glued with bare module
Wirebonding is done
Initial Warm
Done at 200C temperature
Sensor IV
Pull test is done
Final VI
Full Electrical test
Parylene Masking
VI
Mask put on module to prevent parylene to cover unwanted places
Caps on connectors
Parylene Coating
Module get coated with parylene
Parylene unmasking
Masks are taken out
Caps are taken out
VI
Post Parylene Warm
Sensor IV
Tuning of the chip
Pixel Failure Analysis (PFA)
Thermal Cycle
Thermally cycle the for 10 cycles from 400C to -450C
Final Warm
Final Cold
Done at 200C temperature
Sensor IV
Full Electrical test
Done at -150C temperature
Sensor IV
Full Electrical test
Module Loading
Modules are attached with half ring structure
Notes: Half ring picture

## Slide 7
Bare Module reception
20UPGB43200139: an example of passed module
Wentworth AVT702 probe station
Bare module comes from hybridization vendor after flip-chip of sensor and FE-chips
Manually probed at probe station with specialized chuck
Probed 26 bare modules with visual inspection, all with 150 µm sensor
One module has been failed for bad IV (greater than 1.5 µA/cm2)
Automatic probing being developed
Specialize chuck for multi-module probing

## Slide 8
Assembly
For the ITk Strip, see Katie W. Pass’ talk
Glasgow has four assembly stations, with four assembly jig
It takes about 30 minutes to assemble a module, but about 24 hours to cure the glue in jig, wirebonding, pull test, visual inspection needs additional time
assembly jig has been designed to minimize variation due to operator. It controls relative position of bare module and flex PCB and glue thickness.
With four jigs, four modules can be made at a same time
Glasgow has committed to produce 8 modules in a week with wire bonding, pull test and visual inspection

## Slide 9
Quality Control Testing
For the ITk Strip, see Katie W. Pass’ talk
After assembly, a visual inspection is done
An Initial Warm test is done maintaining warm temperature (200C)- a series of tests including module IV
Again Visual inspection is done, connectors on module are masked in preparation for parylene coating of module and shipped to Liverpool for coating Liverpool is going to parylene coat all the modules in UK cluster
Module is shipped back to Glasgow, unmasked and visually inspected again, and made ready for Final Test:
Post-parylene warm test (200C) is done, module is thermally cycled for 10 cycles from +400C to -450C
Final Warm test is done at warm temperature (200C), and Final Cold test is done at cold temperature (-150C) with a source scan with Mini X2 X-ray gun
Masking process
Visual Inspection
Quad Module
Masked Module

## Slide 10
Module testing setup (warm)
Warm parallel Setup
DAQ PC
Card 1
Card 2
Card 3
Card 4
Wiener crate HV supply
HMP 4040
LV
Chiller
DMM 6500
(10 Ch)
Vacuum
DCS
GREEN
DCS
GOLD
DCS
ORANGE
Blue
DCS
Chiller is always set to 140C
Four modules can run parallelly
Four Device Control Setup (DCS) is there to read temperature and humidity of the four modules
Initial warm test can be done in four modules by ~ 3 hours

## Slide 11
Module testing setup (Cold)
DAQ PC
Card 1
Card 2
Card 3
Card 4
DMM 6500
(10 Ch)
HMP 4040
LV
Wiener crate HV supply
Chiller
DCS
Peltier Switch box
Cold parallel Setup
Chiller is always set to - 400C
Four modules can run parallelly
Thermal cycle of 10 (+400C to -450C) takes almost 7 hours
Final test takes almost two days
Mini X2 X-ray gun

## Slide 12
Module test results
Key | Data |  |  | 
MODULE_TEMPERATURE | 20 |  |  | 
ANALYSIS_VERSION | localdb-tools v2.5.0 |  |  | 
QUAD-MODULE_ADC_CALIBRATION | 1 | 2 | 3 | 4
QUAD-MODULE_SLDO | 1 | 2 | 3 | 4
QUAD-MODULE_VCAL_CALIBRATION | 1 | 2 | 3 | 4
QUAD-MODULE_ANALOG_READBACK | 1 | 2 | 3 | 4
QUAD-MODULE_LP_MODE | 1 | 2 | 3 | 4
QUAD-MODULE_INJECTION_CAPACITANCE | 1 | 2 | 3 | 4
QUAD-MODULE_MIN_HEALTH_TEST | 1 | 2 | 3 | 4
QUAD-MODULE_TUNING | 1 | 2 | 3 | 4
QUAD-MODULE_PIXEL_FAILURE_ANALYSIS | 1 | 2 | 3 | 4
MODULE_BAD_PIXEL_NUMBER | 62 |  |  | 
MODULE_ELECTRICALLY_BAD_PIXEL_NUMBER | 57 |  |  | 
MODULE_DISCONNECTED_PIXEL_NUMBER | 5 |  |  | 
QUAD-MODULE_DATA_TRANSMISSION | 1 | 2 | 3 | 4
MODULE_DISABLED_COLUMNS_NUMBER | 0 |  |  | 
In each stage of quality control process, a series of tests are repeated including IV at the different stages
Electrical tests include trimming of different internal circuits and registers, evaluation of internal currents and voltages
Tests also includes scanning of all pixels, tuning , disconnected bump scan and source scan with X-ray gun
To pass, a module has to pass in all the tests (green) at all stages
Threshold scan after tuning
ToT scan
Disconnected Bump scan
Source scan
SLDO

## Slide 13
Module IV yield
For the ITk Strip, see Katie W. Pass’ talk
IV of one good module on all stage
IV of all module on all stage
26 bare modules received, and reception tests done,18 have been assembled into modules and went through initial test. 14 were tested through final stage
At initial test, one module breaks down early
At Final test, another module recorded reduction in breakdown voltage (reduction more then 10 V)
All other module has leakage current less than 1.5 µA/cm2 ,and breakdown voltage more than 200 V

## Slide 14
Module Electrical QC yield
For the ITk Strip, see Katie W. Pass’ talk
One of the modules have one chip (chip 4) disabled
Some of the modules are yet to be finished testing after parylene coating
Overall yield ~73% , which is much better then pre-production yield (53.8%)

## Slide 15
Before and after extensive Thermal cycle
Before TC
After TC
No Difference is detected in terms of IV and Bump delamination
 | Before Thermal Cycle | After Thermal Cycle
MODULE_BAD_PIXEL_NUMBER | 91326 | 91326
MODULE_ELECTRICALLY_BAD_PIXEL_NUMBER | 91321 | 91321
MODULE_DISCONNECTED_PIXEL_NUMBER | 5 | 6
QUAD-MODULE_PIXEL_FAILURE_ANALYSIS | 1 | 2 | 3 | 4
One module was considered to perform long Thermal cycles (normal thermal cycle consist of 10 cycles)
Cycled within +400C to -450C
Cycled 70 time over two days
Intention is to stress the module thermally to see any deviation in module characteristics, specially bump bond delamination

## Slide 16
Before TC
After TC
Before TC
After TC
Disconnected
Bump Scan
Source Scan
Chip 4 has bad communication, do not record any hits in source scan
No significant changes can be seen, other then reduction of noise in chip 4
Before and after extensive Thermal cycle-continued

## Slide 17
 | Before Thermal Cycle | After Thermal Cycle
MODULE_BAD_PIXEL_NUMBER | 6554 | 6553
MODULE_ELECTRICALLY_BAD_PIXEL_NUMBER | 6296 | 6301
MODULE_DISCONNECTED_PIXEL_NUMBER | 258 | 252
QUAD-MODULE_PIXEL_FAILURE_ANALYSIS | 1 | 2 | 3 | 4
Another module was considered to perform long Thermal cycles with broken core column (known as Core Column issue)
Cycled within +400C to -450C
Cycled 70 time over two days
Intention is to stress the module thermally to see any deviation , or increase in more broken columns
No Difference is detected in terms of bump delamination or increase of core column problem
Before and after extensive Thermal cycle-CC module

## Slide 18
Test beam result
https://indico.cern.ch/event/1455346/contributions/6323044/attachments/3006994/5300730/Trento_2025_ITk_Pixel_Testbeam_CKrause.pdf
One module was tested under 120 GeV pion beam at CERN
Module was tested before irradiation
It showed more then 99 % hit efficiency at 20 V bias at sensor
More study will be done after irradiation
More information at poster session
“Up-to-date test beam results of ATLAS ITk Pixel sensors and modules”

## Slide 19
Conclusion
Production has started slowly, but progressively
No significant bump delamination was seen in any module
Though we have a bit less yield (~73%) then the target yield (78%), we are trying to improve
Maximum failure came from core column, software and a bit handling
Have identified some problems in process, and learning more and adapting
With more experience predicting to be much better in future production batches