## Slide 1: DynamiX A prototype high-framerate, high-dynamic-range hard X-ray detector for 4th generation synchrotrons
Simon Knowles (STFC UKRI)
On behalf of the XIDyn Collaboration
iWoRiD 2025

## Slide 2: DynamiX Overview
Charge-integrating hybrid pixellated detector system prototype
Aims to meet challenges of synchrotron upgrades
  Higher fluxes → new ASIC architecture
  Higher energies → CZT sensor material
16 x 16 pixels on 110μm pitch
534kHz (adjustable) framerate to match Diamond orbit
Readout at 14Gbps
DynamiX bonded to 2mm Redlen HF-CdZnTe
Aligning DynamiX at Diamond Light Source (DLS)
Notes: 110um is not a small pitch!

## Slide 3: Content
Motivations
Chip architecture
Results from testing at DLS B16
Summary

## Slide 4: Motivation
Diamond II upgrade
Up to 1011-12 ph/mm2/s
Many beamlines >20keV
Mpixels for 2029
Credit: https://www.diamond.ac.uk/Diamond-II.html
4th generation
First light: 2029
Chart: Energy (keV)

## Slide 5: DynamiX Spec
Parameter | DynamiX
Flux | >109 ph/s/pixel
Pixel Size | 110 μm
Frame Rate | 534 kHz (Synchrotron determined)
Dynamic Range | >1011 ph/mm2/s
Energy | 25keV typical
Noise Level | <single photon
The DynamiX Chip

## Slide 6: Chip Architecture
Incremental digital integration (per pixel):
2 stage charge cancellation circuit
Signal integrated in coarse stage cancelled and remainder passed to fine stage
Coarse and Fine counters output
Notes: Start explanation with why we need charge subtraction then discuss how it works

## Slide 7: Coarse Stage (<6400 ph/pix/frame @ 30keV)
Signal
Coarse threshold
Coarse counts

## Slide 8: Transfer and Fine Stage (<25 ph/pix/frame @ 30keV)
1 orbit of Diamond
8bit Coarse
7bit Fine
8bit Coarse
7bit Fine
NEW
FRAME
NEW
FRAME
Notes: CHARGE REMOVAL AND THRESHOLD ARE

## Slide 9: Results from beam tests at DLS B16

## Slide 10: Setup at DLS B16 Test Beamline
Focusing used to access sub-pixel-sized beam
Beam energies of 20,30,40 and 50keV
Sensor biased to -1000V
Later switched to DCM to access higher fluxes at 20keV

## Slide 11: Charge sharing – raster scanning beam within pixel (20keV, ~20µm beam)
Measured signal in pixel strongly depends on beam alignment to centre of pixel
Evidence of charge sharing to neighbouring pixels in CZT
Destroys ability to resolve single photon peaks  challenging to calibrate without pencil beam centred in-pixel
Notes: PLEASE DELETE/REPLACE PHOTO CREDIT IF CHANGING THE IMAGE.

## Slide 12: Single photon peaks
Able to resolve single-photon peaks (up to 10 single-photon peaks for 20keV beam)
Width σ ≈ 5.7keV (≈1200 e- RMS)
Photodiode-measured fluxes:
  20keV: 1.58 x106 ph/pix/s
  30keV: 2.80 x105 ph/pix/s
20keV
30keV

## Slide 13: Single photon peaks – changing fine packet size
Measurements taken at 20keV, flux ≈1.6x106 ph/s/pix
Can change digitisation resolution at expense of dynamic range
Smaller packets  more packets to remove same charge
Better quality fit with smaller packets  easier to estimate noise
Settled on packet size 100 DAC as this looks like 1 fine count ≈ 1 keV
Notes: PLEASE DELETE/REPLACE PHOTO CREDIT IF CHANGING THE IMAGE.

## Slide 14
20keV beam, moving Al wedge into beam
Linearity with increasing flux
Design spec:
≈4 x 1011 ph/s/mm2 @20keV
Limit of flux from X ray set:
≈3x1010 ph/s/mm2 equiv. @ 20keV
160kV W X-ray set (polychromatic)
Limited by flux available at the beamline, not the DynamiX detector
Notes: PLEASE DELETE/REPLACE PHOTO CREDIT IF CHANGING THE IMAGE.

## Slide 15: Whole array measurements
Imaged a spinning slitted disc (copper and steel) at 533kHz
100kV tube voltage, 30mA tube current
Frames flat-field corrected
Averaged over windows of 10,50 and 100 frames
Notes: PLEASE DELETE/REPLACE PHOTO CREDIT IF CHANGING THE IMAGE.

## Slide 16: 14 Gbps Serialiser Results
Aurora encoded 66b64b CML
14 Gbps working and continuously writing to disc at 534 kfps
Data passes through retimer IC  allows inspection of eye quality
Conversion to optical via Samtec FireFly close to ASIC
Receiver and frame built in FPGA with 100G UDP out
Eye diagrams without and with retiming
Samtec FireFly and retimer IC next to DynamiX ASIC

## Slide 17: Towards the full-scale ASIC, XIDyn
192 x 144 pixel ASIC, combining XIDER and DynamiX designs
Improvements to pixel design, including:
  Increased trim range for calibration of thresholds & packet sizes
  Switch to disconnect front end from pixel during transfer to prevent saturation
Submission planned for autumn this year, with wafers expected back in spring
DynamiX Chip
XIDER Chip

## Slide 18: Summary
Continuous full-frame readout at 534 kHz
Able to resolve single photons with width σ ≈ 5.7keV
Measurements taken in range <1ph/pix/frame  ≈900 ph/pix/frame
Submission of full ASIC in autumn, tileable on 3 sides
Upcoming beamtimes at Diamond & ESRF to go to higher fluxes
Testing at B16 DLS in Feb 2025
Rasting scanning beam within pixel
Single-photon peaks @ 20 keV
Still from imaging spinning slitted disc at 534kfps
The DynamiX ASIC with Redlen HF-CZT

## Slide 19: Additional Slides

## Slide 20
Leakage current shift
(FL = 3 Diamond turns)
1 and 2 photon peaks move to higher channel numbers with increasing flux
With FL = 1 turn, beam 50 µm x 50 µm, excess leakage current density in this flux range ≈ 41 nA/mm2
This effect has been seen in Redlen HF CZT before, e.g. Cline and XIDER
Notes: PLEASE DELETE/REPLACE PHOTO CREDIT IF CHANGING THE IMAGE.

## Slide 21
Calibrating from single photon peaks
(aim: 1 coarse = 64 fine)
1.Determine ADU  keV conversion from photon peak spacing
Notes: PLEASE DELETE/REPLACE PHOTO CREDIT IF CHANGING THE IMAGE.

## Slide 22
Calibrating from single photon peaks
(aim: 1 coarse = 64 fine)
1.Determine ADU  keV conversion from photon peak spacing
2.Change fine threshold to move noise peak to 0
  Find threshold such that 50% counts are >0
Notes: PLEASE DELETE/REPLACE PHOTO CREDIT IF CHANGING THE IMAGE.

## Slide 23
Calibrating from single photon peaks
(aim: 1 coarse = 64 fine)
1.Determine ADU  keV conversion from photon peak spacing
2.Change fine threshold to move noise peak to 0
  Find threshold such that 50% counts are >0
Can also estimate noise:
Look at σ of dark counts where distribution is well away from 0
σ ≈ 5.50 LSB ≈ 5.57 keV ≈ 1210 e-
150fF fine stage feedback, includes leakage + sensor capacitance
Notes: PLEASE DELETE/REPLACE PHOTO CREDIT IF CHANGING THE IMAGE.

## Slide 24
Calibrating from single photon peaks
(aim: 1 coarse = 64 fine)
1.Determine ADU  keV conversion from photon peak spacing
2.Change fine threshold to move noise peak to 0
  Find threshold such that 50% counts are >0
3.Find test pulse setting to give 64 fine counts (≈65 keV)
Notes: PLEASE DELETE/REPLACE PHOTO CREDIT IF CHANGING THE IMAGE.

## Slide 25
Calibrating from single photon peaks
(aim: 1 coarse = 64 fine)
1.Determine ADU  keV conversion from photon peak spacing
2.Change fine threshold to move noise peak to 0
  Find threshold such that 50% counts are >0
3.Find test pulse setting to give 64 fine counts (≈65 keV)
4.Sweep coarse threshold while injecting 64 fine counts
  Find threshold such that 50% chance of triggering
  Set threshold slightly higher to avoid cancelling below baseline
Notes: PLEASE DELETE/REPLACE PHOTO CREDIT IF CHANGING THE IMAGE.

## Slide 26
Calibrating from single photon peaks
(aim: 1 coarse = 64 fine)
1.Determine ADU  keV conversion from photon peak spacing
2.Change fine threshold to move noise peak to 0
  Find threshold such that 50% counts are >0
3.Find test pulse setting to give 64 fine counts (≈65 keV)
4.Sweep coarse threshold while injecting 64 fine counts
  Find threshold such that 50% chance of triggering
  Set threshold slightly higher to avoid cancelling below baseline
5.Inject 3 coarse counts (=3 x 64 fine counts) and sweep coarse packet size
  Find maximum packet corresponding to 2 coarse counts
Notes: PLEASE DELETE/REPLACE PHOTO CREDIT IF CHANGING THE IMAGE.

## Slide 27
Calibrating from single photon peaks
(aim: 1 coarse = 64 fine)
1.Determine ADU  keV conversion from photon peak spacing
2.Change fine threshold to move noise peak to 0
  Find threshold such that 50% counts are >0
3.Find test pulse setting to give 64 fine counts (≈ 65 keV)
4.Sweep coarse threshold while injecting 64 fine counts
  Find threshold such that 50% chance of triggering
  Set threshold slightly higher to avoid cancelling below baseline
5.Inject 3 coarse counts (=3 x 64 fine counts) and sweep coarse packet size
  Find maximum packet corresponding to 2 coarse counts
6.Fine-tune coarse packet size by verifying goodness of fit while injecting increasing charge per frame
Fit only where injected charge < coarse thresh
Notes: PLEASE DELETE/REPLACE PHOTO CREDIT IF CHANGING THE IMAGE.

## Slide 28: Pixel Test Results – Coarse Stage (inbuilt test pulse)
Sweep coarse stage threshold → measure coarse output → repeat to obtain probability of triggering → determine dark noise
Vary input test current → measure coarse ADC output
DNL 0.2 +/- 0.1
INL 0.7 +/- 0.3
Vary input test current → measure coarse and fine counts
Noise certainly above coarse threshold – triggers a count
Noise certainly below coarse threshold – does NOT trigger a count
Injected test current (a.u.)
Injected test current (a.u.)

## Slide 29: Pixel Test Results – Fine Stage (inbuilt test pulse)
1 fine count = 0.25 of 30keV Photon
No current injected → measure counts due to noise
Increase injected current → measure fine ADC output
Fine stage ADC linearity
Coarse =0
Coarse =1
Injected test current (a.u.)
Dark level/ fine counts

## Slide 30
Single photon peaks
Monochromatic beam focused to sub-pixel level ≈50 μm x 50 μm and centred in pixel
Can resolve single photon peaks with a width σ ≈ 5.7keV (~1200 e- RMS)
Misalignment between peaks of same energy:
  μ2x20 keV ≠ μ40 keV
  μ3x20 keV ≠ μ2x30 keV
  …
Thought to be due to flux-dependent excess leakage current in HF-CZT
40 keV
60 keV
80 keV
100 keV
Notes: PLEASE DELETE/REPLACE PHOTO CREDIT IF CHANGING THE IMAGE.