## Slide 1: Detector Developments for the Swiss Light Source Upgrade
A. Bergamaschi, M. Brückner, M. Carulla, R. Dinapoli, S. Ebner, K. Ferjaoui, E. Fröjdh, V. Gautam, D. Greiffenberg, S. Hasanaj, A. Haugdal, J. Heymes, V. Hinger, V. Kedych, T. King, S. Li, C. Lopez-Cuenca, A. Mazzoleni, D. Mezza, K. Moustakas, A. Mozzanica, J. Mulvey, M. Müller, K.A. Paton, C. Posada Soto, C. Ruder, B. Schmitt, P. Sieberer, S. Silletta, D. Thattil, X. Xie, J. Zhang
IWORID20205, Bratislava (SK)
Monday 7th July 2025
Center for Photon Science, Paul Scherrer Institute (Switzerland)
Notes: The Swiss Light Source is nearing the completion of its upgrade to a fourth-generation synchrotron, significantly enhancing its brilliance. This leads to a higher photon flux, while improving the coherence of the beam across the entire energy spectrum, from extreme ultraviolet to 80 keV hard X-rays.
To support the increased photon rates, we are developing the MATTERHORN detector, a single-photon counting pixel detector with a 75 µm pitch. It features faster shaping and utilizes multiple comparators and counters to achieve a count rate of up to 20 MHz per pixel with 90% efficiency. The novel digital readout architecture allows for a frame rate of up to 10 kHz with a single counter per pixel in 16-bit mode. This fast readout technology can also be combined with on-chip digitization in our charge-integrating detectors, enabling the handling of extreme photon rates using dynamic gain switching. The planned upgrade of JUNGFRAU will allow to reach up to 100 million photons per pixel per second at a 10 kHz frame rate. Additionally, we are developing a 1 Megapixel MÖNCH detector with a 25 µm pitch, offering noise levels below 20 e- rms ENC and 2 kHz frame rate. This will improve performance in photon-starved applications like Resonant Inelastic X-ray Scattering (RIXS) and fluorescence emission imaging, where higher brilliance enables greater sensitivity.
To extend the energy range of our detectors, we are improving sensor technologies beyond standard silicon. For hard X-rays, we are testing and characterizing high-Z sensors, primarily GaAs and CdZnTe, to build large-area single-photon counting detectors for hard X-ray diffraction. For soft X-rays below 2 keV, we collaborate with FBK to develop LGAD sensors optimized for soft X-ray detection, enabling single-photon counting down to 500 eV for the first time. The possibility of interpolating using LGAD sensors allows a resolution of approximately 1um and opens the possibility to develop a new hybrid detector technology for RIXS. These advancements will significantly contribute to the experiments at SLS2.0, paving the way for exciting new scientific discoveries.

## Slide 2: Synchrotron applications
Synchrotrons are electron accelerators used to generate X-rays
High X-ray flux
  Energy selection through monochromator
    EUV (<100 eV) to hard X-rays (>100keV)
  High intensity on the detector
  Photon-starved applications possible
  High speed
Coherence
Polarization
Pulsed time structure
Macromolecular Crystallography Leonarski et al.
Operando diffraction Hocine et al.
Full field fluorescence
imaging
Marone et al.
Ptychography
Holler et al.
Magnetic contrast ptychography – Butcher et al.
Mitchell et al, 1999. Structure 7 (5), R111–R121

## Slide 3: SLS 2.0 – a 4th generation synchrotron
Improved emittance, increased brightness
  Higher coherence
  More photons hitting the sample
First photons to the beamlines now
Pilot users at some beamlines Q3/2025
User operation Q3/2026

## Slide 4: Hybrid detectors
Semiconductor sensor
  Direct conversion
Versatile CMOS readout chip
  Fast readout
  Radiation hardness
Bump bonding
  Minimum pixel pitch ~ 20 μm
  Noise due to input capacitance
Large area by tiling

## Slide 5: Everything in-house

## Slide 6: PSI Detector portfolio
 | Microstrip 1D | Pixel 2D | 
 | 25/50 µm pitch | 75 µm pitch | 25 µm pitch
Single photon counting |  |  | 
Charge integrating |  |  | 
MYTHEN
GOTTHARD
MÖNCH 0.3
JUNGFRAU
EIGER

## Slide 7: Readout modes
Noiseless, very large dynamic range
Stable operation, well established calibration
Fluorescence suppression
Frame rate scales with counter depth
Pile-up ≳1MHz/pixel
Soft X-ray detection prevented by noise
Compromise between noise and dynamic range
Works also at pulsed sources
Detection of multiple soft X-rays
Saturation for long integration times
fluorescence
Total
5
fluorescence
Total
5
+1
+1
+1
+1
Total
4
Total
5.5
Photon counting
Charge integrating

## Slide 8: State of the art
Noiseless, very large dynamic range
Stable operation, well established calibration
Fluorescence suppression
Frame rate scales with counter depth
Pile-up ≳1MHz/pixel
Soft X-ray detection prevented by noise
Single photon resolution for hard/tender X-rays
  Detection of multiple soft X-rays
Dynamic range up to 104 photons/pixel/s
Saturations for long integration times
Challenging calibration and cumbersome data processing
fluorescence
Total
5
fluorescence
Total
5
+1
+1
+1
+1
Total
4
Photon counting
Charge integrating with adaptive gain

## Slide 9: Upgraded light sources: many more photons!
Photon counting
Charge integrating with adaptive gain
Noiseless, very large dynamic range
Stable operation, well established calibration
Fluorescence suppression
Frame rate scales with counter depth
Faster shaping, pile-up tracking
Pile-up ≳1MHz/pixel
Soft X-ray detection prevented by noise
Single photon resolution for hard/tender X-rays
  Detection of multiple soft X-ray
Works also at pulsed sources
Dynamic range up to ~104 photons/pixel/s
Supported flux scales with frame rate
Saturations for long integration times
Challenging calibration and cumbersome data processing needs high throughput!

## Slide 10: PSI Detector portfolio
 | Microstrip 1D | Pixel 2D | 
 | 25/50 µm pitch | 75 µm pitch | 25 µm pitch
Single photon counting |  |  | 
Charge integrating |  |  | 
MATTERHORN
MYTHEN
GOTTHARD
MÖNCH1.0
JUNGFRAU2

## Slide 11: MATTERHORN: the next peak for single photon counting
Full size ASIC submitted last week
Dual polarity analog chain
4 independent comparators and gatable counters
  Pile-up tracking
  Pump-multiprobe measurements
Synthesized digital periphery
On-chip PLL
4 x 3.125 GHz serializers / chip
100 Gb/s link / module
 | EIGER | MATTERHORN
Pixel size | 75 x 75 µm2 | 
Module size | 4 x 8 cm2 | 
Thresholds | 1 | 4
Min Threshold | ~ 3 keV | ~ 1.5 keV
Counter depth | 12 bit | 4 x 16 bit
Max. Frame rate (8 bit) | 10 kHz (burst) | 20 kHz (continuous)
Count rate (90% efficiency) | 0.35 Mcounts/pixel/s | 20 Mcounts/pixel/s
Gating | ~ µs | ~ 20 ns
Hocine et al., Materials Today 2020
Aidukas et at., Nature 632, 81–88 (2024).

## Slide 12: MATTERHORN prototypes
Count rate capability with 90% efficiency
      2.5 MHz (1 counter)
      >7.5 MHz (3 counters)
        Below expectation, due to low comparator current used
Preliminary prototype results (MH01-MH02) | 
Pixels | 48 x 48
Min threshold | < 2keV
Noise | 80 e- ENC rms
Threshold dispersion | 33 eV rms
Noise vs. rate capability
(one counter) MH01
Sum counter 0-3
Counter 0
Dose rate dependence of TID Damage in front-end electronics for SLS 2.0 Viveka Gautam
Monday poster session
12
Courtesy E. Frojdh

## Slide 13: JUNGFRAU2: faster, even higher intensities
 | JUNGFRAU | JUNGFRAU2
Pixel size | 75x75 µm2 | 
Module size | 4x8 cm2 | 
Dynamic range (12 keV) | 104 | 
Min Energy* | ~ 800 eV | ~ 600 eV
Max Frame rate | 2.2 kHz | 10 kHz
Max count rate (12 keV) | 22 Mcounts/pixel/s | 100 Mcounts/pixel/s
Storage cells (burst mode) | 16 | 16-64 Rebin-able
* with single photon resolution, standard silicon
Wranik et al. Nat Commun 14, 903 (2023).
At Synchrotrons full duty cycle 2 kHz 10 kHz
2 GB/s/module 10 GB/s/module
190 TB/day/module 1 PB/day/module
5.3 PB/day for a 16M detector 25 PB/day for a 16M detector
Characterization of the JUNGFRAU 1.2 readout ASIC
Vadym Kedych
Thursday 11:00
A 1M charge integrating hybrid pixel detector for electron diffraction
Khalil Ferjaoui
Monday 15:00

## Slide 14: Maximum detectable intensities
Chart: Year
| 2017 | 2021 | 2027
Charge Integrating | 1777 | 3555 | 17777
Single Photon Counting | 0.22 | 33 | 177 | 135 | 533 | 4444
EIGER
MATTERHORN
Retriggering
by DECTRIS
PILATUS
Large pixels
JUNGFRAU2
Hatsui, PIXEL 2022, full duty cycle, burst mode 17.4 kHz
CITIUS
Faster than Moore! ≈2t/1.55y

## Slide 15: Silicon quantum efficiency
High-Z
LGAD
Standard
Silicon
3D Doping Concentration Imaging of Silicon Sensors Using Backside Pulsing
Xiangyu Xie
Wednesday poster session

## Slide 16: High-Z sensors for hard X-rays
Enhancing the Performance of High-Z Sensors for Photon Science Applications
Kirsty Paton
Tuesday 10:10
CZT as a Material for Detecting Hard X-rays at Synchrotrons and FELs
Jonathan Mulvey
Tuesday 9:50

## Slide 17: Soft X-rays detection
n+
Al
p+
Al
x
x x
ASICS
SiO/N
Zhang, et al. JINST 2022; 17(11): C11011.
GOTTHARD 8.0 keV
Andrä, Zhang et al.,
Jour. Synch. Rad. 2019.
Amplification
Shallow absorption of soft X-rays
    Thin entrance window technology with improved doping profile and thin passivation
  Charge per photon comparable to electronic noise
    Exploit internal multiplication of LGADs to increase SNR
n+
n-
Al
p+
Al
x x x x x x x x x x x x x
ASIC
Standard
n-
p+
Characterisation of inverse LGADs in the soft X-ray energy range Shuqi Li
Tuesday 11:00
Hard X-ray sensors
Optimized implant
Passivated surface
Thin passivation (2025)
Carulla, et al.
Sensors. 2024; 24(3): 942.
& unpublished results
Courtesy J. Zhang
p-

## Slide 18: Single photon counting below 1 keV
Largest LGAD ever!
Single photon resolution down to ~500 eV
Effective noise reduction
Cooling required
Scientific papers at the Fe L-edge (700eV) and O K-edge (530 eV)
Image of a FZP acquired using EIGER
SIM beamline SLS
F. Baruffaldi et al., under review.
4 4 cm2 iLGAD
Notes: Here is a reconstruction of the Siemens star with the new module. This is 6 x 6 um and the phase looks sharper. The Siemens star was 150 nm thick and made of permalloy. It looks much better than the previous module and this is the one we will use for the paper on the LGAD Eiger. Do you have a rough estimate of the timeline for that manuscript? I also attached an amplitude XMCD magnetic image from a Pt-Co multilayer (magnetisation out of plane) in which you can see the magnetic domains nicely. That structure is around 2 um and the image was taken with the previous module in February. Please let me know if you need something else.
XMCD: X-ray Magnetic Circular Dichroism
With the new module W9 we obtained a resolution of approximately 6 nm.

## Slide 19: More X-rays for photon-starved applications
fluorescence
Total
5
Total
5
Readout and reset
With fast frame rates and low fluxes
  Energy resolution ≲ keV
  Super-resolution ~µm
Developments to sustain higher fluxes
  Faster frame rates
  Machine learning to handle pile-up
aare – A Flexible Data Analysis Library for Hybrid Pixel Detectors Erik Fröjdh
Wednesday poster session

## Slide 20: Super-resolution
100 μm
25 µm pixels
1 µm bins
Silicon sensor @ 10keV
100 µm
25 µm pixels
5 µm bins
Chiriotti et al., JINST (2022)
GaAs sensor @ 10keV
CdZnTe sensor @ 40 kVp
300 µm
Crystal defect
D. Greiffenberg et al., unpublished
75 µm pixels
5 µm bins
100 µm
LGAD sensor @ 500eV
25 µm pixels
2.5 µm bins
J. Zhang et al., unpublished
High frame rate RIXS spectroscopy using a JUNGFRAU detector with an iLGAD sensor
Viktoria Hinger
Wednesday poster session
Notes: Timepix4 160 Gbps 2x2cm2 250kpixels 55um
event driven readout
Ca. 1.5MHz/pixel
Ca. 4MHz/mm2
Moench 20Gbps 1x1cm2 160kpixels 25um
6 kHz/pixel
0.1 MHz/mm2

## Slide 21: MÖNCH: high spatial and energy resolution
Targeting low noise
  Hyperspectral imaging with super-resolution
  RIXS combined with LGADs
Scalability is challenging
  800 kpixel on a single chip
  1.6 Mpixel module, non negligible gap
Huge data throughput
  ~ kHz frame rate, compromise with ROI
  On-the-fly data reduction is mandatory!
MÖNCH prototype
0.4 x 0.4 cm2
MÖNCH big prototype
1 x 1 cm2
MÖNCH1.0
2.56 x 1.92 cm2
MÖNCH1.0
modules
2.56 x 3.84 cm2
2013
2015
2026
2028
C. Dullin et al., J. Synchrotron Rad. (2018)
Fourier ptychography
Wakonig et al.
Full field fluorescence
imaging
Marone et al.

## Slide 22: Summary
MATTERHORN single photon counting
Faster shaping
Pile-up tracking
Higher frame rate
JUNGFRAU2 and Megapixel MÖNCH
High frame rate
High performance data conversion and compression platform
LGADs for soft X-rays
Soft X-ray single photon counting
RIXS detector with interpolation
High-Z sensors for hard X-rays
Large area single photon counting
High resolution energy dispersive imaging

## Slide 23: Detector plans for SLS2.0
16M MATTERHORN
1.5M MATTERHORN (OMNY)
1.5M JUNGFRAU (Serial laminography)
500k MATTERHORN LGAD
500k MATTERHORN (PEEM)
500k JUNGFRAU CZT (FFFT)
4M MATTERHORN High-Z
1M MÖNCH LGAD
60k MYTHEN3
4M MATTERHORN
4M MATTERHORN High-Z
4M MATTERHORN LGAD
4M JUNGFRAU LGAD
4M MATTERHORN High-Z
16M MATTERHORN
16M JUNGFRAU

## Slide 24
1st Joint X-ray Sensor Workshop From Soft to Hard X-rays
ETH Zurich (Switzerland) 12.-16. January 2026
Bridging Science and Detection Technology for Soft X-ray Applications
Call for abstract is open!
High-Z Pixel Array Detectors for Photon Science
Abstract: hizpad.community@gmail.com
Webpage and registration: https://indico.psi.ch/e/XRaySensors2026
Registration for the two workshops is separate.
Notes: Change QR code

## Slide 25: Thank you!
Back: V. Kedych, M. Müller, P. Sieberer, K. Ferjaoui, M. Brückner, E. Fröjdh, J. Heymes, C. Lopez-Cuenca.
Middle: K. Moustakas, C. Ruder, J. Mulvey, S. Silletta, V. Hinger, A. Mozzanica, M. Carulla, X. Xie,.
Front: R. Dinapoli, A. Bergamaschi, S. Li, D. Thattil, J. Zhang, A. Mazzoleni, V. Gautam, D. Mezza, T. King, B. Schmitt.
Missing: S. Ebner, D. Greiffenberg, S. Hasanaj, K.A. Paton, C. Posada, A. Haugdal.
Dose rate dependence of TID Damage in front-end electronics for SLS 2.0 Viveka Gautam
Monday poster session
3D Doping Concentration Imaging of Silicon Sensors Using Backside Pulsing
Xiangyu Xie
Wednesday poster session
aare – A Flexible Data Analysis Library for Hybrid Pixel Detectors Erik Fröjdh
Wednesday poster session
High frame rate RIXS spectroscopy using a JUNGFRAU detector with an iLGAD sensor
Viktoria Hinger
Wednesday poster session
Characterization of the JUNGFRAU 1.2 readout ASIC
Vadym Kedych
Thursday 11:00
A 1M charge integrating hybrid pixel detector for electron diffraction
Khalil Ferjaoui
Monday 15:00
Enhancing the Performance of High-Z Sensors for Photon Science Applications
Kirsty Paton
Tuesday 10:10
CZT as a Material for Detecting Hard X-rays at Synchrotrons and FELs
Jonathan Mulvey
Tuesday 9:50
Characterisation of inverse LGADs in the soft X-ray energy range Shuqi Li
Tuesday 11:00

## Slide 27: Large detectors
>2.5 mm
Chip
8 cm
4 cm
>0.5 mm

## Slide 28: New counting modes
Reduce shaping time
Depend on fill pattern
Challenging calibration
Photons
Analog and discriminator signal
Paralyzable
Time over Threshold
Bergamaschi, JSR (2011)
Retriggering
Zambon, NIM A (2021)
Pile-up tracking
Andrä, JINST (2019)

## Slide 29: Calibration into number of photons
Fresnel zone plate at XIL beamline, SLS, 92eV
Etched sensor (no Al), module mounted to flange in high vacuum
Gain factors
Pedestals
Bit masking
S. Redford et al., JOURNAL OF INSTRUMENTATION 11, C11013 (2016).

## Slide 30: Adaptive gain methods
Noise increases with exposure time
Fast frame rate extends dynamic range
Challenging calibration
Only qualitative trends!
Linear
Ideal performance
Gain switching
JUNGFRAU, AGIPD, ePix
Non linear
DSSC
Charge removal
MM-PAD, Xider, Dynamix
Detector output
Input signal

## Slide 31: Adaptive gain methods
Only qualitative trends!
Noise increases with exposure time
Fast frame rate extends dynamic range
Challenging calibration
Detector output
Input signal
Linear
Ideal performance
Gain switching
JUNGFRAU, AGIPD, ePix
Non linear
DSSC
Charge removal
MM-PAD, Xider, Dynamix

## Slide 32: Energy resolved imaging at X-ray tubes
Siemens Star with spokes 60-0.5 µm gold on silicon with silicon microspheres
W-anode X-ray tube 40 kV 200 µA
A full spectrum with an energy resolution of about 750 eV FWHM is acquired for each pixel

## Slide 33: Edge subtraction imaging
Siemens Star with spokes 60-0.5 µm gold on silicon with silicon microspheres
W-anode X-ray tube 40 kV 200 µA
Images can be binned in energy
Below the L-edge gold becomes “transparent” to X-rays
Below Au L-edge

## Slide 34: High resolution energy resolved imaging
Siemens Star with spokes 60-0.5 µm gold on silicon with silicon microspheres
W-anode X-ray tube 40 kV 200 µA
Color imaging works also in combination with interpolation.
Below Au L-edge
Interpolated to 1 µm bins