## Slide 1
An overview of
HEXITEC and HEXITEC MHz and its applications
Dr Diana D Caprotti et al
8th of July 2025

## Slide 2
1 Who are we - UKRI-STFC-RAL
An introduction to the UKRI Science and Technologies Facilities Council Rutherford Appleton Laboratory.
2 HEXITEC
Introducing the latest version of the technology, its architecture, initial results and camera status
3 HEXITEC applications
An overview of some of HEXITEC applications - credit to our collaborators
Agenda
4 HEXITECMHz
Introducing the latest version of the technology and latest results

## Slide 3
Who we are

## Slide 4
Who we are
…and around the world
UK Astronomy Technology Centre
Edinburgh, Scotland
Polaris House
Swindon, Wiltshire
Chilbolton Observatory
Stockbridge, Hampshire
Boulby Underground Laboratory
North Yorkshire
Rutherford Appleton Laboratory
Harwell Didcot, Oxfordshire
Daresbury Laboratory
Sci-tech Daresbury Campus, Liverpool City Region

## Slide 5
Who we are
Rutherford Appleton Laboratory
Home to:
Diamond Light Source
ISIS
Central Laser Facility
RAL Space

## Slide 6
Who we are
Technology’s Detector and Electronics Division
ASIC design
CMOS sensors design
Detector Development
Electronics system design
Interconnect
One stop shop for detector systems

## Slide 7
1 Who are we - UKRI-STFC-RAL
An introduction to the UKRI Science and Technologies Facilities Council Rutherford Appleton Laboratory.
2 HEXITEC
An overview of the HEXITEC system
3 HEXITEC applications
An overview of some of HEXITEC applications - credit to our collaborators
Agenda
4 HEXITECMHz
Introducing the latest version of the technology and latest results

## Slide 8
Parameter​ | Value​
Pixel Pitch (mm)​ | 250​
Array Size​ | 80 × 80​
Max Frame Rate (kHz)​ | 9.3​
High Gain (keV)​ | 2 – 200​
Low Gain (keV)​ | 6 – 600 ​
FWHM@60keV (keV)​ | < 1​
Max flux rate (photons s-1 mm-2) | < 2×104
Priority: High resolution hard X-ray spectroscopy
250 µm pixel pitch choice
    Limited charge sharing events
    Moderate spatial resolution
20 × 20 test chip delivered in 2009
Full 80 × 80 system delivered in 2014
2014+ large area systems delivered
The HEXITEC system

## Slide 9
Examples of typical frames of data
2mm CdZnTe sensor @ - 750 V
2 keV noise threshold applied
y = mx + c calibration
HEXITEC records the position and energy of each interaction
The HEXITEC system

## Slide 10
Energy values histogrammed for each pixel across many frames
Result: high resolution spectra!
2 mm CdZnTe @ - 750 V
Am-241 spectrum
FWHM = 700 eV
The HEXITEC system

## Slide 11
80 Pixels
241Am flat field

## Slide 12
The HEXITEC system
Detector Module
Gig E System
2x2 System
2x6 System
2x8 System

## Slide 13
1 Who are we - UKRI-STFC-RAL
An introduction to the UKRI Science and Technologies Facilities Council Rutherford Appleton Laboratory.
2 HEXITEC
An overview of the HEXITEC system
3 HEXITEC applications
An overview of some of HEXITEC applications - credit to our collaborators
Agenda
4 HEXITECMHz
Introducing the latest version of the technology and latest results

## Slide 14
Nuclear Medicine
HEXITEC licensed to MH3D on the alpha-SPECT mini
Target market: radiopharmaceutical development for Targeted Alpha Therapy
Credit: MH3D/ Scintica
EM Zanonni et al (2025): doi: 10.1109/TRPMS.2025.3560558

## Slide 15
K-edge Imaging
Transmission Imaging to differentiate between different Al alloys
  Using machine learning and neural networks to differentiate between different Al alloys
Credit: Valentina Vicini (2024)
DOI: 10.1109/NSS/MIC/RTSD57108.2024.10656027

## Slide 16
Compton Scattering
Credit: C Huang (2025)
J. Phys. Energy 7 (2025) 025009
Measuring and imaging Li+ ion diffusion in operando inside a solid-state Li metal battery
  Compton scattering intensity related to electron density and thus charge in the battery
cathode
anode

## Slide 17
Uranium quantification
Uranium quantification in Nuclear Fuel:
  Determine the mass of Uranium and Uranium Oxide powders using spectral radiography
  σ = 0.62 % with HEXITEC compared to σ = 0.20 % COMPUCEA, but COMPUCEA destroys the sample
  HEXITEC provides a stable and accurate non-destructive characterisation of Uranium
Credit: AJ Gilber et al (2023)
J. Appl. Phys. 134, 024902 (2023); doi: 10.1063/5.0147454
A radiograph of the depleted UO powder

## Slide 18
1 Who are we - UKRI-STFC-RAL
An introduction to the UKRI Science and Technologies Facilities Council Rutherford Appleton Laboratory.
2 HEXITEC
An overview of the HEXITEC system
3 HEXITEC applications
An overview of some of HEXITEC applications - credit to our collaborators
Agenda
4 HEXITECMHz
Introducing the latest version of the technology and latest results

## Slide 19
Designed in TSMC 180 nm
Full spectroscopy @ 2×106 ph s-1 mm-2
Similar energy resolution capabilities to HEXITEC
Parameter​ | HEXITEC | HEXITEC MHz
Pixel Pitch (mm)​ | 250​ | 250​
Array Size​ | 80 × 80​ | 80 × 80​
Max Frame Rate (kHz)​ | 9.3​ | 1000
Digitisation | Off-chip | On-chip
Detector type | Track + hold | Integrating
Energy range (keV) | 2 - 600 | 2 - 300
FWHM@ 60 keV (keV)​ | < 1​ | <1
Max flux rate (photons s-1 mm-2) | < 2×104 | > 106
Power consumption (W) | 1.5 | 15
HEXITEC MHz spec

## Slide 20
Integrating architecture
tint = 900 ns, treset = 100 ns
Front end noise ~ 100 electrons
TDCs shared in a 4×2 ‘super-pixel’
12-bit data out for each pixel
20 × 4.1 Gb s-1 serialisers
Total data rate ≈ 10 GB s-1
HEXITEC MHz architecture
FPGA histogramming of data set
25 MB time slice

## Slide 21
HEXITEC flux results
1×107 ph s-1 mm-2
Occupancy = 50%
Occupancy = 10%
1×106 ph s-1 mm-2

## Slide 22
Hyperspectral Ptychography
Ptychographic phase image of a test target (40 µm field of view). Data was recorded with a monochromatic beam on the single-chip HEXITEC detector.
Match undulator spectrum with the absorption edges of the elements of interest​
HEXITEC-MHz enables simultaneous quantitative element-specific tomographic ptychograph with high spatial resolution

## Slide 23
HEXITEC and HEXITEC MHz are spectroscopic detectors for use with high Z materials.
We are always looking for new collaborations and applications.
Please get in touch via hexitec@stfc.ac.uk
Final remarks
Thank you for listening!