## Slide 1
SC SYNCHROTRON AND GANTRY BASED ON CANTED COSINE THETA MAGNETS
E. Benedetto,
N. Al Harbi, U. Amaldi, D. Bergesio, A. Garonna, V. Ljubicic, P. Riboni, TERA Foundation
Experts:
F. Wenander, J. Pitters, C. Carli, S. Benedetti, M. Martino, CERN
A. Pikin, BNL
Collaborators:
G. Arduini, D. Tommasini, CERN
B. Auchmann, CERN/PSI
L. Brouwer, S. Prestemon, LBNL
Ideas and Technologies for a next generation facility for medical research and therapy with ions, Archamps, FR (ESI), 19-21 June 2018

## Slide 2: Proposal: Compact synchrotron and gantry
C6+ up to 430 MeV/u
  …and other light ions (He)
Full treatment in one cycle
  1010 C6+ (10x10x10 cm3 target)
SC technology (CCT) makes it compact
  Ring: ~27 m length
  Gantry: 5.3 m height

## Slide 3: Proposal: Compact synchrotron and gantry
Compact! compared to PIMMS-derived facilities

## Slide 4: Outline
Magnets technology
Synchrotron
Lattice
Injection/extraction considerations
Gantry
Optics and mechanical
Dose delivery

## Slide 5: Magnets: Canted Cosine Theta
Pure dipole field
  2 tilted solenoids
Nested quadrupoles
  add extra layers

## Slide 6: Magnets: Canted Cosine Theta
LBNL gantry
HL-LHC correctors
PSI 16-T CCT Design
for FCC-hh

## Slide 7: Magnets for the optics
Design choices:
900 bends with nested quads
Powered in series
Bmax=3.5 T
r=1.90 m
Apertures magnets (g.f.r.) diameter:
Gantry: 30 mm
Ring: 60 mm
kf=-kd=~1 m-2
 ~6.6 T/m @ 3.5T
F
F
D
D
PRELIMINARY

## Slide 8: Lattice
gtr >1.43(@430 MeV/n)
Qx~1.67 for slow extraction
kb~0.6 m-1
 4 T/m @ 3.5T
r=1.90 m
3 m
3.6 m
kq1
kq2
(1.67, 1.13)
PRELIMINARY

## Slide 9: Injection, Linac & Source
MultiTurn injection @ 10 MeV/u:
Need ~1010 C+6 in ~100 turns x 0.6 ms (rev. period)
Synergies with CERN:
EBIS MEDeGUN (F. Wenander et al.)
  compact source (1A e- current) w. small emittance: ~0.03p mm mrad, norm.
  C+6 ions, short pulses @180 Hz (developed for CABOTO linac)
  we need longer trap length, i.e. 1m (vs. 25 cm) & no rep. rate requirements
  1010 is on the limit !!!
  OR: source as ~RHIC EBIS (5A e- current, larger intensity but also emittance)
High frequency RFQ for ions (V. Bencini, A. Lombardi et al.)
IH structure (S. Benedetti et al.)

## Slide 10: Extraction
Slow extraction with RF knock-out
Qx~1.67 (3rd order)
ES
MS
Sextupoles:
resonance
phase
Q’x, Q’y
Orbit correctors
PRELIMINARY
Dm~900 between septa

## Slide 11: Gantry
Rotation independent
Achromat
RE11=RE33=1.3
FWHM=4mm @ isocenter
kb=0.8m-2
 ~6T/m @ 3.5T
PRELIMINARY
2.82 m
SMx
SMy
5.3 m

## Slide 12: Gantry mechanical concept
Radius ~5.3 m
< 20 T (without supports)
Magnet:
  gfr (chamber)=15 mm radius
  Outer radius (iron)= 370 mm
warm
cold
warm

## Slide 13: Extraction and dose delivery
Field stability at extraction  reduce intensity spikes
  CCT dipoles windings have large inductance (not the quads)
  All dipoles and nested quads powered in series + current trims
  Switch mode power supplies (>50kHz)
Oblique Raster Scanning method  see A.Garonna’s talk
  Assuming 10x10x10 cm3 volume:
    Dose delivery in 20s FAST!!!
    1010 C6+
  dB/dt ~0.05 T/s SLOW !!!
    Ramp ~60s (10 430 MeV/u)
    Less cryogenics for SC magnets
    Eddy currents ~negligible

## Slide 14: Summary
Proposal of compact design:
  with Canted Cosine Theta magnets
  Bmax=3.5T
  900 bends with nested quads
  Ring 27 m, gantry R=5.3 m
Preliminary study of Synchrotron & Gantry
Synergies with CERN developments:
  Canted Cosine Theta magnets, for HL-LHC
  High frequency RFQ, IH and MEDeGUN EBIS, for medical C6+ linac