## Slide 1
CEC/ICMC - 2025 – Reno (Nevada, USA)
[C3Or4B-01] Large-Scale: Systems, Facilities, and Testing
Operation of CERN’s major tests facility
with upgraded cryogenic infrastructure
for superconducting magnets, power links,
inner triplets String and radio-frequency cavities
for HL-LHC
T. Barbe, T. Dupont, F. Ferrand, V. Gahier, J. Gery, N. Guillotin, R. Mauny, A. Perin, A. Onufrena, O. Pirotte.
Author: Remi Mauny
CERN, Technology Department, Cryogenics group, Geneva 23, CH-1211, Switzerland, remi.mauny@cern.ch.

## Slide 2
RF Crab Cavities Cryo-Module
SM18 presentation
SM18: CERN’s major test facility for superconducting magnets, power links, and radio-frequency cavities
SM18 hall surface : 7200 m2
(Inner Triplet) IT Magnet
Power Link
(Inner Triplet) IT-STRING test bench

## Slide 3
[4.5 K ; 3.5 bar]
[4.5 K
1.6 bar]
[20 K ; 1.2 bar]
precooler
8x 80m3 GHe buffers
Cold Box
[25 g/s]
Dewar
LHe
25 kL
5 Vertical MAGNET test benches
4 Vertical RF Cavities test benches
2 RF Cryomodule test benches (in bunker)
New Cold Compressor Unit
(for IT-STRING)
Back-up
Pumping Line [ ~30 K ; ~15 mbar ]
IT-STRING Pumping Line
[ ~3.3 K ; ~15 mbar ]
CRPL-B (cryoplant)
Compressor Station
Cold Box
[35 g/s]
CRPL-A (Cryoplant)
Compressor Station
Warm Pumping Unit 1
Primary pumps
Roots pumps
Warm Pumping Unit 2
Primary pumps
Roots pumps
New IT-String : 6 magnets + Link
10 Horizontal Magnet & CPS* test benches
SM18 Cryogenics infrastructure
SM18: CERN’s major test facility for superconducting magnets, power links, and radio-frequency cavities
*CPS : Cold Powering System

## Slide 4
nicolas.guillotin@cern.ch
ICEC28-ICMC2022, April 25th – 29th, 2022
2
New Compressor station (CS-B)
New cold box (CB-B)
SM18 Cryogenics infrastructure - issues
CS - HP oil injection issue :
Observation : HP oil bypass valve (CV665) kept going out of adjustment during starting up process.
Consolidation:
HP sensor added
Starting up sequence to start HP oil pump revised
Synchronisation between PV697 & PV666 optimized
DPT(Oil/He) Setpoint decreased from 4 bar to 3 bar.
Photo 1 : Comparison old vs. new valve head
Cold Box - Issue:
inlet & outlet valves of Adsorbers 80 K not tight in line after 1 thermal cycle.
Manufacturer intervention to replace valve heads
25 kL Dewar – Issue (2021):
Flooding in the pit – Dewar tilted by 10°
MAIN Lhe storage

## Slide 5
Test Benches upgrade - CONTEXT
SM18: CERN’s major test facility for superconducting magnets, power links, and radio-frequency cavities
RF CRAB Cavities
Cryomodule
=> To maximize overlap of bunches
Material : Niobium
New Inner Triplet magnets
=> To focus beam
Material : Nb3Sn - NbTi
New CPS
=> To power magnet
Material : MgB2
HL-LHC project => increase Luminosity (= increase collisions) in the Large Hadron Collider
HL-LHC : High Luminosity – Large Hadron Collider
IT-STRING
test bench
To validate collective behavior
Integration in LHC during next Long Shut Down
Test
in
SM 18
*CPS : Cold Powering System

## Slide 6
1. Radio Frequency (RF) Crab Cavities – Test benches upgrades
300 K – 130 K :
Sensitive to mechanical stress during thermal transition
=> ∆T = 50 K max
130 K – 4.5 K :
Must limit Q-Disease => Cool Down as fast as possible
Nominal operation : 2 K
RF Test benches upgraded to fulfill with requirements
Cryogenics process improved
Ready to test RF Crab cavities series
Proto RFD C3M cool down in M7 bunker in August 2024
RF Dipole Crab cavity
made in pure Niobium

## Slide 7
2. Magnet test benches upgrades for Magnets & CPS (Sc-link)
From 2022, 5 Horizontal benches for magnets upgraded :
4 with a shuffling module to test new Inner Triplet magnets
1 to test the new HL-LHC Cold Powering System (72 m long)

## Slide 8
2. Cold Powering System - test bench
In Yellow : upgrade added to the original Magnet test bench
In green : Distribution Feed Box (DFX) : at 4.5 K – it’s where the Link ends-up with connections (splice) between the link in MgB2 to the busbars in Nb-Ti
In Blue : Superconducting (SC) Link; ~72 m ; 19 cables in MgB2
In red : Distribution feed box high (DFHX): where the link starts up with splices MgB2 / HTS < 20 K + 19 Current Leads HTS/Copper
2
1
4
3
2
1
3
4
He flow
4.5 K
< 20 K
< 60 K
Current Leads (x19)
Bypass
~ 290 K
~ 290 K

## Slide 9
Lesson learnt:
15 W permanently
MAGNETS (theoretical position)
Superconducting wires to supply current to magnets :
      Link (in MgB2)
      Bus-bar (in Nb-Ti)
Cold He flow
generated by Heater : 6 g/s
a DFX drawing with the 2 Helium bath in cascade
Constant generated flow of 6 g/s with and without current.
Cryogenic conditions OK with powering up to 94 kA !
2. Cold Powering System – DFX & Results

## Slide 10
3. IT String test bench (1/4)
to STUDY and VALIDATE the COLLECTIVE BEHAVIOUR
of a string if inner triplet magnets with their powering system.
(+ other non-cryogenics systems).
SM18 facility
~ 100 m

## Slide 11
3. IT String test bench (2/4)
3 cool down done for :
SQXL validation
Cooldown procedure
Heat load check
New CCU operation
[6 kW]
25 g/s

## Slide 12
3. IT String test bench 3/4
Cold Compressor
Thermalization
Nominal operation
Bearing Wup
CC off

## Slide 13
Surge detection
Interlock :
3. IT String test bench 4/4
Cold Compressor - Control panel
Cold Compressor operation & control system

## Slide 14
CONCLUSION 1/2
Today status :
Test benches upgraded & validated
Cryogenics infrastructure upgraded
Cryogenics process optimized
LHe @ 4.5 K demand (2024)
Max Liquefaction
capacity: 35 g/s

## Slide 15
CONCLUSION 2/2
Next challenge :
Operation of the IT String test bench with magnets and SC link, scheduled to start in September 2025.
Thank you!

## Slide 16
Thank you!

## Slide 17
Test Facility building (SM18) – Presentation
SM18: CERN’s major test facility for superconducting magnets, power links, and radio-frequency cavities
Precooling / Warming-up [300-80K]
Comp.
Comp.
Comp.
CWU1
GHe / LN2
CWU2
GHe / LN2
[4.5 K ; 3.5 bar]
[4.5 K
1.6 bar]
[300 - 80 K ]
[20 K ; 1.2 bar]
Pure Warm He gas recovery
precooler
ImPure Warm He gas recovery (purges)
3 Compressors [1 - 170 bar]
2 Purifiers
8 GHe buffers 80 m3
Quench buffer
80 m3
Dewar
LHe
25 kL
5 Vertical MAGNET test benches
4 Vertical RF Cavities test benches
2 RF Cryo-module test benches (in bunker)
New Cold Compressor Unit
(for IT-STRING)
Back-up
Pumping Line [ ~30 K ; ~15 mbar ]
IT-STRING Pumping Line [ ~3.3 K ; ~15 mbar ]
New CRPL-B
Compressor Station
Cold Box
[35 g/s]
CRPL-A (Cryoplant)
Compressor Station
Cold Box
[25 g/s]
Warm Pumping Unit 1
Primary pumps
Roots pumps
Warm Pumping Unit 2
Primary pumps
Roots pumps
New IT-String : 6 magnets + SC-Link
8 Horizontal Magnet & SC-Link test benches
CWS
CWU

## Slide 18
nicolas.guillotin@cern.ch
ICEC28-ICMC2022, April 25th – 29th, 2022
2
New Compressor station (CS-B)
New cold box (CB-B)
Cryogenics Infrastructure upgrade : new cryogenics plant (Linde B)
MAIN CARACTERISTICS
1 motor stahl (1.6 MW)
1 double stage compressor Mycom
8 turbines
8 exchangers
2 adsorbers 80 K
1 adsorber 20 K
No precooler LN2
No phase separator
Inlet flow at full capacity : 350 g/s
Liquefaction flow : 35 g/s (Lhe @ 4.5 K, 1.6 bar)
  Installation (Q2 2019 – Q4 2019)
  Commissioning (Q4 2019 – Q3 2020)
Operation (Q4 2020 – now)

## Slide 19
1. Cryogenics Infrastructure upgrade : 25 kl DEWAR ISSUE
Event (June 2021): Flooding in the pit – Dewar tilted by 10°
    1.Emptying
    2.Dewar lifted to maintain Center of gravity to avoid swinging.
    3.Water drained from the pit.
    4.Investigation: dye penetrant test + high resolution 3D laser scan of the neck, endoscopy => experts agreed damage is minor – OK for re-use.
    5. Leak test + pressure test
    6. Back in service in August 2024

## Slide 20
2.0 Radio Frequency (RF) Crab Cavities Cryogenics Module (C3M) upgrade
Introduction & Context
Crab cryo-module cavities' principle
Crab cryo-module cavities location in future LHC