## Slide 1: Design and Development of a State-of-the-Art Cryogenic Testing Platform for Liquid Hydrogen Applications
Presented by: Dr. Shanaka Kristombu1,
Senior Research Fellow,
Sir Winston Churchill Fellow on H2 Storage Infrastructure,
1 The University of Melbourne, Melbourne VIC 3000, Australia

## Slide 2: Testings Need for Large and Small LH2 Storage Tank and Material Development
Mechanical Properties of Materials and Component up 20 K
      Elastic and Strength Properties
      Stress-Strain Curves/Fracture Toughness/Ductility
      Thermal Cycles and Mechanical Cycles Induced Fatigue
    Exposure to LH2 and Cryogenic H2 Gas at Different T and P
      H2 Leak and Permeation through Materials and Connections
      Change in Mechanical Properties, e.g. H2 embrittlement

## Slide 3: Testings Need for LH2 Storage Tank and Material Development
    Insulation and Thermal Properties of Insulation and Structural Materials at 20 K
      Insulation Properties of Vacuumed Spaces with Infill Materials
      Heat Losses through Supports and Components
    Characterising and Managing Boil-off Gas
      Integrated Refrigeration
      Re-liquefaction of Boil-off Gas
      Thermal Intercepts using Vapor Barriers or Active Cooling Supports

## Slide 4
Cost Effective Cryogen Free-Mechanical Testing Cryostats to Characterise Mechanical Properties and Leak

## Slide 5: Mechanical Testing Cryostat at 20 K (in commissioning stage)
343 mm
925 mm

## Slide 6: Required Capacity of the Jig System
Metals (ASTM E8M Subsized Samples)
Composites (ASTM D3039/3039M)
~50 kN
~50 kN
100mmx6mmx6mm (L x W x t)
100mmx6mmx4mm (L x W x t)
Recommended – 250mm x 15 mm x 1 mm for fibre orientation. But no restrictions on the dimensions
We proposing using 115mm x 15 mm x 1 mm
~80 kN

## Slide 7: Sample Sizes
Standards ASTM E8/E8M - Subsize Sample size – Metals - 100mm x 6 mm x 4-6 mm (thickness))
ASTM D3039/D3039M – Composite - 115mm x 15 mm x 1 mm (thickness)
ASTM D638 - Plastics - 115mm x 19 mm x 6 mm (thickness)
127 mm
89 mm
231 mm

## Slide 8: Cryostat Thermal Modelling – Finite Element Method and Analytical Approaches
The second stage of the cryocooler can go up to 10 K, matching heat loss and cryocooler cooling capacity Steady State Equilibrium Temperatures of Cryocooler
Stage 1 = ~ 50 K (~ 21 W)
Stage 2 = ~ 10 K (~ 1 W)
Chart: Stage 2 Heat loss at 20 K [W]
| Conduction | Free molecular flow | Radiation
Heat loss | 1.9 | 0.18 | 0.01
Cooling Power at 20 K = 16 W
Effective Cooling Power at 20 K = 16 W – 2 W (Heat Loss)
= 14 W

## Slide 9: Trial and Validation
The sample chamber goes up to 20 K after 5 hrs.
Cryocooler's second stage can go up to 13 K, and the minimum sample temperature is 18K.

## Slide 10: Cryogenic Mechanical Testing at 20 K (-253 C)
Custom Made Cryostat using 4 K Cryocooler
Capabilities:
Tensile, compression and fatigue testing of material up to 20 K
Expose material to H2 to study cryogenic H2 embrittlement
Thermal cycling testing with and without loading (5 hr cooling cycle)
Combined thermal and loading cycle testing to investigate fatigue failures
Pros
No need for liquid He
Cost effective: Electricity cost < AUD 50
Less cryogenic safety requirement
Cons
Limited load capacity
Cooling time high – 1 or 2 samples per day
Maintenance cost – replace cryocooler after a few years

## Slide 11: Investigating Microcracking and Capturing Strain using the DIC Technique (Under Development)
Carbon fibre – epoxy microcracking with nano-clay (Timmerman et al., n.d.)

## Slide 12: Hydrogen Leak and Permeation Testing at Cryogenic Temperature and Stressed Conditions (Design Stage)
Hydrogen Leak Testing Apparatus
This apparatus will be able to test
  H2 leak due to microcracking of composite tanks
  Vacuum failure and its effects on the insulation and thermal failure
  Performance of different welding/connection for H2 leak
  Insitu testing with expose to H2

## Slide 13
Thermal Insulation Testing to Desing Insulation System for LH2 Storage Applications

## Slide 14: Insulation Testing at 77 K
Flat plate boiloff apparatus – ASTM C1774
Insulation testing cryostat – CS 500
(Fesmire et al., 2015b; ASTM C1774)

## Slide 15: Cryogenic Insulation Systems for 20 K vs 77 K
Thermal conductivity of insulation materials with pressure (Fesmire, 2019)

## Slide 16: Vacuumed Insulation Testing at 20 K
Equipment can measure thermal conductivity at 20 K with a different warm boundary temperature up to 77 K
Combining CS500 results, thermal conductivity values for the 20K to 300 K temperature boundary can be calculated
The method is based on a temperature gradient method using in-built heater
So, There is no need to calculate the cooling power of the cryocooler or the heating power of the heater
Sample and Thermal Measurements Space
Cryo-cooler
Vacuum Chamber

## Slide 17: Vacuumed Insulation Testing at 20 K:
Heat flux through the sample
Temperature map for sample
Heat flux in sample space
Accuracy will be dependent on the sensitivity of the thermal measuring gauges
There are limitations dependent on cryocooler capacity, thermal gradient, and sensitivity of temperature sensors
Use CS 900 Cryostat as final validation

## Slide 18: Sample and Testing Specifications
 | Specification
Type of apparatus | Temperature Gradient Based Electrical Method
Specimen types | Bulk fills, Composites, Powders, Blankets, Flat panels, Sheets, Foams, Multilayer insulation (MLI) system, Foams
Specimen size | Diameter – 200 mm and Thickness – 50 mm
Thermal Boundary | Min Cold Boundary Temperature (CBT) – 20K (cool down using cryocooler) Min Warm Boundary Temperature (WBT) – 77K (limited when P is less than 10 Pa) 300 K
Vacuum Pressure | High vacuum: 10-3 Pa - 10-1 Pa (Limited to MLIs) Medium/Soft vacuum :10-1 Pa – 100 Pa Rough vacuum : 100 Pa – 100,000 Pa
Test limitations | Effective Thermal conductivity (ke) – approximately in the range of 0.1 mK/m-K

## Slide 19
Lab Scale Hydrogen Liquefier and Boil-off Management Methods

## Slide 20: Overview of the Magnetic Refrigeration Cryostat
Heat Exchangers for He
HTS Solenoid
Magnetocaloric Material Bed
Heat Exchanger for HTS
Test Capabilities
Test different magnetocaloric materials and AMR beds
Test liquefaction under different boil-off conditions and efficiencies
Different heat-exchanging mechanisms
1 W cooling capacity at 20 K
4 g/hr LH2 (~60 ml/hr)
Refrigeration using Magneto-caloric effect (MCE) with the potential to achieve higher efficiency > 50%

## Slide 21: Funding Support

## Slide 22: ARENA Research Team
Dr. Sadeep Thilakarathna (UoM)
Prof. Priyan Mendis (UoM)
A/Prof Kevin Li
(UoM)
Mr James Fesmire (GenH2)
Dr William Notardonato (Eta Space)
Dr Luke Djukic
(Omni Tanker)
Mr Susiri Costa (UoM)
Mr Amila Premakumara (UoM)
Mr Upeka Gunarathne (UoM)
Dr Shanaka Baduge (UoM)
PhD Candidates
Mr Harindu Lochana(UoM)

## Slide 23: References
International Energy Agency (IEA) (2019) Global hydrogen demand by sector in the Net Zero Scenario, 2020-2030. Available at: https://www.iea.org/data-and-statistics/charts/global-hydrogen-demand-by-sector-in-the-net-zero-scenario-2020-2030.
DCCEEW (2022). State of Hydrogen 2022. [online] Available at: https://www.dcceew.gov.au/sites/default/files/documents/state-of-hydrogen-2022.pdf [Accessed 2 Dec. 2023].
IRENA (2022). Global Hydrogen Trade to Meet the 1.5°C Climate Goal. [online] IRENA. Available at: https://www.irena.org/publications/2022/Apr/Global-hydrogen-trade-Part-II [Accessed 2 Dec. 2023].
Airbus (2021). Exploring hydrogen - The impact, challenges and options. [online] Airbus. Available at: https://aircraft.airbus.com/en/newsroom/news/2021-09-exploring-hydrogen [Accessed 2 Dec. 2023].
HESC (2023). Commercial Demonstration Project. [online] HESC. Available at: https://www.hydrogenenergysupplychain.com/about-hesc/ [Accessed 2 Dec. 2023].
IEA (2023). Global Hydrogen Review 2023 – Analysis. [online] IEA. Available at: https://www.iea.org/reports/global-hydrogen-review-2023 [Accessed 2 Dec. 2023].
Ghafri, S.Z.A., Munro, S., Cardella, U., Funke, T., Notardonato, W., Trusler, J.P.M., Leachman, J., Span, R., Kamiya, S., Pearce, G., Swanger, A., Rodriguez, E.D., Bajada, P., Jiao, F., Peng, K., Siahvashi, A., Johns, M.L. and May, E.F. (2022). Hydrogen liquefaction: a review of the fundamental physics, engineering practice and future opportunities. Energy & Environmental Science, [online] 15(7), pp.2690–2731. doi:https://doi.org/10.1039/D2EE00099G.
CB&I (2022). Liquid Hydrogen Storage Technologies. [online] Available at: https://www.energy.gov/sites/default/files/2022-03/Liquid%20H2%20Workshop-CBandI.pdf [Accessed 2 Dec. 2023].
‌Shell International B.V. (2022). Hydrogen’s role in Shell’s Journey - Liquid Hydrogen in Emerging Large-Scale Markets. [online] Available at: https://www.energy.gov/sites/default/files/2022-03/Liquid%20H2%20Workshop-Shell.pdf [Accessed 2 Dec. 2023].
Ratnakar, R.R., Gupta, N., Zhang, K., van Doorne, C., Fesmire, J., Dindoruk, B. and Balakotaiah, V. (2021). Hydrogen supply chain and challenges in large-scale LH2 storage and transportation. International Journal of Hydrogen Energy. doi:https://doi.org/10.1016/j.ijhydene.2021.05.025.
GIIGNL (2022). The LNG industry GIIGNL Annual Report. [online] Available at: https://giignl.org/wp-content/uploads/2022/05/GIIGNL2022_Annual_Report_May24.pdf [Accessed 2 Dec. 2023].
Sun, Z., Li, M., Qu, Z., Tian, D. and Zhang, J. (2023). A quasi-2D thermodynamic model for performance analysis and optimization of liquid hydrogen storage system with multilayer insulation and vapor-cooled shield. Journal of Energy Storage, [online] 73, p.109128. doi:https://doi.org/10.1016/j.est.2023.109128.
Fesmire, J.E. (2015). Standardization in Cryogenic Insulation Systems Testing and Performance Data. Physics Procedia, 67, pp.1089–1097. doi:https://doi.org/10.1016/j.phpro.2015.06.205.
Swanger, A. (2018). LARGE SCALE CRYOGENIC STORAGE WITH ACTIVE REFRIGERATION. [online] Available at: https://www.researchgate.net/publication/329071191_LARGE_SCALE_CRYOGENIC_STORAGE_WITH_ACTIVE_REFRIGERATION [Accessed 2 Dec. 2023].
Fesmire, J., Swanger, A., Jacobson, J. and Notardonato, W. (2022). Energy efficient large-scale storage of liquid hydrogen. IOP Conference Series: Materials Science and Engineering, 1240(1), p.012088. doi:https://doi.org/10.1088/1757-899x/1240/1/012088.
Timmerman, J. F., Hayes, B. S., & Seferis, J. C. (n.d.). Nanoclay reinforcement effects on the cryogenic microcracking of carbon fiber/epoxy composites. www.elsevier.com/locate/compsci
‌
‌
‌
‌
‌

## Slide 24: Copyright Notice
(c) The University of Melbourne
2023
This material is protected by copyright. Unless indicated otherwise, the University of Melbourne owns the copyright subsisting in the work. Other than for the purposes permitted under the Copyright Act 1968, you are prohibited from downloading, republishing, retransmitting, reproducing or otherwise using any of the materials included in the work as standalone files. Sharing University teaching materials with third-parties, including uploading lecture notes, slides or recordings to websites constitutes academic misconduct and may be subject to penalties. For more information: Academic Integrity at the University of Melbourne
Requests and enquiries concerning reproduction and rights should be addressed to the University Copyright Office, The University of Melbourne: copyright-office@unimelb.edu.au
The University of Melbourne (Australian University) PRV12150/CRICOS 00116K

## Slide 25: Thank You
[Copyright, CTA or other information/notices]
The University of Melbourne (Australian University) PRV12150/CRICOS 00116K