## Slide 1: Magnetic measurements at ambient temperature on the MBRD2
Lucio Fiscarelli, Mariano Pentella, TM and LMF colleagues
Alessandra Pampaloni, and INFN Genova
WP3 Meeting – 10 July 2024

## Slide 2: Outline
Measurement setup
Performed tests
Results
Conclusions

## Slide 3: Measurement setup
Measurements in b.180
Rotating coil scanner
  Overall diameter 94 mm
  PCB coils
  Measurement radius 35 mm
  Active length 600 mm
  Positioning accuracy ~0.1 mm
  (wire encoder)
  On board tilt sensor
  14 positions to cover the full length of the MBRD
  Reference radius is 35 mm

## Slide 4: Performed tests
Full scan of both apertures at ±15 A
Check at [±5, ±10, ±15, ±18] A on a few positions of V2

## Slide 5: Results: main field
 | V1 |  |  |  |  | V2 |  |  | 
 | CS | CENTER | NCS | INTEGRAL |  | CS | CENTER | NCS | INTEGRAL
Positions | 1-2 | 3-12 | 13-14 |  |  | 1-2 | 3-12 | 13-14 | 
Current (A) | 15 |  |  |  |  | 15 |  |  | 
B1 | 0.0047 | 0.0062 | 0.0047 | 0.0486 |  | 0.0047 | 0.0062 | 0.0047 | 0.0486
TF | 0.3131 | 0.4152 | 0.3120 | 3.2413 |  | 0.3121 | 0.4153 | 0.3126 | 3.2413
30 units more than MBRD1 that was measured at 10 A

## Slide 6: Results: bn
 | V1 |  |  |  |  | V2 |  |  | 
 | CS | CENTER | NCS | INTEGRAL |  | CS | CENTER | NCS | INTEGRAL
Positions | 1-2 | 3-12 | 13-14 |  |  | 1-2 | 3-12 | 13-14 | 
Current (A) | 15 |  |  |  |  | 15 |  |  | 
b2 | -16.19 | -1.80 | -20.51 | -5.63 |  | 22.45 | 3.27 | 22.54 | 7.72
b3 | -3.18 | 2.33 | -5.81 | 0.75 |  | -5.15 | 1.93 | -6.11 | 0.18
b4 | 3.19 | -0.66 | 0.24 | -0.11 |  | -2.07 | 0.90 | 0.37 | 0.49
b5 | 4.11 | 4.51 | -0.55 | 3.88 |  | 3.71 | 4.02 | -1.09 | 3.39
b6 | 1.66 | 0.49 | 0.70 | 0.65 |  | -1.36 | -0.36 | -0.52 | -0.49
b7 | 0.86 | 2.23 | -0.69 | 1.74 |  | 0.87 | 2.11 | -0.85 | 1.63
b8 | 1.40 | 1.12 | 1.19 | 1.16 |  | -1.29 | -1.03 | -1.02 | -1.06
b9 | 0.33 | 1.26 | 0.00 | 1.01 |  | 0.24 | 1.23 | -0.09 | 0.96
b10 | 0.08 | 0.02 | -0.02 | 0.02 |  | 0.04 | -0.03 | 0.11 | 0.00
b11 | -1.40 | -1.43 | -1.67 | -1.45 |  | -1.36 | -1.51 | -1.68 | -1.51
b12 | 1.17 | 1.34 | 1.11 | 1.29 |  | -1.06 | -1.41 | -1.08 | -1.33
b13 | -1.47 | -1.61 | -1.53 | -1.59 |  | -1.45 | -1.68 | -1.56 | -1.64
b14 | 0.62 | 0.71 | 0.59 | 0.69 |  | -0.54 | -0.75 | -0.55 | -0.71
b15 | -0.55 | -0.61 | -0.56 | -0.60 |  | -0.60 | -0.61 | -0.61 | -0.61

## Slide 7: Results: an
 | V1 |  |  |  |  | V2 |  |  | 
 | CS | CENTER | NCS | INTEGRAL |  | CS | CENTER | NCS | INTEGRAL
Positions | 1-2 | 3-12 | 13-14 |  |  | 1-2 | 3-12 | 13-14 | 
Current (A) | 15 |  |  |  |  | 15 |  |  | 
a2 | 0.97 | -0.88 | -0.95 | -0.67 |  | -4.47 | 2.24 | -0.01 | 1.20
a3 | -8.97 | -1.81 | -0.88 | -2.53 |  | -8.97 | -2.18 | -1.46 | -2.88
a4 | 0.39 | -0.70 | 0.21 | -0.47 |  | 0.71 | 0.56 | 0.64 | 0.59
a5 | -1.66 | -0.82 | -0.43 | -0.88 |  | -2.25 | -1.03 | -0.77 | -1.14
a6 | 0.31 | -0.22 | 0.57 | -0.07 |  | 0.20 | -0.06 | -0.02 | -0.02
a7 | -0.21 | -0.31 | -0.14 | -0.27 |  | -0.68 | -0.23 | -0.25 | -0.28
a8 | 0.12 | -0.02 | 0.14 | 0.01 |  | 0.15 | -0.16 | -0.09 | -0.12
a9 | 0.02 | -0.09 | 0.03 | -0.07 |  | -0.21 | -0.14 | -0.06 | -0.14
a10 | 0.06 | 0.02 | 0.18 | 0.05 |  | 0.04 | 0.11 | 0.02 | 0.09
a11 | -0.16 | -0.21 | -0.31 | -0.21 |  | -0.30 | 0.00 | -0.15 | -0.05
a12 | 0.05 | 0.15 | 0.30 | 0.16 |  | -0.01 | 0.08 | -0.08 | 0.05
a13 | -0.02 | -0.19 | -0.21 | -0.17 |  | -0.09 | -0.06 | -0.13 | -0.07
a14 | 0.03 | 0.10 | 0.18 | 0.10 |  | 0.01 | 0.03 | -0.05 | 0.02
a15 | -0.03 | -0.04 | 0.01 | -0.04 |  | -0.12 | -0.12 | -0.12 | -0.12

## Slide 8: Multipoles versus position
V1
V2

## Slide 9: Comparison with measurements from ASG
 | V1 |  |  |  |  | V2 |  |  | 
 | CS | CENTER | NCS | INTEGRAL |  | CS | CENTER | NCS | INTEGRAL
Δ b2 | 4.40 | 1.97 | 8.27 | 2.04 |  | -7.03 | -2.16 | -8.09 | -2.36
Δ b3 | 0.78 | -0.35 | 1.86 | -0.46 |  | 1.00 | 0.02 | 2.93 | -0.08
Δ b4 | -0.81 | 0.72 | 0.82 | 0.68 |  | 1.30 | 0.72 | 1.27 | 0.67
Δ b5 | 0.82 | 0.87 | 1.78 | 0.50 |  | 0.49 | 0.79 | 2.47 | 0.55
Δ b6 | -0.18 | 0.19 | 0.76 | 0.23 |  | 0.99 | 0.54 | 0.70 | 0.58
Δ b7 | 0.22 | -0.30 | -0.23 | -0.47 |  | 0.43 | -0.47 | 0.36 | -0.45
Δ b8 | -0.94 | -0.90 | -0.90 | -0.90 |  | -1.66 | -0.96 | -1.11 | -1.00
Δ b9 | -0.35 | -0.41 | -0.59 | -0.56 |  | 0.02 | -0.08 | 0.39 | -0.14
Δ b10 | -0.92 | -0.65 | 1.09 | -0.39 |  | 0.30 | 0.42 | 0.59 | 0.42
Δ b11 | 1.29 | 1.20 | -0.08 | 1.03 |  | 2.73 | 1.76 | 2.20 | 1.97
Δ b12 | 0.83 | 0.83 | 2.78 | 0.98 |  | 2.16 | 1.54 | 1.80 | 1.70
Δ b13 | 1.56 | 1.46 | 0.57 | 1.36 |  | 3.23 | 1.75 | 2.08 | 2.05
Δ b14 | 0.60 | 0.70 | 1.20 | 0.67 |  | 0.58 | 0.40 | 0.28 | 0.44
Δ b15 | 0.03 | 0.16 | 0.03 | 0.15 |  | 0.11 | 0.19 | 0.06 | 0.18
Differences ASG (18 A) – CERN (15 A)

## Slide 10: Comparison with measurements from ASG
 | V1 |  |  |  |  | V2 |  |  | 
 | CS | CENTER | NCS | INTEGRAL |  | CS | CENTER | NCS | INTEGRAL
Δ a2 | -0.70 | -0.52 | -0.42 | -0.38 |  | 5.48 | -0.52 | 0.60 | 0.14
Δ a3 | 1.90 | -0.43 | -0.73 | -0.25 |  | 0.59 | -0.81 | -0.70 | -0.67
Δ a4 | -1.24 | -1.36 | -0.25 | -0.98 |  | -1.05 | -1.52 | -0.31 | -1.20
Δ a5 | 0.82 | 0.48 | 0.29 | 0.50 |  | 1.13 | 0.74 | 1.06 | 0.83
Δ a6 | 0.66 | 0.17 | 1.09 | 0.44 |  | 0.48 | 0.00 | 1.19 | 0.28
Δ a7 | 0.60 | 0.51 | 0.57 | 0.53 |  | 2.63 | 1.74 | 2.04 | 1.83
Δ a8 | 0.08 | -0.38 | 0.21 | -0.18 |  | 0.11 | -0.65 | -0.21 | -0.41
Δ a9 | -0.25 | -0.39 | -1.36 | -0.48 |  | -1.52 | -0.61 | -0.91 | -0.76
Δ a10 | -0.33 | -0.07 | 0.38 | -0.03 |  | -0.29 | 0.10 | 0.24 | 0.04
Δ a11 | -1.79 | -1.96 | -2.69 | -1.92 |  | -0.83 | 0.39 | 0.24 | 0.14
Δ a12 | 0.41 | 0.60 | 0.86 | 0.57 |  | 0.67 | 1.15 | 1.42 | 1.04
Δ a13 | -0.17 | -0.43 | -0.40 | -0.34 |  | 0.65 | 1.32 | 1.48 | 1.17
Δ a14 | 1.52 | 1.39 | 1.29 | 1.31 |  | 2.58 | 1.70 | 2.02 | 1.79
Δ a15 | 1.34 | 1.16 | 1.29 | 1.15 |  | 0.53 | 1.00 | 1.09 | 0.89
Differences ASG (18 A) – CERN (15 A)

## Slide 11: V1
ASG
CERN

## Slide 12: V2
ASG
CERN

## Slide 13: Change of TF and b2 versus current
35 units

## Slide 14: Conclusions
  MBRD2 measured at CERN, before integration in the cold-mass, at ambient temperature
  Main results
    Transfer function in line with MBRD1
    b2 is -5.6 / 7.7 units
    b3 is 0.8 / 0.2 units
    b5 is 3.9 / 3.4 units
    General agreement with measurements performed at ASG
    Change of TF (~30 units) and b2 (~3 units) from 5 A to 18 A
      Same effect already seen on MBRD1