## Slide 1: Frontend ASIC development for the advanced SiPM camera
Luca Giangrande, Matthieu Heller
FACULTÉ DES SCIENCES
SECTION DPNC

## Slide 2: ASIC features
FACULTÉ DES SCIENCES
SECTION DPNC
Design requirements:
Linear charge transfer in the range [1:250] photons
SNR ≥ 5
FWHM ≤ 3 ns
Recovery time < 30 ns
Power consumption per pixel < 40 mW
Output load: capacitive ≤ 5 pF plus resistive 50 Ω (impedance matching)
Technology node 65nm CMOS
Test features:
Adjustable channel gain, bandwidth and pulse shape
Multiple test points along the channel

## Slide 3: Silicon Photo Multiplier
Hamamatsu LCT5-UVE hexagonal shape with 75x75 μm2 micro cell.
Pixel divided into 4 quarters (each with ~4093 μ-cells).
Each quarter is readout by a dedicated pre-amplifier then the 4 signals are actively added together
Pixels organized in modules of 6 (max distance between frontend and sensor of ~ 4.5 cm).
FACULTÉ DES SCIENCES
SECTION DPNC
Pixel
Two modules
36.9mm
44.2mm
Group A
Group B
ASIC

## Slide 4: FANATIC architecture
Fast ANAlog Transimpedance amplifier Integrated Circuit:
ASIC readout for 8 pixels:
  4 analog inputs per output channel
  Active summation performed in voltage mode
  1 analog output per pixel: 4 channels with single-ended and 4 channels with pseudo-differential outputs
  Adjustable AC coupling capacitors for output pulse shaping
SiPM
FACULTÉ DES SCIENCES
SECTION DPNC

## Slide 5: FANATIC SiPM-test PCBs
FANATIC PCB
GPIO - FPGA PCB
LCT5 UVE
LVR3
FACULTÉ DES SCIENCES
SECTION DPNC

## Slide 6: FANATIC Measurement setup
FACULTÉ DES SCIENCES
SECTION DPNC

## Slide 7: Neutral density filters calibration
Laser source 375 nm: PicoQuant PDL 800-D
Photodiode: Hamamatsu S1337-1010BQ
Motorized filter wheel: Thorlabs FW212C
Neutral density filters: Thorlabs FW212CNEB
Thorlabs integrating sphere
FACULTÉ DES SCIENCES
SECTION DPNC

## Slide 8: Measured pulse
Averaged pulse with laser source
Peak gain: 2.2 mV/p.e.
Peak to peak gain: 3.14 mV/p.e.
Pulse integral (6 ns) gain: 7.7 V*ns/p.e.
Peaking time: 1.1ns
FWHM: 2.9 ns
Recovery: 30.5 ns
FACULTÉ DES SCIENCES
SECTION DPNC

## Slide 9: Data Analysis
Low intensity histogram fit with a generalized Poisson distribution function. Pulse integral over a fixed window 6ns. Transfer characteristics from multiple data sets employing different laser intensities.
FACULTÉ DES SCIENCES
SECTION DPNC

## Slide 10: Measurements summary
Single ended channel performance with 3x3 mm2 LCT5 UVE and 375nm laser source
Linear characteristic (pulse intergral) up to 200 p.e.
Dynamic range 1000 p.e.
 | Pulse Integral 6ns | Pulse peak | Pulse peak2peak
1 p.e. signal | 7.7 mV*nS | 2.2 mV | 3.14 mV
Noise | 1.5 mV*ns | 0.4 mV rms | 0.4 mV rms
SNR | 5.1 | 5.5 | 7.85
FACULTÉ DES SCIENCES
SECTION DPNC

## Slide 11: PRESSEC pre-amplifier ASIC (UB-Unige-UPC)
Preamplifier Readout Electronics for Summing SiPMs Enhanced Circuit
Prototype readout for 4 pixels:
  4 current mode analog inputs per pixel
  3 signal processing channels per pixel: high-gain, low-gain and NSB integration
  1 analog output per pixel (high gain or low gain path)
  1 digital output per ASIC providing the NSB serialized data of all pixels
FACULTÉ DES SCIENCES
SECTION DPNC

## Slide 12: Further developments
Complete the test campaign:
  Acquisition and analysis of differential output to be done, but no major performance difference is expected
  Measurement campaign with larger pixel (LVR 3 6x6 mm2), to be completed. Expected degradation in FWHM and noise from preliminary results
  Temperature and supply voltage dependence
Adapt the buffer designs for PRESSEC (ASIC designed together with UB and UPC)
Technical paper in preparation
FACULTÉ DES SCIENCES
SECTION DPNC

## Slide 13: Backup slides
FACULTÉ DES SCIENCES
SECTION DPNC

## Slide 14
Pseudo-differential frontend schematic
20 Ω
700 pF
873 Ω
873 Ω
8.8 kΩ
4.4mA
2.2mA
700 Ω
3.2 kΩ
350 Ω
4 kΩ
800 Ω
950 Ω
800 Ω
950 Ω
2.2mA
1.3mA
1.4mA
4.4mA
PCB ASIC
FACULTÉ DES SCIENCES
SECTION DPNC

## Slide 15
Pseudo-differential frontend schematic
FACULTÉ DES SCIENCES
SECTION DPNC

## Slide 16
Single-ended frontend schematic
FACULTÉ DES SCIENCES
SECTION DPNC

## Slide 17: FANATIC layout
Differential channels
Single ended channels
Digital state registers
3.58 mm
FACULTÉ DES SCIENCES
SECTION DPNC

## Slide 18: FANATIC measurements with SiPM
Finger plot of single ended channel with laser source and 2x LCT5-UVE 3x3 mm2
FACULTÉ DES SCIENCES
SECTION DPNC
Noise: 260 uV rms
FWHM: 3 ns