## Slide 1: Sensitivity, response and thresholds; first experience during scrubbing and intensity ramp up
17/04/2024
Sara Morales on behalf of the SY-BI-BL section
Joint MPP, ColUSM and BLMTWG

## Slide 2: Outline
Sensitivity of new detectors in the wall (RS09)
  Top energy
  Injection energy/scrubbing
Losses during injection (RS01 and RS09)
New response factors from betatron lossmaps (RS09)
New BLM thresholds families

## Slide 3: Sensitivity of new detectors on the wall
Comparison of signal of “wall” and “old” BLM vs lost intensity from BLM calibration
INJPHYS beam modes a minimum of 1e12 protons per beam in the machine (scrubbing included)
5e9 - pilot
5e9 - pilot
5e9 - pilot
In all cases signal in new BLMs visible from at least losses around 1e6 protons/s

## Slide 4
Comparison of signal of “wall” and ”old” BLM vs lost intensity from BLM calibration
FLATTOP beam modes with a minimum of 1e12 protons per beam in the machine (intensity ramp-up from 12 bunches included)
Sensitivity of new detectors on the wall
In all cases signal in new BLMs visible from losses well below 1e6 protons/s

## Slide 5: Sensitivity of new detectors on the wall
Comparison of signal of “wall” and ”old” BLM vs lost intensity from BLM calibration
SQUEEZE beam modes with a minimum of 1e12 protons in the machine (intensity ramp-up from 12 bunches included)
In all cases signal in new BLMs visible from losses around 3e5 protons/s
500kW -> 5.8e11 protons/s
6 orders of magnitude bewlo max loss allowed for RS09

## Slide 6: Losses during injection
2023
2024

## Slide 7: Maximum loss vs number of bunches
2023
2024

## Slide 8: Response from lossmaps for new detectors in the wall
Response analyzed from 2024 betatron lossmaps with final collimation settings at:
  INJECTION: INJPROT IN and OUT
  TOP energy: FLATTOP, EOS, QCHANGE, and during collisions with 6 different beta star values
  Reduction in response consistent with primary observations by Belen
Response for B2 losses much lower -> We keep only B1 lossmaps

## Slide 9: Response from lossmaps for new detectors in the wall
Response analyzed from 2024 betatron lossmaps with final collimation settings at:
  INJECTION: INJPROT IN and OUT
  TOP energy: FLATTOP, EOS, QCHANGE, and during collisions with 6 different beta star values
  Reduction in response consistent with primary observations by Belen
We keep only the maximum from each set of B1 lossmaps

## Slide 10: Response from lossmaps for new detectors in the wall
Response analyzed from 2024 betatron lossmaps with final collimation settings at:
  INJECTION: INJPROT IN and OUT
  TOP energy: FLATTOP, EOS, QCHANGE, and during collisions with 6 different beta star values
  Reduction in response consistent with primary observations by Belen
We compute the mean for INJECTION and TOP separately, convert to aC/p and do a linear fit with energy to get response curve

## Slide 11
Response analyzed from 2024 betatron lossmaps with final collimation settings at:
  INJECTION: INJPROT IN and OUT
  TOP energy: FLATTOP, EOS, QCHANGE, and during collisions with 6 different beta star values
  Reduction in response consistent with primary observations by Belen
Response from lossmaps for new detectors in the wall

## Slide 12
Family Name | 2022 loss map values per proton [aC, aC/TeV] | Family Name | 2024 loss map values per proton [aC, aC/TeV] | Ratio 2024 / 2022 [offset, slope]
THRI_COL_7_TCPPM | [66,44] | THRI_COLL_7_TCPPM_WALL | [3.8, 4.3] | [0.06,0.098]
THRI_COLL_7_TCP | [ 49, 47 ] | THRI_COLL_7_TCP_WALL | [4.2, 6.4] | [0.09,0.14]
THRI_COLL_7_TCLA_HI | [ 6, 88 ] | THRI_COLL_7_TCLA_HI_WALL | [1.7, 6.1] | [0.3,0.07]
Proposal of BLM thresholds families for detectors in the wall
Maximum allowed power loss from loss maps WITHOUT these BLMs:
H1 : 322.5 kW
V1 : 309.8 kW

## Slide 13
Conclusions
Sensitivity of new BLMs in the wall:
  Injection: capable of observing losses from 1e6 protons/s
  Top Energy: capable of observing losses below 3e5 protons/s
Beam losses at injection seem to be of similar order than in 2023 for 144 bunches
New response factors calculated from loss maps for the 3 new families