## Slide 1: Magnetic measurements on the LMQXFA01 at 1.9 K
Lucio Fiscarelli, Guy Deferne, Mariano Pentella, Piotr Rogacki,
and all TM colleagues
12 September 2024

## Slide 2: Outline
Measurement setups: Stretched Wire and Rotating Coil
Performed tests and results: Nominal current, Machine cycles, Stair-step cycles
Additional results: voltage spikes
Conclusions
Results of alignment tests are nor reported here since test are still ongoing.

## Slide 3: Stretched wire
Standard system used for Q2 and other magnets
PI X-Y tables
FDI integrators
FFMM control software
Copper beryllium wire ⌀ 0.125 mm
Wire length ~21 m
A careful sag correction is required

## Slide 4: Rotating coils
Same system used on the Q2
6-segment rotating shaft
Each segment is 1.3 m + 0.1 m gap
12 FDI integrators
FFMM software for control and online processing
Continuous measurement at 1 turn/s
Not optimal for Q3

## Slide 5: Rotating coils
Anticryostat flange
Longitudinal center
Shaft too short to cover at the same time both magnets
Two measurement positions
Cycles repeated twice

## Slide 6: Performed tests
Performed tests
Integrated gradient with Stretched wire
  Nominal current
Field quality with Rotating Coils
  Machine cycles and Stair-step cycles

## Slide 7: Main field and magnetic length
Current | A | 16230 |  |  |  | 
 |  | A03 | A04 | Integral |  | 
CERN |  |  |  |  |  | 
Integrated gradient | T | 559.53 | 559.62 | 1119.15 | SSW | 1119.03
Central gradient TF | T/m/kA | 8.1790 | 8.1828 |  |  | 
Lm | m | 4.215 | 4.214 |  |  | 
Nodal distance | m | 4.770 |  |  |  | 
FNAL |  |  |  |  |  | 
Integrated gradient | T | 559.70 | 559.95 | 1119.60 |  | 
Difference | units | 8 | 1 | 4 |  | 
Nodal distance | m | 4.772 |  |  |  | 
Difference | mm | -2 |  |  |  | 

## Slide 8: Sag correction on stretched wire
The gradient measured with a vertical displacement (GY) is affected both by gravity and wire diamagnetism: correction is more difficult
The gradient measured with a horizontal displacement (GX) is affected mainly by the wire diamagnetism: more effective extrapolation

## Slide 9: Field quality at nominal 16230 A
Current 16230 A |  |  |  |  |  |  |  | 
 | A04 | A03 | Integral* |  |  | A04 | A03 | Integral*
b3 | 1.39 | 0.46 | 0.93 |  | a3 | -1.36 | 1.35 | -1.36
b4 | -0.09 | -1.18 | -0.64 |  | a4 | 2.03 | 1.76 | 0.13
b5 | -0.40 | 1.62 | 0.61 |  | a5 | -0.93 | 1.90 | -1.42
b6 | -0.47 | -2.87 | -1.67 |  | a6 | -1.19 | -0.28 | -0.45
b7 | -0.15 | 0.20 | 0.03 |  | a7 | 0.27 | -0.12 | 0.19
b8 | 0.12 | 0.09 | 0.10 |  | a8 | -1.49 | -0.57 | -0.46
b9 | 0.06 | 0.13 | 0.10 |  | a9 | -0.08 | 0.05 | -0.07
b10 | 0.27 | 0.33 | 0.30 |  | a10 | -0.15 | -0.03 | -0.06
b11 | 0.09 | 0.14 | 0.11 |  | a11 | -0.01 | 0.05 | -0.03
b12 | -0.02 | -0.01 | -0.01 |  | a12 | 0.07 | 0.03 | 0.02
b13 | 0.17 | 0.14 | 0.15 |  | a13 | -0.11 | 0.19 | -0.15
b14 | -0.95 | -0.97 | -0.96 |  | a14 | 0.15 | 0.02 | 0.07
b15 | 0.00 | 0.00 | 0.00 |  | a15 | 0.00 | 0.01 | -0.01
* reference frame of A04
Harmonics are given in units at the reference radius of 50 mm

## Slide 10: Field quality at nominal
Lines show the expected range at ±(3+1) σ

## Slide 11: Field quality at injection current (960 A)
Current 960 A (precycle at 16230 A and reset at 200 A) |  |  |  |  |  |  |  | 
 | A04 | A03 |  |  |  | A04 | A03 | 
b3 | -1.08 | 1.32 |  |  | a3 | -0.18 | -2.38 | 
b4 | -0.58 | -1.08 |  |  | a4 | -1.85 | 0.26 | 
b5 | 1.11 | 1.26 |  |  | a5 | -0.07 | 1.16 | 
b6 | -1.00 | -3.38 |  |  | a6 | -1.17 | -0.22 | 
b7 | 0.19 | -0.22 |  |  | a7 | 0.24 | -0.52 | 
b8 | 0.20 | 0.14 |  |  | a8 | -1.54 | -0.99 | 
b9 | -0.41 | -0.12 |  |  | a9 | 0.18 | -0.51 | 
b10 | 2.60 | 2.90 |  |  | a10 | -0.22 | 0.01 | 
b11 | 0.00 | 0.09 |  |  | a11 | -0.01 | 0.04 | 
b12 | 0.00 | 0.00 |  |  | a12 | 0.06 | 0.03 | 
b13 | 0.15 | 0.08 |  |  | a13 | -0.07 | 0.08 | 
b14 | -0.56 | -0.57 |  |  | a14 | 0.05 | 0.01 | 
b15 | 0.01 | 0.00 |  |  | a15 | 0.00 | 0.02 | 
Harmonics are given in units at the reference radius of 50 mm

## Slide 12: TF vs current - ramp-up of machine cycle

## Slide 13: Harmonics vs current - ramp-up of machine cycle

## Slide 14: Harmonics vs current - ramp-up of machine cycle

## Slide 15: Decays after ramping – ramp up of stair-step cycle - integral

## Slide 16: Decays after ramping - individual segments on magnet A04
The decays look uniform along the magnet length.

## Slide 17: Quench antenna signals and voltage spikes
First ramp after second cooldown

## Slide 18: Conclusions
  Measurements of integral gradient and field quality performed during the first campaign at CERN
  Main results:
    Integral gradient and nodal distance in agreement with results from AUP
    Field quality within the expected range
    On plateaus after ramping, some harmonics show a decay up to 2/3 units. Similar effect already seen on MQXFB magnets.
    Quench antenna signals don’t show events like the voltage spike seen on other MQXFA magnets. The events at intermediate and high current look of mechanical nature.
  Alignment measurements were not completed during first campaign. They are ongoing now.