## Slide 1
Attempt to measure field change during MQXFA17 ‘spike’ event.
19Sep2024
J. DiMarco, P. Joshi, F. Kurian, V. Teotia

## Slide 2
The full-length ‘Theta Quench Antennas’ (TQA) were switched from DQBuck with amplifiers to UnBucked (UB) configuration in May2024
Thanks to the BNL technical team for their great support doing this work!
TQA are 12 radial, full-length PCB-like, antennas having dipole and quadrupole bucking
These were re-configured to be ‘UnBucked’ in order to have sensitivity to dipole and quadrupole field changes

## Slide 3
The UB TQA were tested with short (90mm), dipole magnet cycled with 60Hz current.
The field from the magnet was measured to be ~ 8mT, magnetic length ~ 0.23m  TQA sees ~1.85 mTm integrated field change.

## Slide 4
Scope signal on TQA is ~15mV P-P for the AC 1.85 mTm field change of the small dipole. measured integrated voltage signal from scope is ~4e-5 Vs.
To compare with what is expected from TQA, the UB quench antenna has sensitivity of about 0.015 m^2, so 1.85 mTm integrated field change should give about 3e-5 Vs signal on the quench antenna.
 Agrees roughly with expected

## Slide 6
MQXFA17 first trip/quench had a ‘current spike’ event.
Voltage
The body quench antennas, which buck dipole and quad have high sensitivity and clearly saw the event.

## Slide 7
Voltage signals on TQA during spike event

## Slide 8
Some small flux changes probably are being seen here
TQA flux signals do not seem to see anything at the level above 1e-6 Vs
These larger changes (still <1e-6) do not involve multiple angles, so are probably not from ‘main field’ changes

## Slide 9
Quadrupole field component determined from TQA flux changes vs time. (Note that the integrated main quadrupole field is ~ 25T at the TQA radius)

## Slide 10
Perhaps in terms of integral field, flux changes are cancelling axially (?)

## Slide 11
Summary/Conclusion:
The azimuthal full-length quench antennas (Theta Quench Antennas (TQA)) were reconfigured to be sensitive to dipole and quadrupole fields.
The reconfigured TQA were tested with a small AC magnetic field to assess their sensitivity – they could clearly see integrated field changes on the order of 1.85e-3 Tm, resulting in TQA signals of about 5e-5 Vs.
TQA signals during quench show < 1e-6 Vs flux change during the quench with a ‘spike event’, and so no large change in field seemed to accompany this event.
The integrated quadrupole field changes during the spike event were measured to be < 2.5e-4 T, compared to the main field integral of ~25 T, which would be a negligible perturbation to the field.
Additional slides

## Slide 13
Flux ‘shift’ in channels over about 30ms) that later see ‘spike’ (as if the later spike is part of a redistribution back to previous level )
Flux transition (over ~ 3ms) that slowly ‘recovers’