## Slide 1
IXPE Gas Pixel Detector test and characterization with the X-ray Calibration Facility
Stefano Tugliani
Simone Maldera, Andrea Frassà, Raffaella Bonino,
Nicolò Cibrario, Luca Latronico, Alessio Gorgi

## Slide 2
2
IXPE GPD test and characterization with the XCF
15/05/2025
IXPE: Imaging X-ray Polarimetry Explorer
Stefano Tugliani
The state-of-art of X-ray polarimetry: Imaging X-ray Polarimetry Explorer, IXPE, a space mission (NASA and ASI collaboration), launched on 9 December 2021 (equatorial orbit at 600 km) that measures the linear polarization of different astrophysical sources in the energy range 2-8 keV
X-ray Astrophysics: study of the high energy Universe.
Electromagnetic Radiation
Characteristics of celestial sources
Information about the geometry both of the emitting matter and of magnetic and gravitational fields.
spatial, spectral, timing and polarization properties
Sensitivity to polarization (Minimum Detectable Polarization) | < 5.5%, 0.5 mCrab, 10 days
Angular resolution | < 30 arcsec
Field of view (FOV) | 10 arcmin
Energy range and resolution | 2 – 8 keV, 20% @ 5.9 keV
Timing accuracy | 20 s
Sensitivity to polarization (Minimum Detectable Polarization) | 
Angular resolution | < 30 arcsec
Field of view (FOV) | 
Energy range and resolution | 
Timing accuracy | 
Orbiting telescopes because of atmospheric opacity to X-rays.
Weisskopf et al. (2021)

## Slide 3
3
IXPE GPD test and characterization with the XCF
IXPE: Imaging X-ray Polarimetry Explorer
Stefano Tugliani
Solar array
Boom w/ thermal Sock deployed
Mirror Module Assembly (x3) (Wolter I)
X-ray Shields (x3) deployed
Forward star tracker
Detector Unit (x3)
15/05/2025

## Slide 4
4
IXPE GPD test and characterization with the XCF
IXPE: Imaging X-ray Polarimetry Explorer
Stefano Tugliani
Detector
Unit
15/05/2025

## Slide 5
5
IXPE GPD test and characterization with the XCF
The Gas Pixel Detector
Stefano Tugliani
V drift = -2800 V
V top = -840 V
V bottom = -400 V
V ASIC = 0 V
ADC
Photoelectron
emission direction
: angle of emission of the photoelectron wrt the polarization direction
10
mm
0.7
mm
ASIC
GEM
DRIFT PLANE
Dymethyl-ether
DME
Interaction point
Interaction point
Interaction point
Emission direction
Detector
Unit
Baldini et al. (2021)
15/05/2025

## Slide 6
6
IXPE GPD test and characterization with the XCF
IXPE polarization measurement
Stefano Tugliani
Event in Detector Coordinates
Event in Celestial Coordinates
Polarization information: recovered on a statistical basis from the azimuthal distribution of the photoelectron directions of emission ().
Photoelectric differential cross section:
The total counts are modulated as:
Muleri & Fabiani (2014), Muleri (2014), Kislat et al. (2015)
A
B
Modulation factor (): response to 100% linearly polarized radiation.
Polarization angle
15/05/2025

## Slide 7
7
IXPE GPD test and characterization with the XCF
GPD polarization measurement
Stefano Tugliani
For each event k:
The azimuthal emission of the photoelectron is modulated as:
For N selected events:
Polarization Degree
Polarization Angle
Ejected
photoelectron
Incoming photon
Polarization direction
x
y
z
Plane // to detector
Kislat et al. (2015), Rankin et al. (2022)
Photoelectron emission direction projection
15/05/2025

## Slide 8
Towards the X-ray Calibration Facility
IXPE
Physics Department Turin (ITA)
4 spare GPDs
3 GPDs (DUs)
8
IXPE GPD test and characterization with the XCF
Stefano Tugliani
GPDs
Only space-qualified detectors currently used for X-ray polarization
After launch:
systematic testing and monitoring on spare GPDs
Space
Ground
Essential in view of future X-ray polarimeters
We need a facility that provides photon beams with various spatial and energy configurations and offering both unpolarized and linearly polarized beams
XCF
Before launch: calibration and qualification of the detectors.
15/05/2025

## Slide 9
9
IXPE GPD test and characterization with the XCF
The X-ray Calibration Facility
Stefano Tugliani
X-ray tube: radiation source
The X-ray Calibration Facility XCF is an irradiation setup at the University of Turin to test and characterize:
IXPE Gas Pixel Detectors, monitoring their performance.
Future position- energy- and polarization-sensitive X-ray detectors.
Radiation source:
Multi-anode X-ray tube (Mc Pherson Mod. 642).
Sealed Micro X-ray tube equipped with a Molybdenum anode with emission at 2.293 keV.
source.
Handling system
Vacuum system:
Pfeiffer scroll and turbo pumps
Handling system:
  PiMicro to move the detectors and the polarizing crystal inside the cylinder
Tugliani et al. (2024a), (2024b)
15/05/2025

## Slide 10
10
IXPE GPD test and characterization with the XCF
The X-ray Calibration Facility
Stefano Tugliani
X-ray tube: radiation source
GPD
CMOS camera
SDD
Handling system
Unpolarized beam line
X-RAY BEAMS MONITORS
Silicon Drift Detector:
Resolution: 2% FWHM @ 5.9 keV Spectrum
CMOS ASI ZWO camera:
2822x4144 4m pixel array
Spectrum and charge map
Mo L
Mo L
Mo escape peak
Tugliani et al. (2024a), (2024b)
15/05/2025

## Slide 11
11
IXPE GPD test and characterization with the XCF
GPD equalization map with XCF
Stefano Tugliani
GPD
Study of the uniformity of the GPD response
GPD 35
Flat field using the unpolarized beam
For each bin [i][j] of the map:
Gain map
Baldini et al. (2021)
15/05/2025

## Slide 12
12
IXPE GPD test and characterization with the XCF
Spurious modulation
Stefano Tugliani
The azimuthal emission of the photoelectron is modulated as:
For N selected events:
[rad]
Entries/bin
[rad]
[rad]
Entries/bin
Entries/bin
Simulated unpolarized beam
Real unpolarized beam
Polarized beam
Rankin et al. (2022)
15/05/2025

## Slide 13
13
IXPE GPD test and characterization with the XCF
Spurious modulation studies with XCF
Stefano Tugliani
90° rotation
Spurious modulation:
Source modulation:
How to measure the spurious modulation?
SM
source
SM
source
Rankin et al. (2022)
15/05/2025

## Slide 14
14
IXPE GPD test and characterization with the XCF
Stefano Tugliani
GPD 30
GPD 35
GPD 30 and GPD 35 have different GEMS
GPD 30 has a higher spurious modulation and more features in the Stokes maps, compared to GPD 35
Additional hint that spurious modulation is correlated to the GEM
Spurious modulation: XCF first results
MEAN SPURIOUS MODULATION:
0.0022 ± 0.0002
MEAN INTRINSEC MODULATION:
0.0531 ± 0.0002
MEAN SPURIOUS MODULATION:
0.0144 ± 0.0002
MEAN INTRINSEC MODULATION:
0.0517 ± 0.0002
15/05/2025

## Slide 15
15
IXPE GPD test and characterization with the XCF
The X-ray Calibration Facility
Stefano Tugliani
Polarized beam line
GPD
CMOS camera
SDD
X-ray tube: radiation source
Plane of incidence
Diffracted polarized radiation
Unpolarized incident radiation
Polarizing crystal
Handling system
Bragg diffraction
Polarizing the beam:
Polarization:
For a totally polarized radiation:
k=0 e ==45°.
Henke et al. (1993), Muleri & Fabiani (2014), Ferrazzoli et al. (2020)
15/05/2025

## Slide 16
16
IXPE GPD test and characterization with the XCF
The X-ray Calibration Facility
Stefano Tugliani
Polarized beam line
GPD
CMOS camera
SDD
X-ray tube: radiation source
Polarizing crystal
Handling system
2xMo L
3xMo L
4xMo L
In L
Sb L
Mo L
Mo L
Mo escape peak
Polarized spectrum with SDD
peaks at integer multiples of E
fluorescence lines of the crystal
Tugliani et al. (2024b)
15/05/2025

## Slide 17
17
IXPE GPD test and characterization with the XCF
Polarizing the beams
Stefano Tugliani
XCF is equipped with a set of 6 crystals which are matched with specific energies so that the Bragg condition is satisfied.
Tugliani et al. (2024b)
InSb 111
2.28 keV
The Bragg arc
Crystal | Energy [keV] | Bragg | P [%]
InSb 111 | 2.28 | 46.6 ° | 99
Ge 111 | 2.62 | 46.5 ° | 99
Si 111 | 2.82 | 44.4 ° | 99
Si 220 | 4.45 | 46.6 ° | 98
Si 400 | 6.32 | 46.3 ° | 99
Ge 422 | 7.43 | 46.2 ° | 99
Crystal | Energy [keV] |  | P [%]
InSb 111 | 2.28 | 46.6 ° | 
Ge 111 | 2.62 | 46.5 ° | 
Si 111 | 2.82 | 44.4 ° | 
Si 220 | 4.45 | 46.6 ° | 
Si 400 | 6.32 | 46.3 ° | 
Ge 422 | 7.43 | 46.2 ° | 
Modulation
15/05/2025

## Slide 18
18
IXPE GPD test and characterization with the XCF
XCF polarized beam studies
Stefano Tugliani
B
A
PRELIMINARY RESULTS
The track length increases with increasing energy and the photoelectron emission direction is better reconstructed
The modulation factor increases with energy
Tugliani et al. (2024a), (2024b)
Low energies
Compact tracks
High energies
Elongated tracks
15/05/2025

## Slide 19
GPD long term studies: secular pressure decrease
Secular decrease of the pressure in the sealed gas cell, due to internal adsorption of DME with a time scale of months.
The pressure decrease can be retrieved analyzing:
Increase of track
length
Increase of gain
Decrease of rate
Tomaiuolo et al. (2024)
Spurious modulation: XCF first results
19
IXPE GPD test and characterization with the XCF
Stefano Tugliani
Source:
5.9 keV
15/05/2025

## Slide 20
20
IXPE GPD test and characterization with the XCF
Stefano Tugliani
Initially conceived as a calibration source to qualify GPDs, the XCF can satisfy evolving requirements to support R&D programs of innovative position, energy and polarization-sensitive X-ray detectors.
Conclusions
Unpolarized beam:
Gain and equalization maps;
Systematics in the detector.
Polarized beam:
Response to polarized radiation;
Calibration.
Tugliani et al. (2024a), (2024b)
15/05/2025

## Slide 21
Thank you for your attention

## Slide 22
22
IXPE GPD test and characterization with the XCF
Stefano Tugliani
References
Weisskopf et al. The Imaging X-Ray Polarimetry Explorer (IXPE): Pre-Launch. 2021. doi: 10.48550/ARXIV.2112.01269.
Baldini et al. “Design, construction, and test of the Gas Pixel Detectors for the IXPE mission”. In: Astroparticle Physics 133 (2021), p. 102628. issn: 0927-6505. doi: https:// doi.org/10.1016/j.astropartphys.2021.102628.
Fabio Muleri Sergio Fabiani. Astronomical X-ray Polarimetry. Aracne, 2014.
Fabio Muleri. “ON THE OPERATION OF X-RAY POLARIMETERS WITH A LARGE FIELD OF VIEW”. In: The Astrophysical Journal 782.1 (Jan. 2014), p. 28. doi: 10.1088/ 0004-637x/782/1/28.
Tugliani et al. “IXPE Gas Pixel Detectors characterization using the X-ray Calibration Facility”. In: Nuovo Cim. C 47.3 (2024), p. 141. doi: 10.1393/ncc/i2024-24141-9.
Tugliani et al. “IXPE Gas Pixel Detector characterization with the X-ray Calibration Facility”. In: Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spec- trometers, Detectors and Associated Equipment 1069 (2024), p. 169800. issn: 0168-9002. doi: https://doi.org/10.1016/j.nima.2024.169800.
Kislat et al. “Analyzing the data from X-ray polarimeters with Stokes parameters”. In: Astroparticle Physics 68 (Aug. 2015), pp. 45–51. doi: 10.1016/j.astropartphys.2015. 02.007.
Ferrazzoli et al. “In-flight calibration system of imaging x-ray polarimetry explorer”. In: Journal of Astronomical Telescopes, Instruments, and Systems 6.04 (Oct. 2020). doi:10.1117/1.jatis.6.4.048002.
B.L. Henke, E.M. Gullikson, and J.C. Davis. “X-Ray Interactions: Photoabsorption, Scattering, Transmission, and Reflection at E = 50-30,000 eV, Z = 1-92”. In: Atomic Data and
Nuclear Data Tables 54.2 (1993), pp. 181–342. issn: 0092-640X. doi: https://doi.org/10.1006/adnd.1993.1013.
John Rankin et al. “An Algorithm to Calibrate and Correct the Response to Unpolarized Radiation of the X-Ray Polarimeter Onboard IXPE”. In: The Astronomical Journal 163.2
(Jan. 2022), p. 39. doi: 10.3847/1538-3881/ac397f.
C. Tomaiuolo et al. " Time-dependent instrumental effects in IXPE: Pressure variation and GEM charging inside GPDs ". In: Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment, 1069 (2024), p. 169881. issn: 0168-9002. doi: https://doi.org/10.1016/j.nima.2024.169881.
15/05/2025

## Slide 23
23
IXPE GPD test and characterization with the XCF
IXPE data processing
Stefano Tugliani
Level 1 data:
Event reconstruction
in detector coordinates
Level 2 data:
Data in sky coordinates
Celestial Coordinates
3 Detector
Units
Star Tracker:
Celestial
Coordinates system
DU1
DU2
DU3
Detector Coordinates
15/05/2025

## Slide 24
24
IXPE GPD test and characterization with the XCF
Polarizing the beams
Stefano Tugliani
2.293 keV (Mo L)
Choose an energy
Find the polarizing crystal
emission line
Rotate the crystal at the Bragg angle
The capillary plate selects
CMOS camera
There is no peak in the intensity
Preferred energy
Preferred angle
Bragg arc
Henke et al. (1993), Muleri & Fabiani (2014), Tugliani et al. (2024b)
15/05/2025

## Slide 25
25
IXPE GPD test and characterization with the XCF
Polarizing the beams
Stefano Tugliani
4.5 keV
CMOS camera
Choose an energy
Find the polarizing crystal
e.g. 4.5 keV
Rotate the crystal at 45° to maximize the polarization
The capillary plate selects
The intensity peaks at the emission line
No preferred energy
No preferred angle
No Bragg arc
In the Bremsstrahlung continuum
Henke et al. (1993), Muleri & Fabiani (2014), Tugliani et al. (2024b)
15/05/2025

## Slide 26
26
IXPE GPD test and characterization with the XCF
XCF polarized beam studies
Stefano Tugliani
InSb 111
Si 220
Si 400
In addiction to the diffracted polarized X-ray, the spectra show features at the fluorescence energy of the crystal: these photons are not polarized.
It is possible to observe also the energy correspondent to the second diffraction order: the obtained is coherent with the one expected at that energy:
2.28 keV
6.32 keV
1.74 keV
Tugliani et al. (2024b)
15/05/2025