## Slide 1
Diagnostics in High Intensity Beams, Part 2
CAS on Intensity Limits in Hadron Beams
Borovets, 19th of June 2025
Peter Forck
Gesellschaft für Schwerionenforschnung (GSI)
p.forck@gsi.de

## Slide 2
Outline for the Lecture Beam Diagnostics
Outline
General remarks to beam instrumentation and diagnostics
Overview for beam current measurement
Various method for transverse profile determination
High-current example 1 and 2
Emittance determination
High-current example 3
Usage of Beam Position Monitors
High-current example 4
Longitudinal Measurements
Beam Loss Monitors

## Slide 3
Outline:
Signal generation  transfer impedance
Capacitive button and linear-cut BPM for high frequencies
Electronics for position evaluation
BPMs for measurement such as closed-orbit, tune and chromaticity
High current example 4 (tune measurement for high currents)
Summary
Pick-Ups for bunched Beams
A Beam Position Monitor is an non-destructive device for bunched beams
It delivers information about the transverse center of the beam:
Trajectory: Position of an individual bunch within a transfer line or synchrotron
Closed orbit: Central orbit averaged over a period much longer than a betatron oscillation
Single bunch position: Determination of parameters like tune, chromaticity, -function
Remarks: - BPMs have a low cut-off frequency  dc-beam can’t be monitored
- The abbreviation BPM and pick-up PU are synonyms

## Slide 4
Pick-Ups for bunched Beams
The image current at the beam pipe is monitored on a high frequency basis
i.e. the ac-part given by the bunched beam.
Beam Position Monitor BPM is the most frequently used instrument!
For relativistic velocities,
the electric field is transversal:

## Slide 5
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
V
-
-
-
-
-
-
-
Principle of Signal Generation of a BPMs, centered Beam
Animation by Rhodri Jones (CERN)
The image current at the wall is
monitored on a high frequency basis
i.e. ac-part given by the bunched beam.

## Slide 6
Model for Signal Treatment of capacitive BPMs
Equivalent circuit
The wall current is monitored by a plate or ring inserted in the beam pipe:
Ubeam
ground
CC
C
R
Ibeam
capacitive coupling
 image current
electronics elements
Uim
At a resistor R the voltage Uim from the image current is measured.
Goal: Connection from beam current to signal strength by transfer impedance Zt ( )
in frequency domain: Uim(ω) = R · Iim(ω) = Zt(ω) · Ibeam(ω)
Result: i.e. complex function
geometry
stray capacitance
frequency response

## Slide 7
Example of Transfer Impedance for Proton LINAC
The high-pass characteristic for typical proton LINAC BPM:
Large signal strength for long bunches  high impedance
Smooth signal transmission important for short bunches  50 
Parameter button BPM at proton LINAC:
pF, cm,  = 90%
for   MHz
for M  kHz
Remark: For   0 it is Zt  0 i.e. no signal is transferred from dc-beams e.g.
 de-bunched beam inside a synchrotron
 for slow extraction through a transfer line
fcut
decrease by
factor 1/ 2 = 3 dB

## Slide 8
Calculation of Signal Shape (here single Bunch)
The transfer impedance is used in frequency domain! The following is performed:
1. Start: Time domain Gaussian function Ibeam(t) having a width of σt
2. FFT of Ibeam(t) leads to the frequency domain Gaussian Ibeam(f) with
3. Multiplication with Zt(f) with fcut = 32 MHz leads to Uim(f) = Zt(f) ∙ Ibeam(f)
4. Inverse FFT leads to Uim(t)
Fourier trans.
inverse Fourier trans.
Remark: Time domain processing via convolution or filters (FIR and IIR) are possible

## Slide 9
Calculation of Signal Shape: Repetitive Bunches
Fourier spectrum is concentrated at acceleration harmonics
with single bunch spectrum as an envelope.
Bandwidth up to typically 10*facc
LINAC beam or synchrotron filled with 8 bunches accelerated with facc= 100 MHz
BPM terminated with R=50   facc<< fcut :
Parameter: R=50   fcut=318 MHz, all buckets filled
C=10 pF, l=1 cm,  =90 %, σt=0.66 ns  σl=18 cm

## Slide 10
Calculation of Signal Shape: Bunch Train with empty Buckets
Fourier spectrum is more complex, harmonics are broader due to sidebands
LINAC or synchrotron filled with empty buckets accelerated with facc= 100 MHz
BPM terminated with R=50   facc<< fcut :
Parameter: R=50   fcut=318 MHz, 2 empty buckets
C=10 pF, l=1 cm,  =90 %, σt=0.66 ns  σl=18 cm

## Slide 11
Cable damping given by:
fe: amplitude damping by 10 dB
Bunch signal is damped; 10-fold lower amplitude, higher frequencies are damped stronger
Bunch signal gets asymmetric; baseline modification i.e. long time reach Uim = 0
 ‘Good cables’ are a precaution for broadband signal transmission
Parameter:
Cable damping, fe= 300 MHz (e.g. 90 m RG 214)
C=10 pF, l=1 cm,  =90 %, σt=0.66 ns  σl=18 cm
Calculation of Signal Shape: Bunch Train with Cable Damping
Effect of cables or other band-limiting electronics

## Slide 12
Calculation of Signal Shape: Filtering of Harmonics
Remark: For numerical calculations, time domain filters (FIR and IIR) are more appropriate
Generally: Ztot(ω) = Hcable(ω) ∙ Hfilter(ω) ∙ Hamp(ω) ∙ … ∙Zt(ω)
Effect of filters, here 4th order Butterworth bandpass filter:
and
Only few frequency components leading to ‘ringing’ due to sharp cutoff
Other filter types more appropriate
Parameter: R=50 , 4th order Butterworth filter at flow=180 MHz & fhigh=220 MHz
C=10 pF, l=1 cm,  =90 %, σt=0.66 ns  σl=18 cm

## Slide 13
-
-
-
-
-
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
V
-
-
-
-
-
-
-
-
+
+
+
+
+
-
-
+
-
+
-
+
-
+
-
-
-
-
+
+
+
-
+
+
-
+
-
-
+
-
+
-
-
-
-
-
+
+
+
-
+
+
-
-
+
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
Animation by Rhodri Jones (CERN)
Principle of Signal Generation of a BPMs: off-center Beam
The image current at the wall is
monitored on a high frequency basis
i.e. ac-part given by the bunched beam.

## Slide 14
Principle of Position Determination by a BPM
The difference voltage between plates gives the beam’s center-of-mass
most frequent application
S(ω,x) is called position sensitivity, sometimes the inverse is used k(ω,x)=1/S(ω,x)
S is a geometry dependent, non-linear function,
Units: S=[%/mm], sometimes S=[dB/mm] or k=[mm].
It is at least:
y from ΔU = Uup - Udown
Typical desired position resolution:
x  0.1 ... 0.3  x of beam width
Example: One turn = 4 bunches @ 35 MeV/u
Uright
Ureft

## Slide 15
Outline:
Signal generation  transfer impedance
Capacitive button BPM for high frequencies
used at most proton LINACs
Capacitive linear-cut BPM for low frequencies
Electronics for position evaluation
BPMs for measurement of closed orbit, tune and further lattice functions
High-current example 4
Summary
Pick-Ups for bunched Beams

## Slide 16
2-dim Model for a Button BPM
‘Proximity effect’: larger signal for closer plate
Ideal 2-dim model: Cylindrical pipe  image current density
via ‘image charge method’ for ‘pencil’ beam:
Image current: Integration over finite button size:
,  = 0o


a
r
button
 =0

## Slide 17
Ideal 2-dim model: Non-linear behavior and hor-vert coupling:
Sensitivity S converts signal to position
with S = [%/mm] or [dB/mm]
i.e. S is the derivative of the curve , here Sx = Sx(x, y) i.e. non-linear.
For this example: central part S=7.4%/mm  k=1/S=14mm
button
 real position
• measured position
Position Map
Horizontal plane
2-dim Model for a Button BPM


a
r
button


## Slide 18
Button BPM Realization
LINACs, e--synchrotrons: 100 MHz < frf < 3 GHz  bunch length  BPM length
 50  signal path to prevent reflections
Example: LHC-type inside cryostat:
24 mm, half aperture a = 25 mm, C = 8 pF
 fcut= 400 MHz, Zt = 1.3  above fcut
Ø24 mm

5cm
N-Connector
Button electrode
Ceramic vacuum insulation
Courtesy C. Boccard (CERN)
 Electronics

## Slide 19
Outline:
Signal generation  transfer impedance
Capacitive button BPM for high frequencies
used at most proton LINACs
Capacitive linear-cut BPM for low frequencies
used at most synchrotrons
Electronics for position evaluation
BPMs for measurement of closed orbit, tune and further lattice functions
High-current example 4
Summary
Pick-Ups for bunched Beams

## Slide 20
Linear-cut BPM for Proton Synchrotrons
Frequency range: 1 MHz < frf < 100 MHz  bunch-length >> BPM length.
Signal is proportional to actual plate length:
In ideal case: linear reading
Linear-cut BPM:
Advantage: Linear, i.e. constant position sensitivity S
 no beam size dependence
Disadvantage: Large size, complex mechanics
high capacitance
Size: 200x70 mm2

## Slide 21
Technical Realization of a linear-cut BPM
Quadrupole
Linear Ampl.
Beam
Technical realization at HIT synchrotron of 46 m length for 7 MeV/u 440 MeV/u
BPM clearance: 180x70 mm2, standard beam pipe diameter: 200 mm.
70mm
180mm
Linear Ampl.
Horizontal
electrode
Right
channel

## Slide 22
70mm
180mm
Linear Ampl.
Horizontal
electrode
Right
channel
Technical Realization of a linear-cut BPM
Technical realization at HIT synchrotron of 46 m length for 7 MeV/u 440 MeV/u
BPM clearance: 180x70 mm2, standard beam pipe diameter: 200 mm.

## Slide 23
Comparison linear-cut and Button BPM
 | Linear-cut BPM | Button BPM
Precaution | Bunches longer than BPM | Bunch length comparable to BPM
BPM length (typical) | 10 to 20 cm length per plane | 1 to 5 cm per button
Shape | Rectangular or cut cylinder | Orthogonal or planar orientation
Bandwidth (typical) | 0.1 to 100 MHz | 100 MHz to 5 GHz
Coupling | 1 M or 1 k (transformer) | 50 
Cutoff frequency (typical) | 0.01… 10 MHz (C=30…100pF) | 0.3… 1 GHz (C=2…10pF)
Linearity | Very good, no x-y coupling | Non-linear, x-y coupling
Sensitivity | Good, care: plate cross talk | Good, care: signal matching
Usage | At proton synchrotrons, frf < 10 MHz | All electron acc., proton Linacs, frf > 100 MHz
Remark: Other types are also some time used: e.g. strip-line, wall current monitors,
inductive antenna, BPMs with external resonator, cavity BPM, slotted wave-guides etc.

## Slide 24
Outline:
Signal generation  transfer impedance
Capacitive button BPM for high frequencies
used at most proton LINACs
Capacitive linear-cut BPM for low frequencies
used at most synchrotrons
Electronics for position evaluation
analog & digital signal conditioning to achieve small signal processing
BPMs for measurement of closed orbit, tune and further lattice functions
High-current example 4
Summary
Pick-Ups for bunched Beams

## Slide 25
Broadband Signal Processing
Hybrid or transformer close to beam pipe for analog ΔU & ΣU generation or Uleft & Uright
Attenuator/amplifier
Filter to get the wanted harmonics and to suppress stray signals
ADC: digitalization  followed by calculation of of ΔU /ΣU
.
Advantage: Bunch-by-bunch observation possible, versatile post-processing possible
Disadvantage: Resolution down to  100 µm for large aperture, i.e. 0.1% of aperture,
resolution is worse than narrowband processing, see below
Challenge: Precise analog electronics with very low drift of amplification

## Slide 26
General: Noise Consideration
1. Signal voltage given by:
2. Thermal noise voltage given by:
Example: GSI-LINAC with frf = 36 MHz
Signal-to-noise ΔUim/Unoise is influenced by:
Input signal amplitude
Thermal noise from amplifiers etc.
Bandwidth Δf
 Restriction of frequency width
as the power is
concentrated at harmonics n frf

## Slide 27
Narrowband Processing for improved Signal-to-Noise
Narrowband processing equals heterodyne receiver (e.g. AM-radio or analog spectrum analyzer)
Attenuator/amplifier
Mixing with accelerating frequency fLO
 IF-output: signal with difference frequency fIF = fLO – fRF
Bandpass filter of the mixed signal (e.g at 10.7 MHz)
Rectifier: synchronous detector
ADC: digitalization → followed calculation of ΔU/ΣU
Advantage: Spatial resolution about 100 time better than broadband processing
Disadvantage: No turn-by-turn diagnosis, due to mixing = ’long averaging time’
Digital
correspondence:
I/Q demodulation

## Slide 28
Analog versus Digital Signal Processing
Digital receiver as modern successor of super heterodyne receiver
Basic functionality is preserved but implementation is very different
Digital transition just after the amplifier & filter or mixing unit
Signal conditioning (filter, decimation, averaging) on FPGA
Advantage of DSP: Versatile operation, flexible adoption without hardware modification
Disadvantage of DSP: non, good engineering skill requires for development, expensive
Analog frequency
translator
Analog demodulator
and filter
ADC
BPM analog
signal
digital
output
Analog frequency
translator
ADC
Digital
filter
Digital
Signal Proc.
digital
output
BPM analog
signal
Modern digital processing
Traditional analog processing
Analog
Digital
Modern instrumentation uses digital techniques with extended functionality.

## Slide 29
Comparison of BPM Readout Electronics (simplified)
Type | Usage | Precaution | Advantage | Disadvantage
Broadband | p-sychr. | Long bunches | Bunch structure signal Post-processing possible Required for transfer lines with few bunches | Resolution limited by noise
Narrowband | all synchr. | Stable beams >100 rf-periods | High resolution | No turn-by-turn Complex electronics
Digital Signal Processing | all | ADC sample typ. 250 MS/s | Very flexible & versatile High resolution Trendsetting technology for future demands | Basically non! Limited time resolution by ADC  under-sampling Man-power intensive
 Stripline BPM

## Slide 30
Outline:
Signal generation  transfer impedance
Capacitive button BPM for high frequencies
used at most proton LINACs
Capacitive linear-cut BPM for low frequencies
used at most synchrotrons
Electronics for position evaluation
analog & digital signal conditioning to achieve small signal processing
BPMs for measurement of closed orbit, tune and further lattice functions
frequent application of BPMs
High-current example 4
Summary
Pick-Ups for bunched Beams

## Slide 31
Trajectory:
The position delivered by an individual bunch within a transfer line or a synchrotron.
Main task: Control of matching (center and angle), first-turn diagnostics
Example: LHC injection 10/09/08 i.e. first day of operation !
Courtesy R. Jones (CERN)
-10
-10
horizontal
vertical
Monitor number (530 BPMs on 27 km)
10
-10
0
x [mm]
10
-10
0
y [mm]
Tune values at LHC: Qh = 64.3, Qv = 59.3
Trajectory Measurement with BPMs

## Slide 32
Close Orbit Measurement with BPMs
Single bunch position averaged over 1000 bunches  closed orbit with ms time steps.
It differs from ideal orbit by misalignments of the beam or components.
Example: GSI-synchrotron at two BPM locations, 1000 turn average during acceleration:
Closed orbit:
Beam position
averaged over many turns
(i.e. betatron oscillations).
The result is the basic tool
for alignment & stabilization
horizontal position at 2 BPMs
vertical position at 2 BPMs
time t = 0 ... 600 ms
0
5
-5
-10
-5
0
2
4
-2
time t = 0 ... 600 ms
0
6
-6
-12
position [mm]
position [mm]
position [mm]
along synchrotron
time t = 0 ... 600 ms
signal strength
horizontal
vertical

## Slide 33
Closed Orbit Feedback: Typical Noise Sources
Courtesy M. Böge, PSI, N. Hubert, Soleil
Beam movement:
Short term (min to 10 ms):
Traffic
Machine (crane) movements
Water & vacuum pumps
50 Hz main power net
Medium term (day to min):
Movement of chambers
due to heating by radiation
Day-night variation
Tide, moon cycle
Long term ( > days):
Ground settlement
Seasons, temperature variation
1
0.1
100
10
1000
Frequency [Hz]
Power spectral density [mm2/Hz]
0.01

## Slide 34
Close Orbit Feedback: BPMs and magnetic Corrector Hardware
Orbit feedback: Synchrotron light source  spatial stability of light beam
Example: SLS-Synchrotron at Villigen, Switzerland
Feedback loop:
1. Position from all BPMs
2. Calculation of corrector setting via
Orbit Response Matrix
3. Change of magnet setting
1.’ New positon measurement …..
 regulation time down to 10 ms
From M. Böge, PSI
Uncorrected orbit: typ.
Corrected orbit: typ. up to  100 Hz bandwidth!
Position from all BPMs
Calculation of corrector strength
Setting of correctors
feedback
Acc. optics
Rule of thumb: To determine the ‘sine-like’ betatron oscillation, 4 BPMs per period required
 4 BPMs per tune value (but detailed investigation required to determine the # of BPMs)

## Slide 35
Graphics by R. Singh, GSI
Tune Measurement: General Considerations
Beam particle’s in-coherent motion  center-of-mass stays constant
Coherent excitations are required for the detection by a BPM
Excitation of all particles by rf  coherent motion
 center-of-mass variation turn-by-turn i.e. center acts as one macro-particle
Notes: Beam with several particles with randomly distributed phases and amplitude will give only very small fluctuations due to finite number of particles (schottky fluctuations) proportional to sqrt(N)
Beam excitation aligns the particle phases, and allows to measure center-of-mass before the beam decohores due to finite spread in frequency. Thus regular excitation needed. Types of excitation were studied as part of this work, Kick excitation (a wide band, or single time impulse excitation), requires high power, and significantly dilutes the phase space.
Chirp excitation, slow method, difficult to use in day to day operations, BTF excitation, to study the amplitude and phase response of the beam.
A compromise of the both the above methods, band limited (or pseudo random noise) for regular excitation without significant blow-up, usable along the ramp.

## Slide 36
Tune Measurement: The Kick-Method in Time Domain
The beam is excited to
coherent betatron oscillation:
 Beam position measured
each revolution (’turn-by-turn’)
 Fourier Trans. gives non-integer tune q.
Short kick compared to revolution.
Here: N = 200 turn  q > 0.003
(tune spreads can be q  0.001!)
The de-coherence time limits the resolution:
N non-zero samples
 General limit of discrete FFT:
Decay is caused by de-phasing,
not by decreasing
single particle amplitude.

## Slide 37
Tune Measurement: Frequency Chirp Measurement
SIS-synchrotron: A wide scan with
both sidebands at h=25th-harmonics:
From the position of the sidebands q = 0.306 is determined.
Advantage: High resolution for tune and tune spread (also for de-bunched beams)
Disadvantage: Long sweep time (up to several seconds).
A detailed scan for the lower sideband
 beam acts like a driven oscillator:
Principle: Slowly scan of the excitation frequency  beam acts as driven oscillator!
(sometimes refer to as Beam Transfer Function BTF measurement)


## Slide 38
Tune Measurement: Gentle Excitation with Wideband Noise
Instead of a sine wave, noise with adequate bandwidth can be applied
 beam picks out its resonance frequency:
Broadband excitation with white noise
of  10 kHz bandwidth
Turn-by-turn position measurement
Fourier transformation of the recorded data
 Continues monitoring with low disturbance
Example: Vertical tune within 4096 turn
duration ≃ 15 ms
at GSI synchrotron 11 → 300 MeV/u in 0.7 s
Advantage:
Fast scan with good time resolution
Disadvantage: Lower tune resolution
vertical tune versus time
Qy  0.02
time [ms]
vert. fractional tune qy
750
600
400
0
200
0.2
0.25
0.3
excitation
noise band
Q  0.05
vertical tune at fixed time  15ms
vertical fractional tune qy
0.2
0.25
0.35
0.3
amplitude [a.u.]
U. Rauch et al., DIPAC 2009

## Slide 39
Chromaticity ξ: Change of tune for off-momentum particle
Two step measurement procedure:
1. Change of momentum p by detuned rf-frequency
2. Excitation of coherent betatron oscillations
and tune measurement
(kick-method, BTF, noise excitation):
Chromaticity Measurement from Closed Orbit Data
Plot of ΔQ/Q as a function of Δp/p
 slope is dispersion ξ.
Example: Measurement at LEP:
From M Minty, F. Zimmermann,
Measurement and Control of charged Particle Beam,
Springer Verlag 2003
 Example 4

## Slide 40
Example for high-intensity effects
Tune spectrum measurement:
High intensity leads to non-linear oscillations
Tune spectra are modified
Coherent tune shift and incoherent tune spread can be determined
Beam parameters are significantly modified
Measurements must be interpreted in an appropriate manner
Tune Spectrum Modification for intense Beams

## Slide 41
Parameter | Typical Value
Circumference | m
Beam current | charges
Injection energy | MeV/u (
Betatron tunes | 4.31 & 3.28
Synchrotron tune | ( kHz)
Trans. size x & y | 6...10 & 3...5 mm
Parameter | Typical Value
Circumference | 
Beam current | 
Injection energy | 
Betatron tunes | 
Trans. size x & y | 6...10 & 3...5 mm
Low current: Tune spectra has symmetric sidebands due to synchrotron motion.
Relevant GSI SIS18 parameters:
Higher current: The „global“ peak moves to the left. The symmetry of the spectrum is broken.
Tune Spectra: Beam Intensity Modification
General questions:
Why is the tune spectrum modified
at increasing beam current ?
Very practical:
Which peak is the coherently shifted tune ?
rf cavity 0.8 < f < 5 MHz
dipole
quadrupole

## Slide 42
Space charge parameter
qsc includes effect of longitudinal oscillations
Incoh. tune spread
caused by repulsion due to electric field and attraction due to magnetic field of moving ions
Coherent tune shift
caused by image current flowing at beam pipe
Transvers impedance
Formulas for coherent Tune Shift and incoh. Tune Spread
Model assumptions: round constant beam & pipe radius, constant or KV distribution, constant synchrotron tune ...
 relation
beam radius
pipe radius
b pipe radius, R synchrotron radius,
A ion mass, q ion charge
N number of ions, Bf bunching factor, Q0 bare tune
W0 kinetic energy per nucleon, 0 velocity, 0 Lorentz factor
a beam radius,  = a2 /<> transverse emittance
here: ax  2 ay
See lecture on ‘Space Charge in circular machine’ by F. Asvesta
Notes: So, how do the space charge effects, modify the tune spectra?
First the more challenging and interesting effect, which was the prime motivation of all these detailed examples.
The distance between the head-tail modes change in a characteristic way, explained by certain theoretical and numerical studies, and our measurements match with the prediction quite well, and rather beautiful interpretation of tune spectra is obtained.
These are the examples of the theoretical curves, and the spectrum on the side for the given beam conditions exactly show, how the characteristic modification occurs.

## Slide 43
M. Blaskiewisz: Fast head-tail instabilities with space charge, Phys. Rev. Acc.. Beams 1, 044201, (1998)
A. Burov, Head-tail modes for strong space charge, Phys. Rev . Acc. Beams 12, 044202, (2009)
O. Boine-Frankenheim et .al., Transverse Schottky noise spectrum for bunches with space charge, Phys. Rev .Acc.. Beams 12, 114201, (2009)
Low intensity: spacing of sidebands:
rough scaling
Findings:
: it is i.e. value of shifted bare tune is of importance for operating
k > 0: head-tail modes come closer
k < 0: almost constant slope (but larger Landau damping  broader and lower spectral lines)
Assumption of analytical description of head-tail modes by M. Blaskiewicz (1998):
Transverse phase space: KV-distribution (i.e. idealized, suited for space charged calculations)
Longitudinal ‘airbag’ phase space: phase  constant momentum only two velocities vs =  vs0
 synchrotron tune: Qs =  vs0 /(R 0 )
Analytical Model: Space Charge Modification of Tune Spectra
Qc = 0.06 Qsc
High intensity: spacing of sidebands:
Qsc incoherent tune spread
Qc coherent tune shift
Example for SIS18 parameter
Notes: So, how do the space charge effects, modify the tune spectra?
First the more challenging and interesting effect, which was the prime motivation of all these detailed examples.
The distance between the head-tail modes change in a characteristic way, explained by certain theoretical and numerical studies, and our measurements match with the prediction quite well, and rather beautiful interpretation of tune spectra is obtained.
These are the examples of the theoretical curves, and the spectrum on the side for the given beam conditions exactly show, how the characteristic modification occurs.

## Slide 44
Red lines: prediction using measured beam parameters:
Measured spectra U73+ beam
Measurements: Moderate Space Charge Parameter
Beam: U73 +, 11.4 MeV/u, (1 ... 12)  108 ions,
Fractional horizontal tune Qfx
Qc = 0.06 Qsc
R. Singh et al., Interpretation of transverse tune spectrum in a heavy-ion synchrotron, Phys. Rev. Acc. Beams 13, 034201 (2013)

## Slide 45
Predicted spectra : red lines
Measurements: Strong Space Charge Parameter
Beam: N7+, 11.56 MeV/u, (3 ... 15)  109 ion,
(smaller transverse size than U73+ beam)
Qc = 0.14 Qsc
Measured spectra N7+ beam
R. Singh et al., Interpretation of transverse tune spectrum in a heavy-ion synchrotron, Phys. Rev. Acc. Beams 13, 034201 (2013)

## Slide 46
Measurements: Determination of coherent Tune Shift Qc
The coherent tune shift Qc from the depicted N7+ spectra and further spectra with Ar18+
For k = 0 : = i.e. value of coherently shifted bare tune as of importance for operating
Results:
peak is not the highest one
Linear scaling of coherent tune shifts versus peak beam current , different from set value
For value of slope the resistive, effective impedance Z can be determined
 estimation of effective beam pipe radius mm & mm
R. Singh et al., Interpretation of transverse tune spectrum in a heavy-ion synchrotron, Phys. Rev. Acc. Beams 13, 034201 (2013)
Measured spectra N7+ beam

## Slide 47
Slope
Tune Spread from Comparison of Space Charge Parameter
General trend well reproduced i.e. reason for modified tune spectra was found
Significant deviation for higher values of qsc (as expected from analytic model)
 More realistic model with higher prediction power should be used
The values qsc from the depicted measurements with U73+ and N7+
Measurement of all parameters:
Predicted
Measured based on the distance between lines
of modes and :  qk 0 1
Findings:
R. Singh et al., Interpretation of transverse tune spectrum in a heavy-ion synchrotron, Phys. Rev. Acc. Beams 13, 034201 (2013)

## Slide 48
11 consecutive turn-by-turn center-of-mass recordings:
Measurement: Time Domain Identification of Head-tail Modes
Frequency sweep allows the observation of bunch center oscillations i.e. head-tail modes
Beam parameter: 15  109 N7+ at 11.5 MeV/u, Qy = 3.75,  = -2.1
Tune spectrogram during sweep:
Results:
Sweep excitation allows excitation of individual head-tail modes
The mode-structure verifies the spectra interpretation
Can be used as a precise determination of chromaticity
Notes: The head tail modes were further identified in time domain, to indeed confirm that the negative modes could not be excited, and the left most mode is actually k=0 mode. On top of that, the mode structure could be used to calculate other beam parameters such as chromaticity.
These are some of the first clean measurements of excited head-tail modes. Normally they occur, when some of them cross their instability thresholds.

## Slide 49
End of example for high-intensity effects
Tune spectrum measurement:
High intensity leads to non-linear oscillations
Tune spectra are modified
Coherent tune shift and incoherent tune spread can be determined
Beam parameters are significantly modified
Measurements must be interpreted in an appropriate manner
Tune Spectrum Modification for intense Beams

## Slide 50
Summary Pick-Ups for bunched Beams
The electric field is monitored for bunched beams using rf-technologies
(’frequency domain’). Beside transformers they are the most often used instruments!
Differentiated or proportional signal: rf-bandwidth  beam parameters
Proton synchrotron: 1 to 100 MHz, mostly 1 M  proportional shape
LINAC: 0.1 to 3 GHz, 50   differentiated shape
Important quantity: Transfer impedance .
Types of capacitive pick-ups:
Linear-cut (p-synch.), button (p-LINAC, e−-LINAC and synch.)
Position reading: Difference signal of two or four pick-up electrodes (BPM):
Non-intercepting reading of center-of-mass → online measurement and control
Synchrotron: Fast reading, ‘bunch-by-bunch’ trajectory, slow reading  closed orbit
 Synchrotron: Excitation of coherent betatron oscillations  tune q,  , (s), D(s)...
Remark: BPMs have high pass characteristic  no signal for dc-beams

## Slide 51
Measurement of longitudinal parameter:
Bunch length measurement at
LINACs with beam of non-relativistic velocities
Bunch length at synchrotrons and tomography
Summary
Longitudinal  transverse correspondences:
position relative to rf  transverse center-of-mass
bunch structure in time  transverse profile
momentum or energy spread  transverse divergence
longitudinal emittance  transverse emittance.
Measurement of longitudinal Parameters

## Slide 52
The Bunch Position measured by a Pick-Up
The bunch position is given relative to the accelerating rf.
e.g. φref=-30o inside a rf cavity
must be well aligned for optimal acceleration
Transverse correspondence: Beam position
Example: Pick-up signal for frf =36 MHz rf at GSI-LINAC:
LINAC
RFQ
source
BPM
pick-up
ADC
acc. frf

## Slide 53
Determination of non-relativistic mean Energy using Pick-Ups
The energy delivered by a LINAC is sensitive to the mechanics, rf-phase and amplitude.
For non-relativistic energies
at proton LINACs time-of-flight (ToF)
with two pick-ups is used:
 the velocity  is measured.
Example: Time-of-flight signal from
two pick-ups at 1.4 MeV/u:
The reading is tscope= 15.82(5) ns with
frf = 36.136 MHz  T = 27.673 ns
L = 1.629(1) m & N = 3
= 0.05497(7)
 Ekin =1.407(3) MeV/u
The accuracy is typically 0.1 %
i.e. comparable to Δp/p
L
LINAC
RFQ
source
BPM
pick-up 1
BPM
pick-up 2
ADC

## Slide 54
Longitudinal Emittance by linear Transformation using a Buncher
Longitudinal focusing:
Variation of the bunch shape by a rf-buncher
 components 5 and 6 from 6-dim phase-space
Transversal correspondence:
Quadrupole variation
 Transfer matrix of buncher & drift:
with focal length
 Variation of buncher amplitude U
 different bunch width at pick-up location

## Slide 55
Result of a longitudinal Emittance Measurement
Example GSI LINAC: Voltage variation at buncher for 11.4 MeV/u Ni14+ beam, 31 m drift:
The structure of short bunches can
be determined with special monitor
This example: The resolution is better
than 50 ps or 2o for 108 MHz
Typical bunch length at proton LINACs:
bunch  10 to 300 ps
Determination of longitudinal emittance possible
Application for synchrotron injection:
Shaping of longitudinal phase space by buncher
i.e. long bunches  low momentum spread
to match to the synchrotron long acceptance
LINAC
RFQ
source
Bunch Shape Monitor
B.
Buncher
BSM
extraction
injection
p synchrotron
de-bunched
multi-turn
injection

## Slide 56
Low Velocity Effect: General Consideration
Lorentz transformation of single point-like charge:
Lorentz boost and transformation of time:
Trans. Elab.-frame of a point charge:
Long. E|| lab.-frame of a point charge:

## Slide 57
Bunch Structure at low Ekin: Not possible with Pick-Ups
Pick-ups are used for:
precise for bunch-center relative to rf
course image of bunch shape
But:
For  << 1  long. E-field significantly modified:
Example: Comparison pick-up – particle counter:
Ar beam of 1.4 MeV/u (β = 5.5%) , frf = 108 MHz
 the pick-up signal is insensitive
to bunch ’fine-structure’
ampl.
R
  R / 
1/ frf = 9.2 ns
LINAC
RFQ
source
Bunch Shape Monitor
B.
Buncher
BSM

## Slide 58
Bunch Structure using secondary Electrons for low Ekin Protons
Secondary e− liberated from a wire carrying the time information.
→ Bunch Shape Monitor (BSM)
Working principle:
insertion of a 0.1 mm wire at  10 kV
emission of secondary e− within less than 10 ps
secondary e− are accelerated
toward an rf-deflector
rf-deflector as ’time-to-space’ converter
detector with a thin slit
slow shift of the phase
resolution  10 ps  1o @ 280 MHz
Measurements are comparable
to that obtained with particle detectors.
SEM: secondary electron multiplier

## Slide 59
Realization of Bunch Shape Monitor at CERN LINAC2
ion beam
rf-deflector
movable
HV wire
Flange Ø 150 mm
electron detector
electrons
Example: The bunch shape behind RFQ with120 keV/u:
back: Conclusion
LINAC
RFQ
source
Bunch Shape Monitor
B.
Buncher
BSM

## Slide 60
Longitudinal Emittance using tomographic Reconstruction
original object
Iterative projection
& back-projection
after sufficient
iterations
1st back-
projection
Filtered back projection:
Iterative process by redistributing
the 2-dim image and considering the
differences to the previous iteration step.
Tomography: Generation of a n-dim image
from many lower-dim projections
Medical application: 2-dim reconstruction using sufficient 1-dim projections

## Slide 61
Longitudinal Emittance using tomographic Reconstruction
original object
Iterative projection
& back-projection
after sufficient
iterations
1st back-
projection
Bunch observation:
Each revolution, the bunch shape differs
due to synchrotron oscillations.
Fulfilled condition: fsynch << frev .
Application at accelerators:
Transverse & longitudinal
emittance reconstruction in transfer lines
Longitudinal bunch evolution in synchr.
Tomography: Generation of a n-dim image
from many lower-dim projections
Medical application: 2-dim reconstruction using sufficient 1-dim projections
extraction
injection
synchrotron
Current trans.
FCT or WCM

## Slide 62
Results of tomographic Reconstruction at a Synchrotron I
Typical bucket filling.
Important knowledge for bunch ’gymnastics’.
Bunches from 500 turns at CERN PS and phase space for the first time slice,
measured with a Wall Current Monitor:
extraction
injection
synchrotron
Current trans.
FCT or WCM

## Slide 63
Results of tomographic Reconstruction at a Synchrotron II
Bunches from 500 turns at CERN PS and phase space for time slice of shortest bunch,
measured with a Wall Current Monitor WCM:
Fast rise of rf-amplitude for ‘bunch-rotation’:
Bunch is compressed in time
Filamentation due to amplitude- dependent synchrotron frequency
extraction
injection
synchrotron
Current trans.
FCT or WCM
 Conclusion

## Slide 64
Outline for the Lecture Beam Diagnostics
Outline
General remarks to beam instrumentation and diagnostics
Overview for beam current measurement
Various method for transverse profile determination
High-current example 1 and example 2
Emittance determination
High-current example 3
Usage of BPMs and longitudinal measurements
High-current example 4
Beam Loss Detection

## Slide 65
Beam Loss Monitors
When energetic beam particles penetrates matter, secondary particles are emitted:
this can be e−, , protons, neutrons, excited nuclei, fragmented nuclei...
 Spontaneous radiation and permanent activation is produced.
 Large variety of Beam Loss Monitors (BLM) depending on the application.
Protection: Sensitive devices e.g. super-conducting magnets to prevent quenching (energy absorption by electronic stopping)
 interlock signal for fast beam abortion.
Operation: Alignment of the beam to prevent for activation
 optimal transmission to the target.
Several loss monitors are used for various secondaries
Different count rate & time resolution
Some applications for usage

## Slide 66
Halo Development and Cleaning at LINACs
Halo development at LINACs caused by:
Higher-order magnet fields (e.g. aberration)
Transverse mis-match
Off-momentum particles due to wrong acceleration
Space charge forces
Collimators:
Cut the beam transverse tails
to prevent loss downstream
Example SNS LINAC:
Scraping at 3 MeV
Profile measurement at 40 MeV
Problem: Activation prevents hands-on maintenance
Goal: Halo cutting at low energy to limit activation
A. Aleksandrov et al., PAC 05
M. Plum, in JAS School on Beam Loss and Acc. Protection 2014, CERN-2016-002
-6
-4
0
-2
4
2
6
Position in units of
10-3
10-1
10-2
100
norm. current
Halo
Operation:
Trans. & long. matching to prevent for halo generation
BLMs for detection at loss location ( mis-match origin)
See lecture ‘Collimation’ by N. Fuster

## Slide 67
Basic Idea of Beam Loss Monitors
Basic idea for Beam Loss Monitors BLM:
A loss beam particle must collide with the vacuum chamber or other insertions
Interaction leads to some shower particle:
e−, , protons, neutrons, excited nuclei, fragmented nuclei
Detection of these secondaries by an appropriate detector outside of beam pipe
Relative cheap detector installed at many locations
Remark: Due to grazing angle a thin vacuum chamber might be a ‘thick target’
BLM detector
front-end
electronics
digitalization
& fast analysis
interlock
display
lost beam particle
Secondary products:
Electromagnetic or hadronic shower products
Charged particles
Neutrons or 
vacuum pipe
beam
See lectures ‘Beam Loss Mechanisms & Machine Protection’ by A. Nordt &
‘Beam Loss Consequences’ by G. Lerner

## Slide 68
Scintillators as Beam Loss Monitors
Plastics or liquids are used:
Detection of charged particles
by electronic stopping
Detection of neutrons
by elastic collisions n on p in plastics
and fast p electronic stopping.
Scintillator + photo-multiplier:
counting (large PMT amplification)
or analog voltage ADC (low PMT amplification)
Radiation hardness:
plastics 1 Mrad = 104 Gy
liquid 10 Mrad = 105 Gy
Scintillator
2x2x5 cm3
Photo-multiplier
inside
HV base

## Slide 69
Ionization Chamber as Beam Loss Monitors
Energy loss of charged particles in gases → electron-ion pairs → current meas.
W is average energy for creation
for one e- -ion pair:
Gas | Ionization Pot. [eV] | W-Value [eV]
Ar | 15.7 | 26.4
N2 | 15.5 | 34.8
O2 | 12.5 | 30.8
Air |  | 33.8
shower particle
Sealed tube Filled with Ar or N2 gas:
Creation of Ar+-e− pairs,
average energy W = 32 eV/pair
measurement of this current
Slow time response
due to  10 μs drift time of Ar+.
Per definition: Direct measurement of dose !
15 cm
Notes: 69

## Slide 70
Ionization Chamber as BLM: TEVATRON and CERN Type
15 cm
38 cm
Parameter | TEVATRON, RHIC | CERN type
Size | 15cm,  6 cm | 50 cm,  9 cm
Gas | Ar at 1.1 bar | N2 at 1.1 bar
# of electrodes | 3 | 61
Voltage | 1000 V | 1500 V
Reaction time | 3 µs | 0.3 µs
TEVATRON, RHIC type
CERN type
4000 BLMs at LHC each  6m

## Slide 71
Comparison of different Types of BLMs
Different detectors are sensitive to various physical processes
very different count rate, but basically proportional to each other
Typical choice of the detector type:
Ionization Chamber:
Advantage:
- Measurement of absolute dose
Disadvantage:
- Low signal (low , eff, no neutron detection),
- Sometimes slow, ion drift time 10 ... 100 µs
 Often used at many location for interlock
Scintillator
2x2x5 cm3
Photo-multiplier
inside
HV base
Scintillator:
Advantage:
- Fast current reading or particle counting
- Can be fabricated in any shape, cheap
Disadvantage:
- Need calibration in many cases
- Might suffer from radiation
 Often used at sensitive insertions
38 cm

## Slide 72
Conclusion for Beam Diagnostics Course
Diagnostics is the ’sensory organ’ for the beam.
It required for operation and development of accelerators
Several categories of demands leads to different installations:
Quick, non-destructive measurements leading to a single number or simple plots
Complex instrumentation used for hard malfunction and accelerator development
Automated measurement and control of beam parameters i.e. feedback
High-current effects might modify traditional methods and their interpretation
The goal and a clear interpretation of the results is a important design criterion.
General comments:
Quite different technologies are used, based on various physics processes
Accelerator development goes parallel to new diagnostics demands
 Diagnostics contributes to accelerator development.
Thank you for your attention!
Do you have some remarks?

## Slide 73
Backup slides

## Slide 74
Stripline BPM: General Idea
For short bunches, the capacitive button deforms the signal
Relativistic beam   1  field of bunches nearly TEM wave
Bunch’s electro-magnetic field induces a traveling pulse at the strips
Assumption: Bunch shorter than BPM, Zstrip = R1 = R2=50  and vbeam = cstrip
LHC stripline BPM, l = 12 cm
From C. Boccard, CERN

## Slide 75
Stripline BPM: General Idea
For relativistic beam with   1 and short bunches:
Bunch’s electro-magnetic field induces a traveling pulse at the strip
Assumption: lbunch<< l, Zstrip=R1=R2=50  and vbeam=cstrip
Signal treatment at upstream port 1:
Signal at downstream port 2: Beam induced charges cancel with traveling charge from port 1
 Signal depends on direction  can distinguish between counter-propagation beams
t=0: Beam induced charges at port 1:
 half to R1, half toward port 2
t=l/c: Beam induced charges at port 2:
 half to R2, but due to different sign,
it cancels with the signal from port 1
 half signal reflected
t=2∙l/c: reflected signal reaches port 1
If beam repetition time equals 2∙l/c: reflected preceding port 2 signal cancels the new one:
 no net signal at port 1

## Slide 76
Stripline BPM: Transfer Impedance
The signal from port 1 and the reflection from port 2 can cancel  minima in Zt .
For short bunches Ibeam(t)  Ne ∙ (t):
Zt show maximum at l=c/4f=λ/4 i.e. ‘quarter wave coupler’ for bunch train
 l has to be matched to vbeam
No signal for l=c/2f=λ/2 i.e. destructive interference with subsequent bunch
Around maximum of |Zt|: phase shift φ=0 i.e. direct image of bunch
fcenter=1/4 ∙ c/l ∙ (2n-1). For first lope: flow=1/2∙fcenter, fhigh=3/2 ∙ fcenter i.e. bandwidth 1/2∙fcenter
Precise matching at feed-through required t o preserve 50  matching.
Stripline length l=30 cm, α=100

## Slide 77
Stripline BPM: Transfer Impedance
The signal from port 1 and the reflection from port 2 can cancel  minima in Zt .
For bunches of length :
Stripline length l=30 cm, α=100
Zt(ω) decreases for higher frequencies
If total bunch is too long ±3σt > l destructive interference leads to signal damping
Cure: length of stripline has to be matched to bunch length
Further advantage: Linear phase propagation  good for coupled bunch feedback

## Slide 78
Comparison: Stripline and Button BPM (simplified)
 | Stripline | Button
Idea | traveling wave | electro-static
Requirement | Careful Zstrip = 50  matching | 
Signal quality | Less deformation of bunch signal | Deformation by finite size and capacitance
Bandwidth | Broadband, but minima | Highpass, but fcut < 1 GHz
Signal strength | Large Large longitudinal and transverse coverage possible | Small Size <3cm, to prevent signal deformation
Mechanics | Complex | Simple
Installation | Inside quadrupole possible improving accuracy | Compact insertion
Directivity | YES | No
From . S. Vilkins, D. Nölle (DESY)
FLASH BPM inside quadrupole

## Slide 79
button
Beam envelope
Image envelope
Estimation of finite Beam Size Effect for Button BPM
Ideal 2-dim model:
Due to the non-linearity, the beam size enters in the position reading.
Finite beam size:
Calculation of signal response
at different location
‘Averaging’ of image position
 Cannot be corrected !
Remark: For most LINACs: Linearity is less important, because beam has to be centered
Position correction as feed-forward for next macro-pulse.
Image center
beam center

## Slide 80
Ibeam
Time Domain  Frequency Domain: Instrumentation
Time domain: Recording of a voltage as a function of time:
Frequency domain: Displaying of a voltage as a function of frequency:
Instrument:
Spectrum Analyzer
Instrument:
Oscilloscope
courtesy company Keysight
voltage
time
frequency
Photos and idea by Piotr Kowina
Fields seems unstructured
Other perspective:
Fields is structured

## Slide 81
Ibeam
Time Domain  Frequency Domain: Mathematics
Time domain: Recording of a voltage as a function of time:
Frequency domain: Displaying of a voltage as a function of frequency:
Mathematics for function :
Fourier Transformation:
Fourier Transformation:
Contains amplitude & phase
as (complex in math. sense)
The same information is displayed differently
Law of Convolution: For a convolution in time:
  convolution be expressed as multiplication of FTs

## Slide 82
Close Orbit Feedback: Results
Orbit feedback:
Example: 12 beam positions at GSI-SIS during ramping from 8.6 to 500 MeV/u for Ar18+
Procedure:
1. Position from all 12 BPMs
2. Calculation of corrector setting on fast (FPGA-based) electronics
3. Submission to corrector magnets
4. New position measurement
 regulation time down to 10 ms
feedback off
time [ms]
400
300
100
0
position [mm]
20
10
0
-10
200
feedback on
time [ms]
0
100
300
400
position [mm]
20
10
0
-10
200
Rule of thumb:
Movement related to tune i.e. ‘natural oscillations by periodic focusing’
To determine the ‘sine-like’ oscillation, 4 BPMs per peroid are required
 4 BPMs per tune value (but detailed investigation required to determine the # of BPMs)
position x(t) at 12 BPMs

## Slide 83
β-Function Measurement from Bunch-by-Bunch BPM Data
Excitation of coherent betatron oscillations:
 Time-dependent position reading results the phase advance between BPMs
The phase advance is:
 - function from
Remark: Determination of -function with 3 BPMs:
See e.g.: R. Tomas et al., Phys. Rev. Acc. Beams 20, 054801 (2017)
A. Wegscheider et al., Phys. Rev. Acc. Beams 20, 111002 (2017)
Conclusion

## Slide 84
‘Beta-beating’ from Bunch-by-Bunch BPM Data
Result concerning ‘beta-beating’:
Model doesn’t fit reality completely
e.g. caused by misalignments
Corrections executed
Increase of the luminosity
Remark:
Measurement accuracy depends on
BPM accuracy
Numerical evaluation method
Example: ‘Beta-beating’ at BPM with measured meas & calculated model
for each BPM at BNL for RHIC (proton-proton or ions circular collider with 3.8 km length)
From X. Shen et al.,
Phys. Rev. Acc. Beams 16, 111001 (2013)
Conclusion

## Slide 85
From J. Wenninger: CAS on BD, CERN-2009-005 & J. Wenninger CERN-AB-2004-009
Dispersion Measurement from Closed Orbit Data
Dispersion D(si ): Change of momentum p by detuned rf-cavity
 Position reading at one location :
 Result from plot of xi as a function of Δp/p  slope is local dispersion D(si )
Example: Dispersion measurement D(s) at BPMs at CERN SPS
Theory-experiment correspondence
after correction of
BPM calibration
quadrupole calibration
: model expectation
green histogram: measurement
6.9 km
Example 4

## Slide 86
However: Not only noise contributes but additionally beam movement by betatron oscillation
 broadband processing i.e. turn-by-turn readout for tune determination.
Comparison: Filtered Signal ↔ single Turn
Example: GSI Synchr.: U73+, Einj = 11.5 MeV/u Eout = 250 MeV/u within 0.5 s, 109 ions
Position resolution < 30 μm (BPM diameter d=180 mm)
average over 1000 turns corresponding to 1 ms
or 1 kHz bandwidth
Turn-by-turn data have much larger variation
1000 turn average for closed orbit
Variation < 10 μm (sufficient for application)
Single turn e.g. for tune
Variation  150 μm

## Slide 87
)
Amplitude:
Frequency:
Jk : Bessel‘s functions
s : synchrotron frequency
Phase:
Spectrum: frequency modulation
Longitudinal Spectrum of non-interacting Particles

## Slide 88
Sideband amplitude given by
Bessel-function Jk of order k
b =Q0 betatron freq. 0 rev. freq. a bunch length
Transverse Spectrum of non-interacting Particles
Spectrum comprises:
betatron = amplitude x(t)
synchrotron = phase I(t) modulation
Finite chromaticity   0 modifies sideband’s heights !
Frequency of harmonics m & mode k
Chromaticity :

## Slide 89
Non -zero chromaticity
 intra-bunch movement
Non -zero chromaticity  intra-bunch movement
Reason: Head p > 0 ; tail p < 0
Eigen-functions:
Rigid oscillations for zero chromaticity
Example: five consecutive turns
Turn 1 2 3 4
Time Domain Mode Structure: Non-interacting Particle Model
Time  / ns
Time  / ns
Chromatic tune freq.
SIS18 injection:
Q0 = 3.27,  = 0.92
Bunch center-of-mass as can be determined by BPMs
1
2
3
4
5
 = 0
Time  / ns
Time  / ns
 = 1
 = 2
k = 0
k = 1