## Slide 1: Hands-On CAS Intensity Limitations in Hadron Beams
22nd & 23rd June 2025
Alexandre Lasheen, Christine Vollinger, Elena de la Fuente, Heiko Damerau, Jake Flowerdew, Leandro Intelisano, Michela Neroni
(all at CERN)

## Slide 2: Case Study: SPS Beam Wire Scanners (BWS)
Brief recap of how rotational beam wire scanners works
Super Proton Synchrotron (SPS) BWS failures & beam parameters during scrubbing run
BWS modifications in CST & consequences
All the wire scanner design has been carried out by experts in the Beam Instrumentation Group at CERN and is their intellectual property.
BIG THANKS for letting us use this instrument as a CAS Hands-on example
and to the TF colleagues who let us use all their contributions!

## Slide 3: Case Study: SPS Beam Wire Scanners (BWS)
Brief recap of how rotational beam wire scanners works
Fast rotation of the wire around 270 degrees at a speed of up to 20 ms-1
Cutting through the beam leads to particle scattering, detected by downstream scintillators
From the scattering products, geometrical beam size and transverse emittance is determined.
SPS is equipped with 4 BWS, 2 horizontal scanners & 2 vertical scanners
Beam
Motor mechanism
(outside vacuum)
Titanium holders
(“forks”)
Motor driven
hollow shaft
Beam vacuum pipe
Cylindrical
vacuum tank
Carbon fiber (“wire”)

## Slide 4: Case Study: SPS Beam Wire Scanners (BWS)
Brief recap of how rotational beam wire scanners works

## Slide 5: Case Study: SPS Beam Wire Scanners (BWS)
Brief recap of how rotational beam wire scanners works
Particularly worrying:
The wire broke without contact to the beam in the so-called parking position.
Beam
Motor mechanism
(outside vacuum)
Titanium holders
(“forks”)
Motor driven
hollow shaft
Beam vacuum pipe
Cylindrical
vacuum tank
Carbon fiber (“wire”)
Beam

## Slide 6: Case Study: SPS Beam Wire Scanners (BWS)
SPS BWS failures & beam parameters during scrubbing run
BWS have shown excellent performance with more than 70’000 scans during the first year!
Until… 12-April-23, when all 4 BWS failed simultaneously during a scrubbing run… WHAT HAPPENED?
What was observed after BWS opening:
No wire left, No remnants found.
(2) No other damage to the BWS body.
(3) Traces on the wire holding “forks” and the connecting parts.
(4) No full logging data on the exact time of breakage...
Makes it difficult to identify the exact beam parameters.

## Slide 7: Case Study: SPS Beam Wire Scanners (BWS)
Changed installation configuration:
1 vertical WS and 1 horizontal BWS
Wire scanners are installed in pairs: 2 horizontal and 2 vertical together.
Area of vertical BWS installation was not accessible due to radiation levels.

## Slide 8: Measured thermal gradient on MKP-S
4 BWS broke?
2 newly installed BWS broke
 We use the MKP-S heating rate as an indicator of the beam parameters.
2 spare BWS were installed, beam settings of the incident were repeated: wires broke almost “instantaneously”
but this time all beam parameters were logged.
4x72 bunches up to
2e11 ppb and σz ~1.6ns, Long FT
4x72 bunches up to
1.8e11 ppb and σz ~1.6 ns
Short FT

## Slide 9: Comparisons of bunch parameters
SPS RF settings were restored to reproduce beam conditions during BWS breakage (22-Apr),
Bunch profiles were acquired with MD scope and Head-tail monitor,
Beam spectra around 800 MHz were recorded with a spectrum analyzer connected to AEW pick-up (= wall current monitor).
Mean bunch lengths are very similar despite a slightly larger spread of maximum bunch length for higher intensity
Beam spectra acquired during both MDs should be sufficiently close
BQM bunch length (s)
MD 04.05.2023
MD 17.05.2023
Logged data
min
mean
max
4x72 bunches with
min
mean
max
BQM bunch length (s)

## Slide 10: Task Force: SPS Beam Wire Scanners
EM Simulations targeting bci, and mitigation measures,
Update on operational beams, calculation of beam spectra
Heating and losses study,
RF measurements on a spare,
Thermal-mechanical simulations, post failure analysis, preparation for next intervention.
Not all this work of the task force can be repeated in this case study, of course, but we (you) will do:
EM Simulations targeting mitigation measures (CST hands-on),
Calculation of different beam spectra (Python hands-on),
Heating/Loss study (CST & Python hands-on “combined”).

## Slide 11: Case Study: SPS Beam Wire Scanners (BWS)
Final solution – watch out – spoiler alert!

## Slide 12: Case Study: SPS Beam Wire Scanners (BWS)
Quick mitigation was required without impacting on the BWS functionality, without compromising the vacuum features, without leading to excessive heating,…
Decision to introduce a new loss mechanism by inserting (already know, vacuum tested & approved) lossy ferrite material with the goal to pull EM-fields “away” from the wire, re-distributing the losses.
Impact of ferrite installation on calculated impedance

## Slide 13: Case Study: SPS Beam Wire Scanners (BWS)
Impedance traces were used to determine the overall ferrite impact and the field distributions in the BWS tank.
H-field distribution on the wire surface was used as criteria to see the impact on the wire loss reduction.
Overall, the use of 5 tiles or 6 tiles showed to be the best solutions, however, each tile produces heating in the tank (so less ferrite use is desired), and
Both solutions have insufficient impact at low frequency!

## Slide 14: Case Study: SPS Beam Wire Scanners (BWS)
RF-coupler impact
Beam spectrum was overlayed with the calculated impedance.
Several spectra were tested for compliance.
Heating in the ferrites was estimated.
Decision: additional installation of an RF-coupler is required to impact on the one low frequency mode that could not be tackled with ferrite insertion.

## Slide 15: Case Study: SPS Beam Wire Scanners (BWS)
H-BWS: 6 Tiles
V-BWS: 5 Tiles + coupler
Power loss studies on the wire were carried out, to decide on the final solution.

## Slide 16: Some observations during MDs after installation of the two mitigated BWS
~2.2e11 ppb at FT with 4x72 bunches, standard FT
Case Study: SPS Beam Wire Scanners (BWS)
beam intensity
average bunch length
Temperature in the BWS

## Slide 17: Case Study: SPS Beam Wire Scanners (BWS)
RF - COUPLER
FERRITES

## Slide 18: SPS Beam Wire Scanners Task Force
Contributions from:
Alexandre Lasheen, Alexandre Mariet, Ana Guerrero, Anthony Harrison, Benoit Salvant, Carlo Zannini, Christine Vollinger, Elena de la Fuente, Federico Carra, Federico Roncarolo, Francesco Velotti, Giannis Papazoglou, Giulia Papotti, Heiko Damerau, Ivan Karpov, Jonathan Emery, Jose Antonio Ferreira Somoza, Kevin Li, Leonardo Sito, Michela Neroni, Michael Stephen Sullivan, Nabil El-Kassem, Raymond Veness, Thibaut Lefevre, William Andreazza
… and hopefully, I did not forget somebody.
Thank you for your attention!