## Slide 1: Impedance mitigation of SPS rotary beam wire scanners with CST Studio Suite
Heiko Damerau, Elena de la Fuente Garcia, Jake Flowerdew, Leandro Intelisano, Alexandre Lasheen, Michela Neroni*, Christine Vollinger
Based on the work of the SPS BWS Task Force
Many thanks to: A. Mostacci, L. Sito and all the SPS BWS Task Force contributors
*michela.neroni@cern.ch

## Slide 2: References
    IPAC24 papers:
      R. Veness et al., “Overview of beam intensity issues and mitigations in the CERN-SPS fast wire scanners”, 10.18429/JACoW-IPAC2024-WEPG26
      L. Sito et al., “Impedance and thermal studies of the CERN SPS wire scanners and mitigation of wire heating”, 10.18429/JACoW-IPAC2024-WEPG29
      C. Vollinger et al., “Mitigation of beam coupling impedance for the wire scanners in the CERN Super Proton Synchrotron”, 10.18429/JACoW-IPAC2024-THPC51
    Additional material and presentations:
      SPS BWS Task Force meeting - 02062023
      SPS BWS Task Force meeting – 02082023
      SPS BWS Task Force meeting – 28082023
      SPS BWS Task Force meeting - 14112023

## Slide 3: CST simulation files
“1-SPS_WS_Wake_NotMitigated.cst” contains the component geometry.
“1-SPS_WS_Wake_NotMitigated_solved.cst” contains Wake Impedance results.
“1-SPS_WS_Wake_NotMitigated_FieldMon.cst” contains electric and magnetic Field Monitors solved.
Each folder contains three files
N.B. Please do not run the simulations on your laptop during the tutorial!

## Slide 4: 1- Not mitigated case

## Slide 5
Rotating shaft (stainless-steel)
Carbon wire
Particle beam
SPS beam wire scanner geometry as before the 2023 breakage, with fork in parking position:
Stainless steel case
Rotary Beam Wire Scanner – 3D model

## Slide 6
Simplification and/or rebuilding of geometrical shapes in CST
Rotary Beam Wire Scanner – 3D model

## Slide 7
Simplification and/or rebuilding of geometrical shapes in CST
Fork in scan position
Rotary Beam Wire Scanner – 3D model
The modelling is not part of this tutorial  please have a look at the Simulation tutorial from RF CAS: “Numerical analysis of RF problems”

## Slide 8
Rotary Beam Wire Scanner – 3D model
You can access components and materials from the Navigation Tree

## Slide 9: Simulation settings
Definition of Particle beam
Mesh settings
Definition of Frequency range, Background material and Boundary conditions
Definition of Wakelength and frequency sampling of the impedance trace
Disclaimer: an example of simulation will be provided; different settings could lead to an equally accurate result.

## Slide 10: Simulation settings (1/4)
Definition of Particle Beam:
We select a sigma of 100 mm
To click on “Wakefields”
Tip
You can try to change the sigma to explore the higher frequency range.
Try to change the injection scheme to Analytics and see the changes.
The bunch length will define the maximum simulation frequency. In CST:
assuming an ultrarelativistic bunch.
To have a maximum frequency of 1 GHz

## Slide 11: Simulation settings (1/4)
Definition of integration method:
To choose Indirect testbeams as integration method
N.B. If the structure does not allow the use of Indirect testbeams, like in this case, the solver will automatically select Direct integration method.

## Slide 12: Simulation settings (2/4)
Time domain solvers use Hexahedral mesh type. Mesh settings:
10 cells per wavelength (default option)
Sub-optimal mesh definition introduced to reduce computational time.
Mesh convergence study, gradually increasing the number of mesh cells, is always needed to achieve accurate results.

## Slide 13: Simulation settings (2/4)
Mesh settings:
10 cells per wavelength (default option)
Local refinement on the wire  3 cells across object in each direction

## Slide 14: Simulation settings (3/4)
Frequency range from 0 to 1 GHz
Background material  PEC
Boundaries  - electric boundaries (along x and y
      - open* along z
*The open boundary extends the touching geometry virtually to infinity so that waves can pass this boundary with minimal reflections

## Slide 15: Simulation settings (4/4)
Setup of Wakefield solver:
Simulated wakelength = 200 m

## Slide 16: Simulation settings (4/4)
Selecting 10001 samples for impedance trace post-processing
The simulation can now be started.

## Slide 17: Longitudinal impedance result
Up to 10 kΩ
Impedance trace is not fully resolved and a longer wakelength would be needed. Truncation error leads to ripples according to the Fourier Transformation.

## Slide 18: Longitudinal impedance result
The largest impedance resonance can be represented with a resonator model (, , )
Used as impedance source for power loss computation (case study, part 2 by A. Lasheen et al.)

## Slide 19: Longitudinal wake potential
The wake potential in CST corresponds to the wake function of a charge distribution (nomenclature used in A. Mostacci lectures)
An additional convergency check can be done by looking at the Energy under 1D results. Good convergency reached when energy is below -40 dB (rule of thumb)

## Slide 20
Computational resources vs. settings
© CST, Dassault Systèmes
Linear dependence on the number of mesh cells with the Time Domain solver
Linear dependence on the wakelength
Trade-off between the total number of mesh cells and the wakelength is needed!

## Slide 21: Field monitors – settings (1/2)
Definition of E-field and H-field monitors at the resonant frequencies.
Example:
GHz
GHz
GHz
GHz

## Slide 22: Field monitors – settings (2/2)
Shorter wakelength can be selected to solve field monitors.
Select “Store per solid” to obtain volume losses

## Slide 23
Field monitors - results
H-field monitor at GHz

## Slide 24
Field monitors - results
H-field monitor at GHz

## Slide 25
Material and solid losses
A summary of the computed losses can be found under 1D Results

## Slide 26
Surface losses
Surface losses are computed starting from H-field monitors and material properties:
Power dissipated in the resonant structure. It is directly related to the magnetic field H:
Energy stored in the resonant structure
Perturbation approach:
computed assuming all the metals perfect conductors.
is the surface resistance, related to material properties
is the skin depth
is the electrical conductivity

## Slide 27
Surface losses
Percentage of losses for each material/solid can be extracted and they are used to rescale the impedance with respect to each material

## Slide 28
Thermal losses
The file will contain the Electric Volume Losses and Surface Losses computed for each field monitor at the specific frequency

## Slide 29: 2- Adding 5 Ferrite Tiles

## Slide 30
Rotating shaft (stainless-steel)
Carbon wire
Adding 5 ferrite tiles – 3D model
Particle beam
Ferrite
Magnetic permeability versus frequency for ferrite material:
Assumption: placing ferrite tiles in regions with high magnetic field

## Slide 31: Simulation settings
Definition of Particle beam  same as Not Mitigated case
Mesh settings
Definition of Frequency range, Background material and Boundary conditions  same as Not Mitigated case
Definition of Wakelength and frequency sampling of the impedance trace

## Slide 32: Simulation settings (2/4)
Time domain solvers use Hexahedral mesh type. Mesh settings:
20 cells per wavelength
Local refinement on the wire  3 cells across object in each direction

## Slide 33: Simulation settings (4/4)
Setup of Wakefield solver:
Simulated wakelength = 100 m
To click on ”Special”  number of samples set to 6001

## Slide 34: Longitudinal impedance result
Field monitors can be defined at the frequencies of resonance
Low frequency peak not fully resolved, the actual shunt impedance is higher

## Slide 35: Comparison of longitudinal impedance traces

## Slide 36
Field monitors - results
H-field monitor at GHz
H-field monitor at GHz
Losses are moved from the wire to the ferrite tiles

## Slide 37: 3- Adding 6 Ferrite Tiles

## Slide 38
Rotating shaft (stainless-steel)
Carbon wire
Adding 6 ferrite tiles – 3D model
Particle beam
Ferrite
Assumption: adding one more ferrite tile (more losses), just below the carbon wire

## Slide 39: Simulation settings
Definition of Particle beam  same as Not Mitigated case
Mesh settings
Definition of Frequency range, Background material and Boundary conditions  same as Not Mitigated case
Definition of Wakelength and frequency sampling of the impedance trace

## Slide 40: Simulation settings (2/4)
Time domain solvers use Hexahedral mesh type. Mesh settings:
20 cells per wavelength
Local refinement on the wire  3 cells across object in each direction

## Slide 41: Simulation settings (4/4)
Setup of Wakefield solver:
Simulated wakelength = 100 m
To click on ”Special”  number of samples set to 6001

## Slide 42: Longitudinal impedance result
Field monitors can be defined at the frequencies of resonance

## Slide 43: Comparison of longitudinal impedance traces
Not much gain on the impedance reduction when adding one more ferrite tile.
Low frequency resonance is not touched with respect to the 5 tiles case.

## Slide 44: 3- Adding 5 Ferrite Tiles and RF coupler

## Slide 45
Adding 5 Ferrite Tiles and RF coupler – 3D model
Rotating shaft (stainless-steel)
Carbon wire
Particle beam
Ferrite
Vacuum
RF coupler
Assumption: to use ferrite tiles and to add an RF coupler at the available port

## Slide 46: Simulation settings
Definition of Particle beam  same as Not Mitigated case
Mesh settings
Definition of Frequency range, Background material and Boundary conditions  same as Not Mitigated case
Definition of Wakelength and frequency sampling of the impedance trace
Addition of a waveguide port

## Slide 47: Simulation settings (2/5)
Time domain solvers use Hexahedral mesh type. Mesh settings:
20 cells per wavelength
Local refinement on the wire  3 cells across object in each direction

## Slide 48: Simulation settings (4/5)
Setup of Wakefield solver:
Simulated wakelength = 100 m
To click on ”Special”  number of samples set to 6001

## Slide 49: Simulation settings (5/5)
Select the face at the coupler port  from the CST working plane, click “F” on your keyboard (shortcut)
Define a waveguide port

## Slide 50: Longitudinal impedance result
Field monitors can be defined at the frequencies of resonance

## Slide 51: Comparison of longitudinal impedance traces
The RF coupler is very effective in the low frequency range and it reduces further the impedance around 800 MHz.

## Slide 52: Power outgoing from the coupler port

## Slide 53
Field monitors - results
H-field monitor at low frequency (around 160 MHz)
Not mitigated
Mitigated case

## Slide 54
Field monitors - results
H-field monitor at impedance peak around 800 MHz
Not mitigated
Mitigated case

## Slide 55: Impedance comparison

## Slide 56: Comparison of longitudinal impedance traces

## Slide 57: Supporting slides
With material from A. Mostacci lecture on “Bench Measurements and Simulations of Beam Coupling Impedance”

## Slide 58: The CST wakefield solver

## Slide 59
The CST wakefield solver
Creation of a new CST file and selection of wakefield solver.

## Slide 60
Select Wakefield solver
Default background material (PEC) and units
You can click to make modifications
The CST wakefield solver
Creation of a new CST file and selection of wakefield solver.

## Slide 61
Definition of a particle beam

## Slide 62: Selection of integration method

## Slide 63: Mesh settings

## Slide 64: Definition of wakelength
of a charge distribution