## Slide 1
Resonant Converter
The Power Supply for the SWLS Project
Lucas Carnevalli de Almeida
CNPEM’s Superconducting Wavelength Shifter (SWLS)
Technical Design Review
4–5th November 2024
Power Converters Group
On behalf of the Technology Unit

## Slide 2
Summary
Motivation
The Resonant Converter
Results
  Article Results
  Preliminary Version
Next Steps
References

## Slide 3
Motivation
Figure 1. Danfysik Power Converter
SCW Sources: Danfysik MP8443 (Figure 1)
SPECIFICATIONS:
3 kW, 300 A and 10 V, 25 ppm stability and 1A minimum current offset;
ISSUES:
Obsolete hardware and software; No technical support and low reliability;
Resonant Converter: In-house development
TARGET:
300 A output with voltage less than 10 V, ripple less than 10 ppm (with 3 kHz band) and similar minimum current; New technologies were implemented, and all features were developed in-house;
Figure 2. Final PCB Design for the Resonant Converter

## Slide 4
The Resonant Converter
Specifications for the project
Output current: ≤300 A;
Output voltage: ≤10 V;
Long-term stability and ripple of the output current: ≤20 ppm;
Is composed mainly of parasitic elements, which can be done by using a secondary-side resonance technique (Figure 3) [3];
Zero Current Switching (ZCS) to reduce switching losses (Figure 4).
Figure 3. Example of the Simplified Resonant Converter Circuit [5]
Figure 4. Simplified Resonant Converter Waveforms [5];

## Slide 5: Results - Article

## Slide 6
The AC RMS ripple of the output current was , which represents
Article Results
Figure 6. AC RMS ripple [5];

## Slide 7
Output current of was reached (Figure 7 and Figure 8);
Figure 7. Output current as a function of the switching frequency in open-loop operation [5]
Figure 8. Waveforms in closed-loop operation (current of the resonant inductor; gate-source voltage of the high-side MOSFET; output current) [5]
Article Results
f(x) = Ax + B
However, the current output wasn’t completely linear, and it didn’t reach zero

## Slide 8: Preliminary Version
Lower current output
IMPROVEMENTS
Figure 9. Preliminary version output
278A
Higher oscillations
Figure 10. Converter Installation

## Slide 9: Preliminary Version
CLAMP CIRCUIT
It was designed to allow us to achieve lower output currents while effectively reducing intermodal oscillation.
The following parameters presented in the Figure 12 are defined:
  T_on: ON pulse width of the main MOSFETs (M1 and M2) DT_1: dead-time to turn on the main keys (M1 and M2) DT_2: dead-time to turn on the clamp keys (M3 and M4) T: Switching period
Figure 11. Main switches with Clamp switches
Figure 12. Switching Diagram

## Slide 10
Comparing Current Output
Chart: frequency [hz]
| 290 | 500 | 1000 | 1500 | 2000 | 2500 | 3000 | 3500 | 4000 | 4500 | 5000 | 5500 | 6000 | 6500 | 7000 | 7500 | 8000 | 8500 | 9000 | 9500 | 10000 | 10500 | 11000 | 11500 | 12000
50 | 0.05 | 0.09 | 0.185 | 0.281 | 0.375 | 0.486 | 1.323 | 2.795 | 4.756 | 6.846 | 9.125 | 11.504 | 13.94 | 16.455 | 18.835 | 21.362 | 23.784 | 26.111 | 28.784 | 31.179 | 33.104 | 35.495 | 38.35 | 41.063
100 | 0.204 | 0.348 | 0.701 | 1.583 | 5.169 | 9.755 | 14.715 | 19.868 | 25.093 | 30.334 | 35.548 | 40.725 | 45.953 | 51.113 | 56.021 | 61.172 | 66.219 | 70.417 | 76.082 | 80.963 | 84.255 | 88.297 | 94.439 | 100.254
150 | 0.445 | 0.769 | 1.895 | 7.708 | 15.072 | 22.905 | 30.917 | 38.948 | 47.009 | 54.945 | 62.818 | 70.584 | 78.463 | 86.122 | 93.365 | 100.948 | 108.484 | 114.345 | 122.948 | 130.228 | 134.187 | 139.915 | 148.323 | 157.763
200 | 0.803 | 1.348 | 5.928 | 15.108 | 25.308 | 35.758 | 46.282 | 56.788 | 67.285 | 77.625 | 87.785 | 97.896 | 108.226 | 118.164 | 127.558 | 137.548 | 147.328 | 154.575 | 165.752 | 175.065 | 179.872 | 187.108 | 198.505 | 210.228
250 | 1.252 | 2.085 | 10.407 | 22.242 | 34.836 | 47.615 | 60.398 | 73.156 | 85.908 | 98.452 | 111.315 | 123.592 | 134.806 | 146.672 | 158.105 | 169.742 | 180.512 | 189.511 | 202.325 | 213.118 | 218.345 | 226.532 | 239.798 | 254.958
300 | 1.757 | 3.075 | 14.495 | 28.525 | 43.125 | 57.835 | 72.528 | 87.174 | 101.839 | 116.324 | 130.698 | 144.872 | 159.372 | 173.317 | 186.596 | 200.355 | 214.165 | 224.989 | 239.278 | 252.222 | 258.223 | 267.0017 | 282.226 | 298.518
300 A
Linear current operation
300 A output reached, even with load limitation
And more!
ZOOM
Figure 13. Current Output of the Preliminary Version

## Slide 11
Chart: Frequency [hz]
| 290 | 500 | 1000 | 1500 | 2000 | 2500 | 3000 | 3500 | 4000 | 4500 | 5000 | 5500 | 6000 | 6500 | 7000 | 7500 | 8000 | 8500 | 9000 | 9500 | 10000 | 10500 | 11000 | 11500 | 12000
50 | 0.05 | 0.09 | 0.185 | 0.281 | 0.375 | 0.486 | 1.323 | 2.795 | 4.756 | 6.846 | 9.125 | 11.504 | 13.94 | 16.455 | 18.835 | 21.362 | 23.784 | 26.111 | 28.784 | 31.179 | 33.104 | 35.495 | 38.35 | 41.063
100 | 0.204 | 0.348 | 0.701 | 1.583 | 5.169 | 9.755 | 14.715 | 19.868 | 25.093 | 30.334 | 35.548 | 40.725 | 45.953 | 51.113 | 56.021 | 61.172 | 66.219 | 70.417 | 76.082 | 80.963 | 84.255 | 88.297 | 94.439 | 100.254
150 | 0.445 | 0.769 | 1.895 | 7.708 | 15.072 | 22.905 | 30.917 | 38.948 | 47.009 | 54.945 | 62.818 | 70.584 | 78.463 | 86.122 | 93.365 | 100.948 | 108.484 | 114.345 | 122.948 | 130.228 | 134.187 | 139.915 | 148.323 | 157.763
200 | 0.803 | 1.348 | 5.928 | 15.108 | 25.308 | 35.758 | 46.282 | 56.788 | 67.285 | 77.625 | 87.785 | 97.896 | 108.226 | 118.164 | 127.558 | 137.548 | 147.328 | 154.575 | 165.752 | 175.065 | 179.872 | 187.108 | 198.505 | 210.228
250 | 1.252 | 2.085 | 10.407 | 22.242 | 34.836 | 47.615 | 60.398 | 73.156 | 85.908 | 98.452 | 111.315 | 123.592 | 134.806 | 146.672 | 158.105 | 169.742 | 180.512 | 189.511 | 202.325 | 213.118 | 218.345 | 226.532 | 239.798 | 254.958
300 | 1.757 | 3.075 | 14.495 | 28.525 | 43.125 | 57.835 | 72.528 | 87.174 | 101.839 | 116.324 | 130.698 | 144.872 | 159.372 | 173.317 | 186.596 | 200.355 | 214.165 | 224.989 | 239.278 | 252.222 | 258.223 | 267.0017 | 282.226 | 298.518
Minimum Current Output
Figure 14. Minimum Current Output of the Preliminary Version
This minimum value will still decrease, since
PFC Implementation Using the superconducting coil
<2A

## Slide 12
Ripple of Output Current
In addition, an oscillation of more than 10 ppm was observed during operation, due to the low current in the DC-Link supplied by Magna Power.
However, this problem should be solved in the future with the implementation of the Feed Forward in the resonant converter control code.
Figure 15. Ripple with 300A of output current and 300V of DC-Link
Figure 16. Ripple oscillation

## Slide 14
Efficiency and Losses
Chart: FREQUENCY [Hz]
| 290 | 500 | 1000 | 1500 | 2000 | 2500 | 3000 | 3500 | 4000 | 4500 | 5000 | 5500 | 6000 | 6500 | 7000 | 7500 | 8000 | 8500 | 9000 | 9500 | 10000 | 10500 | 11000 | 11500
50 | 0.0025714 | 0.0046286 | 0.00925 | 0.01124 | 0.015 | 0.00972 | 0.022389 | 0.055517 | 0.057462 | 0.10187 | 0.12623 | 0.19816 | 0.2267 | 0.27812 | 0.3018 | 0.34043 | 0.37541 | 0.41524 | 0.46248 | 0.42849 | 0.52983 | 0.43418 | 0.44486 | 0.50463
100 | 0.0032931 | 0.0060154 | 0.011619 | 0.027974 | 0.065498 | 0.085897 | 0.34066 | 0.28412 | 0.44191 | 0.49396 | 0.46219 | 0.52609 | 0.59551 | 0.54369 | 0.59258 | 0.65409 | 0.59487 | 0.62912 | 0.63142 | 0.64317 | 0.64082 | 0.64099 | 0.7024 | 0.69157
150 | 0.0043652 | 0.0088618 | 0.022498 | 0.080248 | 0.15529 | 0.24338 | 0.38188 | 0.40397 | 0.52171 | 0.55195 | 0.57564 | 0.57339 | 0.5494 | 0.58747 | 0.62552 | 0.64226 | 0.66187 | 0.68839 | 0.68828 | 0.70378 | 0.7292 | 0.70767 | 0.73177 | 0.71228
200 | 0.007975 | 0.0066938 | 0.035637 | 0.12401 | 0.22976 | 0.39323 | 0.46704 | 0.51555 | 0.57082 | 0.64596 | 0.63383 | 0.65214 | 0.68046 | 0.64345 | 0.70403 | 0.67669 | 0.73608 | 0.71182 | 0.72741 | 0.76631 | 0.76666 | 0.76008 | 0.76523 | 0.76767
250 | 0.01108 | 0.01051 | 0.051904 | 0.18662 | 0.3486 | 0.49488 | 0.55759 | 0.50032 | 0.58439 | 0.66112 | 0.69661 | 0.70212 | 0.68752 | 0.70759 | 0.72767 | 0.71582 | 0.76349 | 0.7712 | 0.75504 | 0.76592 | 0.79293 | 0.78439 | 0.78209 | 0.80061
300 | 0.0098978 | 0.010952 | 0.077606 | 0.19601 | 0.34484 | 0.43292 | 0.52795 | 0.54298 | 0.60974 | 0.62173 | 0.67476 | 0.6658 | 0.69665 | 0.72521 | 0.72329 | 0.74721 | 0.75857 | 0.73557 | 0.74618 | 0.79343 | 0.76598 | 0.78791 | 0.80417 | 0.82191
Figure 17. Efficiency of the Preliminary Version
Figure 18. Loss Distribution for 300 V DC-Link
Increasing Dead Time (Clamp and main switches)
Decreasing parasitic inductance

## Slide 15
Comparing the improvements made
Preliminary Version
+ Clamp Circuit
+ Dead Time - Inductance
Figure 19. Improvements Comparison

## Slide 16
Preliminary Version – Final results and installation
Identify changes needed for the final version Make adjustments for future tests with the superconductive coil.
FINAL VERSION
Manufacture and optimize the final plate version; Increase the power density of the project; Improve installed components; Assemble the racks for delivery.
NEXT STEPS
Figure 20. Final Version tests and optimization

## Slide 17
References
[1] J. G. Galvez et al., “Preliminary design of a cryogen-free 6.6 T superconducting wavelength shifter,” IEEE Transactions on Applied Superconductivity, to be published, 2023.
[2] Y. G. Kang and A. K. Upadhyay, “Analysis and design of a half-bridge parallel resonant converter,” IEEE Transactions on Power Electronics, vol. 3, no. 2, pp. 174-182, 1988.
[3] K.-H. Liu and F. C. Lee, “Secondary-side resonance for high-frequency power conversion,” in 1986 IEEE Applied Power Electronics Conference and Exposition, pp. 83-89, 1986.
[4] G. O. Brunheira, J. P. Bonaldo, J. N. H. da Rosa, C. Rodrigues and J. A. Pomilio, “Digital control system for high precision power supplies of the new Brazilian synchrotron source,” in 2013 Brazilian Power Electronics Conference, pp. 326-331, 2013.
[5] W. Contesini, L. de Oliveira Porto, B. E. Limeira and F. S. de Oliveira, "Prototype of a Resonant Converter for a Superconducting Magnet with 300 A and 10 V Output Operating in Discontinuous Conduction Mode," 2023 IEEE 8th Southern Power Electronics Conference and 17th Brazilian Power Electronics Conference (SPEC/COBEP), Florianopolis, Brazil, 2023, pp. 1-8, doi: 10.1109/SPEC56436.2023.10408187.

## Slide 18
Lucas Carnevalli de Almeida
lucas.carnevalli@cnpem.br
THANK YOU